## Worksheet
|  | Johnson Gage and Inspection, Inc. |  | Page 2 of 2 | [13 empty] | INSPECTOR NOTE: PRINT PAGE 1 ONLY | [6 empty] | µF | µΩ | °C
| Supplementary Data Sheet to Report # | [16 empty] | This section is for Creation of Datasheet Only | [6 empty] | µV | @ | °F
| Hioki BT3563-01 |  | Date: | [11 empty] | Secondary Unit for Range field |  |  |  |  | Note: You can now use "O" instead of Ω, "DGC" instead |  |  |  |  | µA |  |  | RESULT | Special Character Filter
[10 empty] | REPEAT READINGS | [9 empty] | of "°C", and "DGF" instead of "°F" if you want to. | [7 empty] | CONV.
Function | Range | Signal Applied | Result | Min | Max | Uncertainty @k=2 | Coil Used | Meter Used | Notes: | 1 | 2 | 3 | 4 | 5 |  | Resolution | Range | Primary Unit | DC/AC | Nominal | tolerance | max | min | (leave blank) | Frequency value | Secondary Unit |  | TO NUMBER
Resistance | 3 mΩ | 1 mΩ | 1 | 0.9940mΩ | 1.0060mΩ | * |  | 34420A | [7 empty] | 0.0000 | 3 | mΩ |  | 1 | 0.006 | 1.0060 | 0.9940 |  |  |  |  | 1 | mΩ | 1 | 1.006 | 0.994 | 1 | 1 | 1
| 30 mΩ | 20 mΩ | 20 | 19.895mΩ | 20.105mΩ | 3 µΩ |  | 34420A | [7 empty] | 0.000 | 30 | mΩ |  | 20 | 0.105 | 20.105 | 19.895 |  |  |  |  | 20 | mΩ | 20 | 20.105 | 19.895 | 1 | 1 | 1
| 300 mΩ | 200 mΩ | 200 | 198.95mΩ | 201.05mΩ | 4 µΩ |  | 34420A | [7 empty] | 0.00 | 300 | mΩ |  | 200 | 1.05 | 201.05 | 198.95 |  |  |  |  | 200 | mΩ | 200 | 201.05 | 198.95 | 1 | 1 | 1
| 3 Ω | 2 Ω | 2 | 1.9895Ω | 2.0105Ω | 64 µΩ |  | 34420A | [7 empty] | 0.0000 | 3 | Ω |  | 2 | 0.0105 | 2.0105 | 1.9895 |  |  |  |  | 2 | Ω | 2 | 2.0105 | 1.9895 | 1 | 1 | 1
| 30 Ω | 20 Ω | 20 | 19.895Ω | 20.105Ω | 600 µΩ |  | 34420A | [7 empty] | 0.000 | 30 | Ω |  | 20 | 0.105 | 20.105 | 19.895 |  |  |  |  | 20 | Ω | 20 | 20.105 | 19.895 | 1 | 1 | 1
| 300 Ω | 200 Ω | 200 | 198.95Ω | 201.05Ω | 5.9 mΩ |  | 34420A | [7 empty] | 0.00 | 300 | Ω |  | 200 | 1.05 | 201.05 | 198.95 |  |  |  |  | 200 | Ω | 200 | 201.05 | 198.95 | 1 | 1 | 1
| 3000 Ω | 2000 Ω | 2000 | 1989.5Ω | 2010.5Ω | 59 mΩ |  | 34420A | [7 empty] | 0.0 | 3000 | Ω |  | 2000 | 10.5 | 2010.5 | 1989.5 |  |  |  |  | 2000 | Ω | 2000 | 2010.5 | 1989.5 | 1 | 1 | 1
Voltage | 6 V | 5 VDC | 5 | 4.99947VDC | 5.00053VDC | 0.3 mV | [9 empty] | 0.00000 | 6 | V | DC | 5 | 0.00053 | 5.00053 | 4.99947 |  |  |  |  | 5 | V | 5 | 5.0005 | 4.9995 | 1 | 1 | 1
| 60 V | 50 VDC | 50 | 49.9947VDC | 50.0053VDC | 3.9 mV | [9 empty] | 0.0000 | 60 | V | DC | 50 | 0.0053 | 50.0053 | 49.9947 |  |  |  |  | 50 | V | 50 | 50.0053 | 49.9947 | 1 | 1 | 1
| 600 V | 500 VDC | 500 | 499.947VDC | 500.053VDC | 34 mV | [9 empty] | 0.000 | 600 | V | DC | 500 | 0.053 | 500.053 | 499.947 |  |  |  |  | 500 | V | 500 | 500.053 | 499.947 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1

## Unc. Calculator
RED=Key Calculator Component |  | 5502A | Standard1
GREEN=Formula Uses a Table Lookup |  | 3458A | Standard2 |  | Troubleshooter Lookup | [11 empty] | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | [38 empty] | ACV Freq. Rng Low | ACI Freq. Rng Low | ACA10T. Rng low | ACA20T. Rng low | ACA50T. Rng Low
ORANGE= END USER RAW INPUTS |  | 34420A | Standard3 |  | Rdg | L.up ID | Standard | Base | Mult | 2nd B | 2nd Mult | Rep. Mult | Unc. Units | Result |  | Standard6 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [13 empty] | 1000000 | [23 empty] | Standard1 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719
BLUE= PART OF MAIN LOOKUP ID |  | Transmille | Standard4 |  | 0.003 | DCA0.000010 | Standard1 | 5.8001 | 1.6239 | 5.8001 | 1.5456 | 0 | µA | 5.8049717µA |  | Standard5 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [37 empty] | Standard2 | #N/A | $I$821:$I$1000 | #N/A | #N/A | #N/A
|  |  | Standard5 | [12 empty] | Standard4 | $D$4:$D$28 | $D$80:$D$119 | $D$29:$D$58 | $D$120:$D$149 | #N/A | $D$150:$D$164 | $D$165:$D$169 | $D$59:$D$64 | $D$65:$D$69 | $D$70:$D$79 | #N/A | #N/A | #N/A | [37 empty] | Standard3 | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  | Standard6 | [12 empty] | Standard3 | #N/A | #N/A | #N/A | #N/A | $D$39:$D$73 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [37 empty] | Standard4 | $I$80:$I$119 | $I$120:$I$149 | #N/A | #N/A | #N/A | Col. Index # for frequency Lookup | Input Unit | Unc units
|  |  | Standard7 | [12 empty] | Standard2 | $D$33:$D$61 | #N/A | $D$115:$D$167 | $D$821:$D$1000 | $D$168:$D$221 | #N/A | $D$1001:$D$1002 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Unit Multiplier | [13 empty] | 2ndary Unit Multiplier | [22 empty] | Standard5 | #N/A | #N/A | #N/A | #N/A | #N/A |  | V11 | CN11
[16 empty] | Standard1 | $D$4:$D$32 | $D$189:$D$348 | $D$33:$D$66 | $D$349:$D$538 | $D$121:$D$188 | $D$720:$D$785 | $D$786:$D$809 | $D$67:$D$87 | $D$88:$D$104 | $D$105:$D$120 | $D$539:$D$609 | $D$610:$D$659 | $D$660:$D$719 | Base unit | [13 empty] | Base Unit | [18 empty] | RESOLUTION CONVERTER |  |  |  | Standard6 | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  | USER INPUTS |  |  |  |  | Function Selector | [8 empty] | MAIN INPUT FUNCTION LOOKUP | [12 empty] | V | V | Amps | Amps | kOhms | mF | kHz | Amps | Amps | Amps | Amps | Amps | Amps |  |  |  |  | LOOKED UP UNC. COMPONENTS |  |  |  |  | REPEAT USER INPUTS |  |  |  |  | CONVERTED REPEAT ENTRIES | [5 empty] | "0.000 to 0.001" | [10 empty] | Repeat Multiplier Ratio
Main Lookup ID | Standard Used | Unc. | Res. | In. Value | Units | Range 2 | Units | L.upbldr | L.upbldr | L.upbldr | L.upbldr | L.upbldr | L.upbldr | Func.Sel. | Main L.up | Converted Rdg | DCV | ACV | DCA | ACA | Resistance | Capacitance | Frequency | DCA 10 TURN | DCA 20 TURN | DCA 50 TURN | ACA 10 TURN | ACA 20 TURN | ACA 50 TURN | DCV | ACV | DCA | ACA | Resistance | Capacitance | Frequency | DCA 10 TURN | DCA 20 TURN | DCA 50 TURN | ACA 10 TURN | ACA 20 TURN | ACA 50 TURN | Primary Conv. Factor | kHz | Converted Rdg | 2ndary L. Up | Base | Mult | 2nd B. | 2nd Mult. | Rept Mult | R1 | R2 | R3 | R4 | R5 | R1(Mult) | R2(Mult) | R3(Mult) | R4(Mult) | R5(Mult) | Unc. Units | Input | Conv. Txt | Result |  | Unc. Table Address for L. Ups | ACV | ACA | ACA 10 TURN | ACA 20 TURN | ACA 50 TURN |  | Input | Uncert. | Multiplier
Resistance0.00000000010 | 34420A | * | 1e-10 | 1 | mΩ |  |  | m | O | DC | O |  | Resistance | Resistance | 0 | 1e-06 | #N/A | #N/A | #N/A | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1e-06 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.0000 | 0.0001 | 1e-10 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 2 | 0.001 | #N/A
Resistance0.0000000010 | 34420A | 3 µΩ | 1e-09 | 20 | mΩ |  |  | m | O | DC | O |  | Resistance | Resistance | 0 | 2e-05 | #N/A | #N/A | #N/A | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1e-06 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 |  |  | 3 | 120 | 2.3273 | 80.8131 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µΩ | 0.000 | 0.001 | 1e-09 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | 0.001 | 1e-06
Resistance0.000000010 | 34420A | 4 µΩ | 1e-08 | 200 | mΩ |  |  | m | O | DC | O |  | Resistance | Resistance | 0 | 0.0002 | #N/A | #N/A | #N/A | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1e-06 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 |  |  | 3 | 120 | 3.9637 | 79.3768 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µΩ | 0.00 | 0.01 | 1e-08 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 4 | 0.001 | 1e-06
Resistance0.00000010 | 34420A | 64 µΩ | 1e-07 | 2 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0 | 0.002 | #N/A | #N/A | #N/A | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 63.6928 | 62.6594 | 54.2074 | 47.1632 | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µΩ | 0.0000 | 0.0001 | 1e-07 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 5 | 1 | 1e-06
Resistance0.0000010 | 34420A | 600 µΩ | 1e-06 | 20 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0 | 0.02 | #N/A | #N/A | #N/A | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 600.322 | 96.4808 | 579.4447 | 6.4836 | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µΩ | 0.000 | 0.001 | 1e-06 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 6 | 1 | 1e-06
Resistance0.000010 | 34420A | 5.9 mΩ | 1e-05 | 200 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0 | 0.2 | #N/A | #N/A | #N/A | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 5.8513 | 0.24805 | 5.7999 | 0.00065169 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.00 | 0.01 | 1e-05 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 7 | 1 | 0.001
Resistance0.00011 | 34420A | 59 mΩ | 0.0001 | 2000 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 1 | 2 | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 58.5135 | 0.24789 | 57.9996 | 0.00048274 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.0 | 0.1 | 0.0001 | Standard3 | $B$4:$P$72 | #N/A | #N/A | #N/A | #N/A | #N/A | 8 | 1 | 0.001
DCV0.000013 | 5502A | 0.3 mV | 1e-05 | 5 | V |  |  |  | V | DC | DCV |  |  | DCV | 3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.007 | 0.058 | 0.0059614 | 0.05773 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.00000 | 0.00001 | 1e-05 | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 9 | 1 | 0.001
DCV0.000133 | 5502A | 3.9 mV | 0.0001 | 50 | V |  |  |  | V | DC | DCV |  |  | DCV | 33 | 50 | 33 | 33 | 11 | 11 | 33 | 33 | 12 | 30 | 20 | #N/A | 30 | 22 | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.75 | 0.063 | 0.57872 | 0.063505 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.0000 | 0.0001 | 0.0001 | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 10 | 1 | 0.001
DCV0.001330 | 5502A | 34 mV | 0.001 | 500 | V |  |  |  | V | DC | DCV |  |  | DCV | 330 | 500 | 330 | 330 | 11 | 11 | 330 | 33 | 120 | 110 | 220 | 150 | 110 | 220 | 150 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 2.4 | 0.063 | 1.7517 | 0.063494 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 11 | 1 | 0.001
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 12 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 13 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 14 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 15 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 16 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 17 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 18 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 19 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 20 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 21 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 22 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 23 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 24 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 25 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 26 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 27 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 28 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 29 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 30 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 31 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 32 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 33 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 34 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 35 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 36 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 37 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 38 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 39 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 40 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 41 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 42 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 43 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 44 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 45 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 46 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 47 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 48 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 49 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 50 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 51 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 52 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 53 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 54 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 55 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 56 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 57 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 58 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 59 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 60 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 61 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 62 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 63 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 64 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 65 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 66 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 67 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 68 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 69 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 70 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 71 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 72 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 73 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 74 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 75 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 76 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 77 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 78 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 79 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 80 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 81 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 82 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 83 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 84 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 85 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 86 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 87 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 88 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 89 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 90 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$807 | $I$189:$I$348 | $I$349:$I$538 | $I$539:$I$609 | $I$610:$I$659 | $I$660:$I$719 | 91 |  | #N/A
Note on S1: Lookup table thatTranslates Standard Name to Generic name. i.e. "5502A" to "Standard1". See Colum "CQ"
Note on BA7: Columns AU-BA convert the reading entered into the standard unit of the table for each function
Note on BV7: Converts the 2ndary units (usually frequency) to the standard table format (kHz)
Note on T9: User inputs. Resolution is converted from "0.0" format to "0.1" format in Column CA-CE, and Column T references that
Note on Y9: Columns Y-AD select which function (column C) is used based on the inputs
Note on AH9: Columns AG- AM determine in what range the primary unit falls (Ex: 5VDC falls in the 3.3-30V Range, so it returns 3.3, as it is set to show the "Range Low" part as a marker
Note on DB9: This is where the ratio between the Input units and the Uncertainty Units is created to generate the repeat multiplier in CC11
Note on Q10: Constructed input ID in the form of Function,resolution,InputRangeLOW,InputRange2LOW
Note on T10: From Column CE
Note on U10: Returns the Absolute value of the reading (filters out Negative Numbers)
Note on Z10: This is where special characters get filtered out and converted to whatever we designate. (i.e. Ω changes to O)
Note on AE10: Determines which category in Column "C" the reading falls under
Note on AF10: Column AE Retrieves the appropriate "Range Low" result from Columns AG-AM, based on the Function Selector in Column AD. It is an Hlookup formula, desgined to return the Nth row down, where N is the number of rows
Note on AG10: Column AF Converts Column AE to whatever the standard unit is in the table, based on a multiplier in Column BI
Note on BU10: Looks up the correct converted Unit based on the function selector
Note on BV10: 2ndary unit conv. factor
Note on BX10: Returns the appropriate "Range 2 Low" value (usually frequency) from the table. Part of the INPUT ID that looks up the Uncertainty components
Note on CQ10: Final Resolution Input: Part of the INPUT ID used for uncertainty Lookup

## Standard1
Updated: | 2018-04-08
|  | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.00000010 | DCV | 0 | 0.33 | 1e-07 | V | µV |  |  |  | 3.6 | 69 | 3.4651 | 69.28
| DCV0.0000010 | DCV | 0 | 0.33 | 1e-06 | V | µV |  |  |  | 3.6 | 69.116 | 3.5122 | 69.156
| DCV0.000010 | DCV | 0 | 0.33 | 1e-05 | V | µV |  |  |  | 6.77 | 61.242 | 6.7306 | 61.297
| DCV0.00010 | DCV | 0 | 0.33 | 0.0001 | V | µV |  |  |  | 58.109 | 17.038 | 58.091 | 16.985
| DCV0.0010 | DCV | 0 | 0.33 | 0.001 | V | µV |  |  |  | 580.02 | 1.808 | 580.01 | 1.7825
| DCV0.010 | DCV | 0 | 0.33 | 0.01 | V | µV |  |  |  | 5800 | 0.18855 | 5800 | 0.17835
| DCV0.0000010.33 | DCV | 0.33 | 3.3 | 1e-06 | V | µV |  |  |  | 5.9 | 58 | 5.7885 | 57.731
| DCV0.000010.33 | DCV | 0.33 | 3.3 | 1e-05 | V | µV |  |  |  | 6.8245 | 57.473 | 6.514 | 57.537
| DCV0.00010.33 | DCV | 0.33 | 3.3 | 0.0001 | V | µV |  |  |  | 47.838 | 47.647 | 47.36 | 47.676
| DCV0.0010.33 | DCV | 0.33 | 3.3 | 0.001 | V | µV |  |  |  | 577.17 | 10.806 | 577 | 10.703
| DCV0.010.33 | DCV | 0.33 | 3.3 | 0.01 | V | µV |  |  |  | 5799.8 | 1.1454 | 5799.7 | 1.1003
| DCV0.10.33 | DCV | 0.33 | 3.3 | 0.1 | V | µV |  |  |  | 58000 | 0.12814 | 5799.8 | 0.11006
| DCV0.000013 | DCV | 3 | 33 | 1e-05 | V | mV |  |  |  | 0.007 | 0.058 | 0.0059614 | 0.05773
| DCV0.00013 | DCV | 3 | 33 | 0.0001 | V | mV |  |  |  | 0.018974 | 0.057393 | 0.015087 | 0.05748
| DCV0.0013 | DCV | 3 | 33 | 0.001 | V | mV |  |  |  | 0.46341 | 0.046595 | 0.45849 | 0.046625
| DCV0.013 | DCV | 3 | 33 | 0.01 | V | mV |  |  |  | 5.7713 | 0.010319 | 5.7696 | 0.010216
| DCV0.13 | DCV | 3 | 33 | 0.1 | V | mV |  |  |  | 57.998 | 0.0010937 | 57.997 | 0.0010486
| DCV13 | DCV | 3 | 33 | 1 | V | mV |  |  |  | 580 | 0.00012297 | 580 | 0.00010489
| DCV0.000133 | DCV | 33 | 330 | 0.0001 | V | mV |  |  |  | 0.75 | 0.063 | 0.57872 | 0.063505
| DCV0.00133 | DCV | 33 | 330 | 0.001 | V | mV |  |  |  | 0.75 | 0.063263 | 0.64626 | 0.063324
| DCV0.0133 | DCV | 33 | 330 | 0.01 | V | mV |  |  |  | 4.6657 | 0.053552 | 4.6183 | 0.053589
| DCV0.133 | DCV | 33 | 330 | 0.1 | V | mV |  |  |  | 57.656 | 0.012921 | 57.639 | 0.012819
| DCV133 | DCV | 33 | 330 | 1 | V | mV |  |  |  | 579.97 | 0.0013701 | 579.96 | 0.0013249
| DCV1033 | DCV | 33 | 330 | 10 | V | mV |  |  |  | 5800 | 0.00015062 | 5800 | 0.00013254
| DCV0.001330 | DCV | 330 | 1000 | 0.001 | V | mV |  |  |  | 2.4 | 0.063 | 1.7517 | 0.063494
| DCV0.01330 | DCV | 330 | 1000 | 0.01 | V | mV |  |  |  | 3.0262 | 0.062773 | 2.7036 | 0.062797
| DCV0.1330 | DCV | 330 | 1000 | 0.1 | V | mV |  |  |  | 50.232 | 0.037958 | 49.963 | 0.037334
| DCV1330 | DCV | 330 | 1000 | 1 | V | mV |  |  |  | 578.9 | 0.0052931 | 578.86 | 0.0047972
| DCV10330 | DCV | 330 | 1000 | 10 | V | mV |  |  |  | 5799.9 | 0.00068179 | 5799.9 | 0.00048139
| DCA0.0000000010 | DCA | 0 | 0.00033 | 1e-09 | A | µA |  |  |  | 0.025 | 170 | 0.02311 | 173.17
| DCA0.000000010 | DCA | 0 | 0.00033 | 1e-08 | A | µA |  |  |  | 0.025 | 171.65 | 0.023816 | 171.66
| DCA0.00000010 | DCA | 0 | 0.00033 | 1e-07 | A | µA |  |  |  | 0.062455 | 111.22 | 0.062385 | 111.01
| DCA0.0000010 | DCA | 0 | 0.00033 | 1e-06 | A | µA |  |  |  | 0.58048 | 15.571 | 0.58045 | 15.378
| DCA0.000010 | DCA | 0 | 0.00033 | 1e-05 | A | µA |  |  |  | 5.8001 | 1.6239 | 5.8001 | 1.5456
| DCA0.00010 | DCA | 0 | 0.00033 | 0.0001 | A | µA |  |  |  | 58 | 0.1859 | 58 | 0.15457
| DCA0.000000010.00033 | DCA | 0.00033 | 0.0033 | 1e-08 | A | µA |  |  |  | 0.059 | 150 | 0.058075 | 150.03
| DCA0.00000010.00033 | DCA | 0.00033 | 0.0033 | 1e-07 | A | µA |  |  |  | 0.076723 | 145.61 | 0.07387 | 146.15
| DCA0.0000010.00033 | DCA | 0.00033 | 0.0033 | 1e-06 | A | µA |  |  |  | 0.56851 | 71.494 | 0.56635 | 71.248
| DCA0.000010.00033 | DCA | 0.00033 | 0.0033 | 1e-05 | A | µA |  |  |  | 5.799 | 8.7697 | 5.7982 | 8.5256
| DCA0.00010.00033 | DCA | 0.00033 | 0.0033 | 0.0001 | A | µA |  |  |  | 58 | 0.95318 | 58 | 0.85455
| DCA0.00000010.0033 | DCA | 0.0033 | 0.033 | 1e-07 | A | µA |  |  |  | 0.2 | 150 | 0.1928 | 146.04
| DCA0.0000010.0033 | DCA | 0.0033 | 0.033 | 1e-06 | A | µA |  |  |  | 0.43579 | 139.77 | 0.42593 | 139.98
| DCA0.000010.0033 | DCA | 0.0033 | 0.033 | 1e-05 | A | µA |  |  |  | 5.6412 | 61.573 | 5.6359 | 61.499
| DCA0.00010.0033 | DCA | 0.0033 | 0.033 | 0.0001 | A | µA |  |  |  | 57.982 | 7.2082 | 57.98 | 7.1459
| DCA0.0010.0033 | DCA | 0.0033 | 0.033 | 0.001 | A | µA |  |  |  | 580 | 0.74105 | 580 | 0.71593
| DCA0.0000010.033 | DCA | 0.033 | 0.33 | 1e-06 | A | µA |  |  |  | 3 | 120 | 2.9414 | 115.33
| DCA0.000010.033 | DCA | 0.033 | 0.33 | 1e-05 | A | µA |  |  |  | 5.2852 | 109.48 | 5.266 | 109.51
| DCA0.00010.033 | DCA | 0.033 | 0.33 | 0.0001 | A | µA |  |  |  | 56.995 | 42.585 | 56.984 | 42.551
| DCA0.0010.033 | DCA | 0.033 | 0.33 | 0.001 | A | µA |  |  |  | 579.89 | 4.761 | 579.88 | 4.7411
| DCA0.010.033 | DCA | 0.033 | 0.33 | 0.01 | A | µA |  |  |  | 5800 | 0.4827 | 5800 | 0.47471
| DCA0.000010.33 | DCA | 0.33 | 3 | 1e-05 | A | mA |  |  |  | 0.052 | 0.44 | 0.051 | 0.43873
| DCA0.00010.33 | DCA | 0.33 | 3 | 0.0001 | A | mA |  |  |  | 0.060248 | 0.43614 | 0.060204 | 0.43607
| DCA0.0010.33 | DCA | 0.33 | 3 | 0.001 | A | mA |  |  |  | 0.50421 | 0.32729 | 0.50424 | 0.32696
| DCA0.010.33 | DCA | 0.33 | 3 | 0.01 | A | mA |  |  |  | 5.784 | 0.058643 | 5.7841 | 0.058299
| DCA0.10.33 | DCA | 0.33 | 3 | 0.1 | A | mA |  |  |  | 57.998 | 0.0060523 | 57.998 | 0.0059106
| DCA0.00013 | DCA | 3 | 11 | 0.0001 | A | mA |  |  |  | 0.61 | 0.69 | 0.57897 | 0.69271
| DCA0.0013 | DCA | 3 | 11 | 0.001 | A | mA |  |  |  | 0.65926 | 0.68757 | 0.6565 | 0.68751
| DCA0.013 | DCA | 3 | 11 | 0.01 | A | mA |  |  |  | 5.0083 | 0.45871 | 5.0057 | 0.45792
| DCA0.13 | DCA | 3 | 11 | 0.1 | A | mA |  |  |  | 57.868 | 0.065168 | 57.867 | 0.064474
| DCA0.000111 | DCA | 11 | 20.5 | 0.0001 | A | mA |  |  |  | 0.87 | 1.2 | 0.86628 | 1.1547
| DCA0.00111 | DCA | 11 | 20.5 | 0.001 | A | mA |  |  |  | 0.88429 | 1.1544 | 0.88526 | 1.1541
| DCA0.0111 | DCA | 11 | 20.5 | 0.01 | A | mA |  |  |  | 2.6424 | 1.1019 | 2.6473 | 1.1008
| DCA0.111 | DCA | 11 | 20.5 | 0.1 | A | mA |  |  |  | 55.6 | 0.36123 | 55.616 | 0.35906
| DCA 10 turn0.00011 | DCA 10 turn | 1 | 11 | 0.0001 | A | mA |  |  |  | 56 | 3.4 | 55.012 | 3.3767
| DCA 10 turn0.0011 | DCA 10 turn | 1 | 11 | 0.001 | A | mA |  |  |  | 56 | 3.3729 | 55.236 | 3.3728
| DCA 10 turn0.011 | DCA 10 turn | 1 | 11 | 0.01 | A | mA |  |  |  | 75.465 | 3.0449 | 75.403 | 3.0441
| DCA 10 turn0.11 | DCA 10 turn | 1 | 11 | 0.1 | A | mA |  |  |  | 579.54 | 0.70852 | 579.51 | 0.70625
| DCA 10 turn11 | DCA 10 turn | 1 | 11 | 1 | A | mA |  |  |  | 5800 | 0.073274 | 5799.9 | 0.07231
| DCA 10 turn0.00111 | DCA 10 turn | 11 | 30 | 0.001 | A | mA |  |  |  | 3.3 | 2.4 | 0.024561 | 2.0002
| DCA 10 turn0.0111 | DCA 10 turn | 11 | 30 | 0.01 | A | mA |  |  |  | 3.3 | 2.4 | 0.71231 | 1.9865
| DCA 10 turn0.111 | DCA 10 turn | 11 | 30 | 0.1 | A | mA |  |  |  | 44.41 | 1.3054 | 44.468 | 1.3001
| DCA 10 turn111 | DCA 10 turn | 11 | 30 | 1 | A | mA |  |  |  | 577.9 | 0.17531 | 577.94 | 0.17187
| DCA 10 turn1011 | DCA 10 turn | 11 | 30 | 10 | A | mA |  |  |  | 5799.8 | 0.018636 | 5799.8 | 0.017253
| DCA 10 turn10011 | DCA 10 turn | 11 | 30 | 100 | A | mA |  |  |  | 58000 | 0.0022788 | 58000 | 0.0017254
| DCA 10 turn0.0130 | DCA 10 turn | 30 | 110 | 0.01 | A | mA |  |  |  | 6 | 2.4 | 0.90569 | 1.9962
| DCA 10 turn0.130 | DCA 10 turn | 30 | 110 | 0.1 | A | mA |  |  |  | 14.114 | 2.3253 | 29.791 | 1.8011
| DCA 10 turn130 | DCA 10 turn | 30 | 110 | 1 | A | mA |  |  |  | 569.15 | 0.46761 | 569.15 | 0.46649
| DCA 10 turn1030 | DCA 10 turn | 30 | 110 | 10 | A | mA |  |  |  | 5798.9 | 0.048871 | 5798.9 | 0.048384
| DCA 10 turn10030 | DCA 10 turn | 30 | 110 | 100 | A | mA |  |  |  | 58000 | 0.0050353 | 58000 | 0.0048403
| DCA 10 turn0.01110 | DCA 10 turn | 110 | 205 | 0.01 | A | mA |  |  |  | 19 | 2.4 | 0.88362 | 2.0021
| DCA 10 turn0.1110 | DCA 10 turn | 110 | 205 | 0.1 | A | mA |  |  |  | 19 | 2.4 | 12.193 | 1.9666
| DCA 10 turn1110 | DCA 10 turn | 110 | 205 | 1 | A | mA |  |  |  | 516.04 | 0.95327 | 516.1 | 0.95081
| DCA 10 turn10110 | DCA 10 turn | 110 | 205 | 10 | A | mA |  |  |  | 5792.2 | 0.11031 | 5792.2 | 0.10898
| DCA 10 turn100110 | DCA 10 turn | 110 | 205 | 100 | A | mA |  |  |  | 57999 | 0.011451 | 57999 | 0.010915
| DCA 20 turn0.00120 | DCA 20 turn | 20 | 60 | 0.001 | A | mA |  |  |  | 5.9 | 2.4 | 0.069752 | 2.0016
| DCA 20 turn0.0120 | DCA 20 turn | 20 | 60 | 0.01 | A | mA |  |  |  | 5.9 | 2.4 | 0.61737 | 1.9948
| DCA 20 turn0.120 | DCA 20 turn | 20 | 60 | 0.1 | A | mA |  |  |  | 30.889 | 1.9835 | 39.05 | 1.573
| DCA 20 turn120 | DCA 20 turn | 20 | 60 | 1 | A | mA |  |  |  | 575.93 | 0.27416 | 575.92 | 0.27332
| DCA 20 turn1020 | DCA 20 turn | 20 | 60 | 10 | A | mA |  |  |  | 5799.6 | 0.027997 | 5799.6 | 0.027652
| DCA 20 turn10020 | DCA 20 turn | 20 | 60 | 100 | A | mA |  |  |  | 58000 | 0.0029038 | 58000 | 0.0027655
| DCA 20 turn0.00160 | DCA 20 turn | 60 | 220 | 0.001 | A | mA |  |  |  | 19 | 2.4 | 1.6742 | 2.0017
| DCA 20 turn0.0160 | DCA 20 turn | 60 | 220 | 0.01 | A | mA |  |  |  | 19 | 2.4 | 18.119 | 1.944
| DCA 20 turn0.160 | DCA 20 turn | 60 | 220 | 0.1 | A | mA |  |  |  | 541.39 | 0.85539 | 541.39 | 0.8538
| DCA 20 turn160 | DCA 20 turn | 60 | 220 | 1 | A | mA |  |  |  | 5795.4 | 0.097995 | 5795.5 | 0.09714
| DCA 20 turn1060 | DCA 20 turn | 60 | 220 | 10 | A | mA |  |  |  | 58000 | 0.010072 | 58000 | 0.009729
| DCA 20 turn10060 | DCA 20 turn | 60 | 220 | 100 | A | mA |  |  |  | 580000 | 0.0011102 | 580000 | 0.00097292
| DCA 20 turn0.01220 | DCA 20 turn | 220 | 410 | 0.01 | A | mA |  |  |  | 19 | 2.4 | 2.5543 | 2.0124
| DCA 20 turn0.1220 | DCA 20 turn | 220 | 410 | 0.1 | A | mA |  |  |  | 19 | 2.4 | 8.2583 | 2.0034
| DCA 20 turn1220 | DCA 20 turn | 220 | 410 | 1 | A | mA |  |  |  | 407.64 | 1.4723 | 407.64 | 1.4707
| DCA 20 turn10220 | DCA 20 turn | 220 | 410 | 10 | A | mA |  |  |  | 5768.7 | 0.22084 | 5768.8 | 0.2194
| DCA 20 turn100220 | DCA 20 turn | 220 | 410 | 100 | A | mA |  |  |  | 57997 | 0.022667 | 57997 | 0.022081
| DCA 50 turn0.00155 | DCA 50 turn | 55 | 150 | 0.001 | A | mA |  |  |  | 5.7 | 2.4 | 1.4795 | 1.9907
| DCA 50 turn0.0155 | DCA 50 turn | 55 | 150 | 0.01 | A | mA |  |  |  | 5.7 | 2.4 | 1.685 | 1.9897
| DCA 50 turn0.155 | DCA 50 turn | 55 | 150 | 0.1 | A | mA |  |  |  | 5.7 | 2.4 | 20.758 | 1.8992
| DCA 50 turn155 | DCA 50 turn | 55 | 150 | 1 | A | mA |  |  |  | 554.35 | 0.65863 | 554.33 | 0.65803
| DCA 50 turn1055 | DCA 50 turn | 55 | 150 | 10 | A | mA |  |  |  | 5797.2 | 0.070784 | 5797.2 | 0.070489
| DCA 50 turn10055 | DCA 50 turn | 55 | 150 | 100 | A | mA |  |  |  | 58000 | 0.0071727 | 57996 | 0.0070542
| DCA 50 turn0.01150 | DCA 50 turn | 150 | 550 | 0.01 | A | mA |  |  |  | 17 | 2.4 | 2.6539 | 1.9968
| DCA 50 turn0.1150 | DCA 50 turn | 150 | 550 | 0.1 | A | mA |  |  |  | 17 | 2.4 | 9.5683 | 1.987
| DCA 50 turn1150 | DCA 50 turn | 150 | 550 | 1 | A | mA |  |  |  | 398.66 | 1.7061 | 432.58 | 1.4759
| DCA 50 turn10150 | DCA 50 turn | 150 | 550 | 10 | A | mA |  |  |  | 5772.1 | 0.24055 | 5772 | 0.23923
| DCA 50 turn100150 | DCA 50 turn | 150 | 550 | 100 | A | mA |  |  |  | 57997 | 0.024694 | 57997 | 0.02415
| DCA 50 turn0.01550 | DCA 50 turn | 550 | 1025 | 0.01 | A | mA |  |  |  | 92 | 2.4 | 119.97 | 1.8868
| DCA 50 turn0.1550 | DCA 50 turn | 550 | 1025 | 0.1 | A | mA |  |  |  | 92 | 2.4 | 122.12 | 1.8855
| DCA 50 turn1550 | DCA 50 turn | 550 | 1025 | 1 | A | mA |  |  |  | 92 | 2.4 | 322.89 | 1.7672
| DCA 50 turn10550 | DCA 50 turn | 550 | 1025 | 10 | A | mA |  |  |  | 5638.1 | 0.50461 | 5638.2 | 0.50217
| DCA 50 turn100550 | DCA 50 turn | 550 | 1025 | 100 | A | mA |  |  |  | 57983 | 0.053297 | 57983 | 0.052219
| Resistance0.0000010 | Resistance | 0 | 0.011 | 1e-06 | kOhms | mΩ |  |  |  | 12 | 140 | 11.576 | 138.26
| Resistance0.000010 | Resistance | 0 | 0.011 | 1e-05 | kOhms | mΩ |  |  |  | 12.938 | 125.77 | 12.935 | 125.21
| Resistance0.00010 | Resistance | 0 | 0.011 | 0.0001 | kOhms | mΩ |  |  |  | 59.144 | 29.346 | 59.141 | 28.778
| Resistance0.0010 | Resistance | 0 | 0.011 | 0.001 | kOhms | mΩ |  |  |  | 580.12 | 3.175 | 580.12 | 2.9417
| Resistance0.0000010.011 | Resistance | 0.011 | 0.033 | 1e-06 | kOhms | mΩ |  |  |  | 18 | 140 | 17.34 | 138.45
| Resistance0.000010.011 | Resistance | 0.011 | 0.033 | 1e-05 | kOhms | mΩ |  |  |  | 18.266 | 133.34 | 18.261 | 133.19
| Resistance0.00010.011 | Resistance | 0.011 | 0.033 | 0.0001 | kOhms | mΩ |  |  |  | 60.489 | 46.163 | 60.483 | 45.899
| Resistance0.0010.011 | Resistance | 0.011 | 0.033 | 0.001 | kOhms | mΩ |  |  |  | 580.26 | 4.9787 | 580.25 | 4.8632
| Resistance0.0000010.033 | Resistance | 0.033 | 0.11 | 1e-06 | kOhms | mΩ |  |  |  | 18 | 100 | 17.339 | 103.86
| Resistance0.000010.033 | Resistance | 0.033 | 0.11 | 1e-05 | kOhms | mΩ |  |  |  | 18.221 | 101.31 | 18.218 | 101.09
| Resistance0.00010.033 | Resistance | 0.033 | 0.11 | 0.0001 | kOhms | mΩ |  |  |  | 60.259 | 41.183 | 60.259 | 40.719
| Resistance0.0010.033 | Resistance | 0.033 | 0.11 | 0.001 | kOhms | mΩ |  |  |  | 580.23 | 4.6418 | 580.23 | 4.4307
| Resistance0.0000010.11 | Resistance | 0.11 | 0.33 | 1e-06 | kOhms | mΩ |  |  |  | 25 | 300 | 24.812 | 296.87
| Resistance0.000010.11 | Resistance | 0.11 | 0.33 | 1e-05 | kOhms | mΩ |  |  |  | 25.358 | 294.79 | 25.35 | 294.41
| Resistance0.00010.11 | Resistance | 0.11 | 0.33 | 0.0001 | kOhms | mΩ |  |  |  | 61.466 | 184.35 | 61.46 | 183.05
| Resistance0.0010.11 | Resistance | 0.11 | 0.33 | 0.001 | kOhms | mΩ |  |  |  | 580.25 | 24.237 | 580.25 | 23.449
| Resistance0.000010.33 | Resistance | 0.33 | 1.1 | 1e-05 | kOhms | mΩ |  |  |  | 28 | 100 | 23.826 | 103.48
| Resistance0.00010.33 | Resistance | 0.33 | 1.1 | 0.0001 | kOhms | mΩ |  |  |  | 52.988 | 87.716 | 52.887 | 87.614
| Resistance0.0010.33 | Resistance | 0.33 | 1.1 | 0.001 | kOhms | mΩ |  |  |  | 577.21 | 17.351 | 577.2 | 17.171
| Resistance0.010.33 | Resistance | 0.33 | 1.1 | 0.01 | kOhms | mΩ |  |  |  | 5799.7 | 1.8197 | 5799.7 | 1.7449
| Resistance0.000011.1 | Resistance | 1.1 | 3.3 | 1e-05 | kOhms | Ω |  |  |  | 0.25 | 0.1 | 0.23105 | 0.10391
| Resistance0.00011.1 | Resistance | 1.1 | 3.3 | 0.0001 | kOhms | Ω |  |  |  | 0.25 | 0.10308 | 0.23681 | 0.10304
| Resistance0.0011.1 | Resistance | 1.1 | 3.3 | 0.001 | kOhms | Ω |  |  |  | 0.605 | 0.064275 | 0.60454 | 0.064066
| Resistance0.011.1 | Resistance | 1.1 | 3.3 | 0.01 | kOhms | Ω |  |  |  | 5.8014 | 0.008354 | 5.8013 | 0.0082072
| Resistance0.00013.3 | Resistance | 3.3 | 11 | 0.0001 | kOhms | Ω |  |  |  | 0.17 | 0.1 | 0.12477 | 0.10332
| Resistance0.0013.3 | Resistance | 3.3 | 11 | 0.001 | kOhms | Ω |  |  |  | 0.46565 | 0.083937 | 0.46492 | 0.08383
| Resistance0.013.3 | Resistance | 3.3 | 11 | 0.01 | kOhms | Ω |  |  |  | 5.7683 | 0.015342 | 5.7683 | 0.015182
| Resistance0.13.3 | Resistance | 3.3 | 11 | 0.1 | kOhms | Ω |  |  |  | 57.997 | 0.0016041 | 57.997 | 0.0015381
| Resistance0.000111 | Resistance | 11 | 33 | 0.0001 | kOhms | Ω |  |  |  | 1.3 | 0.1 | 1.1565 | 0.10389
| Resistance0.00111 | Resistance | 11 | 33 | 0.001 | kOhms | Ω |  |  |  | 1.3 | 0.1023 | 1.2456 | 0.10229
| Resistance0.0111 | Resistance | 11 | 33 | 0.01 | kOhms | Ω |  |  |  | 5.6627 | 0.052539 | 5.6618 | 0.052465
| Resistance0.111 | Resistance | 11 | 33 | 0.1 | kOhms | Ω |  |  |  | 57.978 | 0.0061973 | 57.978 | 0.0061535
| Resistance0.00133 | Resistance | 33 | 110 | 0.001 | kOhms | Ω |  |  |  | 1.3 | 0.13 | 1.2349 | 0.12649
| Resistance0.0133 | Resistance | 33 | 110 | 0.01 | kOhms | Ω |  |  |  | 4.3546 | 0.10789 | 4.3517 | 0.10781
| Resistance0.133 | Resistance | 33 | 110 | 0.1 | kOhms | Ω |  |  |  | 57.522 | 0.022123 | 57.523 | 0.022002
| Resistance133 | Resistance | 33 | 110 | 1 | kOhms | Ω |  |  |  | 579.95 | 0.0022919 | 579.95 | 0.0022413
| Resistance0.001110 | Resistance | 110 | 330 | 0.001 | kOhms | Ω |  |  |  | 12 | 0.14 | 11.564 | 0.13853
| Resistance0.01110 | Resistance | 110 | 330 | 0.01 | kOhms | Ω |  |  |  | 12.345 | 0.13708 | 12.34 | 0.13706
| Resistance0.1110 | Resistance | 110 | 330 | 0.1 | kOhms | Ω |  |  |  | 55.082 | 0.080188 | 55.079 | 0.080108
| Resistance1110 | Resistance | 110 | 330 | 1 | kOhms | Ω |  |  |  | 579.52 | 0.010063 | 579.52 | 0.010012
| Resistance0.01330 | Resistance | 330 | 1100 | 0.01 | kOhms | Ω |  |  |  | 16 | 0.17 | 12.172 | 0.17279
| Resistance0.1330 | Resistance | 330 | 1100 | 0.1 | kOhms | Ω |  |  |  | 38.622 | 0.15614 | 38.58 | 0.15612
| Resistance1330 | Resistance | 330 | 1100 | 1 | kOhms | Ω |  |  |  | 571.17 | 0.039215 | 571.17 | 0.039139
| Resistance10330 | Resistance | 330 | 1100 | 10 | kOhms | Ω |  |  |  | 5799.1 | 0.0040748 | 5799.1 | 0.0040418
| Resistance0.011100 | Resistance | 1100 | 3300 | 0.01 | kOhms | kΩ |  |  |  | 0.19 | 0.00017 | 0.17333 | 0.00017318
| Resistance0.11100 | Resistance | 1100 | 3300 | 0.1 | kOhms | kΩ |  |  |  | 0.19 | 0.00017216 | 0.17904 | 0.00017213
| Resistance11100 | Resistance | 1100 | 3300 | 1 | kOhms | kΩ |  |  |  | 0.55512 | 0.00011804 | 0.55496 | 0.00011789
| Resistance101100 | Resistance | 1100 | 3300 | 10 | kOhms | kΩ |  |  |  | 5.7933 | 1.6589e-05 | 5.7933 | 1.6468e-05
| Resistance0.13300 | Resistance | 3300 | 11000 | 0.1 | kOhms | kΩ |  |  |  | 0.33 | 0.00069 | 0.29035 | 0.00069271
| Resistance13300 | Resistance | 3300 | 11000 | 1 | kOhms | kΩ |  |  |  | 0.3729 | 0.00068714 | 0.37261 | 0.00068714
| Resistance103300 | Resistance | 3300 | 11000 | 10 | kOhms | kΩ |  |  |  | 4.8626 | 0.00044971 | 4.8624 | 0.00044965
| Resistance1003300 | Resistance | 3300 | 11000 | 100 | kOhms | kΩ |  |  |  | 57.852 | 6.2354e-05 | 57.852 | 6.2302e-05
| Resistance0.111000 | Resistance | 11000 | 33000 | 0.1 | kOhms | kΩ |  |  |  | 2.9 | 0.0012 | 2.887 | 0.0011547
| Resistance111000 | Resistance | 11000 | 33000 | 1 | kOhms | kΩ |  |  |  | 2.9012 | 0.0011544 | 2.901 | 0.0011544
| Resistance1011000 | Resistance | 11000 | 33000 | 10 | kOhms | kΩ |  |  |  | 4.2497 | 0.0011258 | 4.2488 | 0.0011258
| Resistance10011000 | Resistance | 11000 | 33000 | 100 | kOhms | kΩ |  |  |  | 54.576 | 0.00049848 | 54.576 | 0.00049842
| Resistance133000 | Resistance | 33000 | 110000 | 1 | kOhms | kΩ |  |  |  | 3.5 | 0.0058 | 3.4659 | 0.0057735
| Resistance1033000 | Resistance | 33000 | 110000 | 10 | kOhms | kΩ |  |  |  | 3.5801 | 0.0057727 | 3.5791 | 0.0057727
| Resistance10033000 | Resistance | 33000 | 110000 | 100 | kOhms | kΩ |  |  |  | 14.462 | 0.0056975 | 14.465 | 0.0056974
| Resistance100033000 | Resistance | 33000 | 110000 | 1000 | kOhms | kΩ |  |  |  | 503.99 | 0.0032606 | 503.99 | 0.0032605
| Resistance1110000 | Resistance | 110000 | 330000 | 1 | kOhms | kΩ |  |  |  | 120 | 0.0058 | 115.47 | 0.0057735
| Resistance10110000 | Resistance | 110000 | 330000 | 10 | kOhms | kΩ |  |  |  | 120 | 0.0057734 | 115.52 | 0.0057734
| Resistance100110000 | Resistance | 110000 | 330000 | 100 | kOhms | kΩ |  |  |  | 120 | 0.0057671 | 118.42 | 0.0057671
| Resistance1000110000 | Resistance | 110000 | 330000 | 1000 | kOhms | kΩ |  |  |  | 371.66 | 0.0052445 | 371.67 | 0.0052444
| Resistance10330000 | Resistance | 330000 | 1100000 | 10 | kOhms | kΩ |  |  |  | 930 | 0.017 | 577.19 | 0.017321
| Resistance100330000 | Resistance | 330000 | 1100000 | 100 | kOhms | kΩ |  |  |  | 930 | 0.01732 | 578 | 0.01732
| Resistance1000330000 | Resistance | 330000 | 1100000 | 1000 | kOhms | kΩ |  |  |  | 930 | 0.017297 | 611.79 | 0.017297
| Resistance10000330000 | Resistance | 330000 | 1100000 | 10000 | kOhms | kΩ |  |  |  | 3454.2 | 0.015468 | 3454.1 | 0.015468
| ACV0.0000010.0010.01 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 0.01 | 0.045 | kHz | 25 | 1700 | 23.108 | 1731.8
| ACV0.000010.0010.01 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 0.01 | 0.045 | kHz | 25 | 1717.9 | 23.784 | 1717.6
| ACV0.00010.0010.01 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 0.01 | 0.045 | kHz | 61.983 | 1124.3 | 61.97 | 1122.7
| ACV0.0010.0010.01 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 0.01 | 0.045 | kHz | 580.38 | 157.29 | 580.38 | 156.08
| ACV0.010.0010.01 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 0.01 | 0.045 | kHz | 5800 | 16.176 | 5800 | 15.689
| ACV0.0000010.0010.045 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 0.045 | 10 | kHz | 24 | 1200 | 23.109 | 1154.4
| ACV0.000010.0010.045 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 0.045 | 10 | kHz | 24 | 1142.3 | 23.798 | 1141.8
| ACV0.00010.0010.045 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 0.045 | 10 | kHz | 62.21 | 672.33 | 62.198 | 670.35
| ACV0.0010.0010.045 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 0.045 | 10 | kHz | 580.43 | 85.997 | 580.42 | 84.786
| ACV0.010.0010.045 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 0.045 | 10 | kHz | 5800.1 | 8.9929 | 5800.1 | 8.5055
| ACV0.0000010.00110 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 10 | 20 | kHz | 25 | 1700 | 23.108 | 1731.8
| ACV0.000010.00110 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 10 | 20 | kHz | 25 | 1717.9 | 23.784 | 1717.6
| ACV0.00010.00110 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 10 | 20 | kHz | 61.983 | 1124.3 | 61.97 | 1122.7
| ACV0.0010.00110 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 10 | 20 | kHz | 580.38 | 157.29 | 580.38 | 156.08
| ACV0.010.00110 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 10 | 20 | kHz | 5800 | 16.176 | 5800 | 15.689
| ACV0.0000010.00120 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 2300 | 23.108 | 2309.1
| ACV0.000010.00120 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 20 | 50 | kHz | 24 | 2294.4 | 23.77 | 2294.1
| ACV0.00010.00120 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 20 | 50 | kHz | 61.72 | 1616.4 | 61.708 | 1615
| ACV0.0010.00120 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 20 | 50 | kHz | 580.32 | 247.57 | 580.31 | 246.37
| ACV0.010.00120 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 20 | 50 | kHz | 5800 | 25.313 | 5800 | 24.826
| ACV0.0000010.00150 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 50 | 100 | kHz | 40 | 4000 | 38.113 | 4041.3
| ACV0.000010.00150 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 50 | 100 | kHz | 40 | 4031.5 | 38.517 | 4032
| ACV0.00010.00150 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 50 | 100 | kHz | 68.374 | 3416 | 68.283 | 3416.2
| ACV0.0010.00150 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 50 | 100 | kHz | 580.85 | 729.42 | 580.8 | 727.73
| ACV0.010.00150 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 50 | 100 | kHz | 5800.1 | 75.135 | 5800.1 | 74.409
| ACV0.0000010.001100 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 100 | 500 | kHz | 70 | 12000 | 69.286 | 11547
| ACV0.000010.001100 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 100 | 500 | kHz | 70 | 11541 | 69.497 | 11542
| ACV0.00010.001100 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 100 | 500 | kHz | 88.56 | 11078 | 88.414 | 11080
| ACV0.0010.001100 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 100 | 500 | kHz | 581.06 | 4649.5 | 580.96 | 4646.8
| ACV0.010.001100 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 100 | 500 | kHz | 5800.1 | 529.57 | 5800 | 527.92
| ACV0.0000010.0330.01 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 0.01 | 0.045 | kHz | 33 | 550 | 23.103 | 577.33
| ACV0.000010.0330.01 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 0.01 | 0.045 | kHz | 33 | 549.05 | 23.531 | 576.27
| ACV0.00010.0330.01 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 0.01 | 0.045 | kHz | 55.17 | 503.68 | 55.071 | 503.89
| ACV0.0010.0330.01 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 0.01 | 0.045 | kHz | 577.51 | 123.09 | 577.47 | 123.1
| ACV0.010.0330.01 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 0.01 | 0.045 | kHz | 5799.8 | 12.725 | 577.47 | 12.725
| ACV0.0000010.0330.045 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 0.045 | 10 | kHz | 24 | 350 | 23.104 | 346.39
| ACV0.000010.0330.045 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 0.045 | 10 | kHz | 24 | 345.06 | 23.624 | 345.18
| ACV0.00010.0330.045 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 0.045 | 10 | kHz | 58.533 | 274.73 | 58.43 | 274.91
| ACV0.0010.0330.045 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 0.045 | 10 | kHz | 579.4 | 50.596 | 579.36 | 50.6
| ACV0.010.0330.045 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 0.045 | 10 | kHz | 5800 | 5.1337 | 5799.9 | 5.1338
| ACV0.0000010.03310 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 10 | 20 | kHz | 24 | 810 | 23.102 | 808.27
| ACV0.000010.03310 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 10 | 20 | kHz | 24 | 807.25 | 23.467 | 807.34
| ACV0.00010.03310 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 10 | 20 | kHz | 52.172 | 737.66 | 52.078 | 737.88
| ACV0.0010.03310 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 10 | 20 | kHz | 574.82 | 223.06 | 574.77 | 223.07
| ACV0.010.03310 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 10 | 20 | kHz | 5799.5 | 23.648 | 5799.4 | 23.648
| ACV0.0000010.03320 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 20 | 50 | kHz | 47 | 1200 | 46.192 | 1154.7
| ACV0.000010.03320 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 20 | 50 | kHz | 47 | 1154.1 | 46.406 | 1154.2
| ACV0.00010.03320 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 20 | 50 | kHz | 65.852 | 1107.1 | 65.784 | 1107.2
| ACV0.0010.03320 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 20 | 50 | kHz | 571.22 | 452.24 | 571.18 | 452.11
| ACV0.010.03320 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 20 | 50 | kHz | 5799 | 50.926 | 5798.9 | 50.848
| ACV0.0000010.03350 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 50 | 100 | kHz | 200 | 2700 | 196.3 | 2655.8
| ACV0.000010.03350 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 50 | 100 | kHz | 200 | 2655.6 | 196.37 | 2655.7
| ACV0.00010.03350 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 50 | 100 | kHz | 202.7 | 2641.3 | 202.66 | 2641.3
| ACV0.0010.03350 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 50 | 100 | kHz | 582.13 | 1931.7 | 582.03 | 1931.7
| ACV0.010.03350 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 50 | 100 | kHz | 5796.8 | 307.99 | 5796.8 | 307.81
| ACV0.0000010.033100 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 100 | 500 | kHz | 390 | 5800 | 381.05 | 5773.5
| ACV0.000010.033100 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 100 | 500 | kHz | 390 | 5773.4 | 381.1 | 5773.4
| ACV0.00010.033100 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 100 | 500 | kHz | 390 | 5766 | 384.24 | 5766.1
| ACV0.0010.033100 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 100 | 500 | kHz | 642.89 | 5199.9 | 642.71 | 5200.1
| ACV0.010.033100 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 100 | 500 | kHz | 5783.1 | 1368.2 | 5783 | 1367.9
| ACV0.000010.330.01 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 0.01 | 0.045 | kHz | 0.11 | 0.56 | 0.10361 | 0.56176
| ACV0.00010.330.01 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 0.01 | 0.045 | kHz | 0.11 | 0.55987 | 0.1089 | 0.55992
| ACV0.0010.330.01 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 0.01 | 0.045 | kHz | 0.4828 | 0.46399 | 0.48243 | 0.46402
| ACV0.010.330.01 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 0.01 | 0.045 | kHz | 5.7719 | 0.10322 | 5.7718 | 0.10315
| ACV0.10.330.01 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 0.01 | 0.045 | kHz | 57.997 | 0.010625 | 57.997 | 0.010595
| ACV0.000010.330.045 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 0.045 | 10 | kHz | 0.07 | 0.35 | 0.069482 | 0.34636
| ACV0.00010.330.045 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 0.045 | 10 | kHz | 0.079475 | 0.34378 | 0.079164 | 0.34384
| ACV0.0010.330.045 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 0.045 | 10 | kHz | 0.52701 | 0.24771 | 0.52666 | 0.24769
| ACV0.010.330.045 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 0.045 | 10 | kHz | 5.7894 | 0.041307 | 5.7893 | 0.041234
| ACV0.10.330.045 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 0.045 | 10 | kHz | 57.999 | 0.0041989 | 57.999 | 0.0041685
| ACV0.000010.3310 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 10 | 20 | kHz | 0.07 | 0.81 | 0.069395 | 0.80826
| ACV0.00010.3310 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 10 | 20 | kHz | 0.074971 | 0.80677 | 0.074789 | 0.80681
| ACV0.0010.3310 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 10 | 20 | kHz | 0.43434 | 0.71615 | 0.43398 | 0.7162
| ACV0.010.3310 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 10 | 20 | kHz | 5.7428 | 0.20326 | 5.7427 | 0.20319
| ACV0.10.3310 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 10 | 20 | kHz | 57.994 | 0.021429 | 57.994 | 0.021398
| ACV0.000010.3320 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 20 | 50 | kHz | 0.07 | 1.2 | 0.069357 | 1.1547
| ACV0.00010.3320 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 20 | 50 | kHz | 0.073767 | 1.1535 | 0.073402 | 1.1536
| ACV0.0010.3320 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 20 | 50 | kHz | 0.38083 | 1.0735 | 0.38004 | 1.0736
| ACV0.010.3320 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 20 | 50 | kHz | 5.6889 | 0.39089 | 5.6885 | 0.39077
| ACV0.10.3320 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 20 | 50 | kHz | 57.988 | 0.043127 | 57.988 | 0.043055
| ACV0.000010.3350 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 50 | 100 | kHz | 0.24 | 2.7 | 0.23098 | 2.6558
| ACV0.00010.3350 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 50 | 100 | kHz | 0.24 | 2.6553 | 0.23261 | 2.6554
| ACV0.0010.3350 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 50 | 100 | kHz | 0.38861 | 2.6138 | 0.38742 | 2.6141
| ACV0.010.3350 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 50 | 100 | kHz | 5.3726 | 1.6156 | 5.3718 | 1.6155
| ACV0.10.3350 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 50 | 100 | kHz | 57.935 | 0.23007 | 57.935 | 0.2299
| ACV0.000010.33100 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 100 | 500 | kHz | 1.1 | 5.8 | 1.0392 | 5.7735
| ACV0.00010.33100 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 100 | 500 | kHz | 1.1 | 5.7733 | 1.04 | 5.7733
| ACV0.0010.33100 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 100 | 500 | kHz | 1.1025 | 5.757 | 1.1012 | 5.7573
| ACV0.010.33100 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 100 | 500 | kHz | 4.9058 | 4.8509 | 4.9039 | 4.851
| ACV0.10.33100 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 100 | 500 | kHz | 57.708 | 1.1137 | 57.707 | 1.1134
| ACV0.00013.30.01 | ACV | 3.3 | 33 | 0.0001 | V | mV | 0.01 | 0.045 | kHz | 0.93 | 0.58 | 0.92505 | 0.57732
| ACV0.0013.30.01 | ACV | 3.3 | 33 | 0.001 | V | mV | 0.01 | 0.045 | kHz | 0.98935 | 0.57563 | 0.98883 | 0.57564
| ACV0.013.30.01 | ACV | 3.3 | 33 | 0.01 | V | mV | 0.01 | 0.045 | kHz | 4.86 | 0.48311 | 4.8592 | 0.4831
| ACV0.13.30.01 | ACV | 3.3 | 33 | 0.1 | V | mV | 0.01 | 0.045 | kHz | 57.705 | 0.11031 | 57.705 | 0.11026
| ACV13.30.01 | ACV | 3.3 | 33 | 1 | V | mV | 0.01 | 0.045 | kHz | 579.97 | 0.011366 | 579.97 | 0.011347
| ACV0.00013.30.045 | ACV | 3.3 | 33 | 0.0001 | V | mV | 0.045 | 10 | kHz | 0.7 | 0.35 | 0.69482 | 0.34636
| ACV0.0013.30.045 | ACV | 3.3 | 33 | 0.001 | V | mV | 0.045 | 10 | kHz | 0.7925 | 0.34383 | 0.79164 | 0.34384
| ACV0.013.30.045 | ACV | 3.3 | 33 | 0.01 | V | mV | 0.045 | 10 | kHz | 5.2673 | 0.24773 | 5.2666 | 0.24769
| ACV0.13.30.045 | ACV | 3.3 | 33 | 0.1 | V | mV | 0.045 | 10 | kHz | 57.893 | 0.041279 | 57.893 | 0.041234
| ACV13.30.045 | ACV | 3.3 | 33 | 1 | V | mV | 0.045 | 10 | kHz | 579.99 | 0.004187 | 579.99 | 0.0041685
| ACV0.00013.310 | ACV | 3.3 | 33 | 0.0001 | V | mV | 10 | 20 | kHz | 0.7 | 0.81 | 0.69395 | 0.80826
| ACV0.0013.310 | ACV | 3.3 | 33 | 0.001 | V | mV | 10 | 20 | kHz | 0.74845 | 0.8068 | 0.74789 | 0.80681
| ACV0.013.310 | ACV | 3.3 | 33 | 0.01 | V | mV | 10 | 20 | kHz | 4.3407 | 0.7162 | 4.3398 | 0.7162
| ACV0.13.310 | ACV | 3.3 | 33 | 0.1 | V | mV | 10 | 20 | kHz | 57.427 | 0.20323 | 57.427 | 0.20319
| ACV13.310 | ACV | 3.3 | 33 | 1 | V | mV | 10 | 20 | kHz | 579.94 | 0.021417 | 579.94 | 0.021398
| ACV0.00013.320 | ACV | 3.3 | 33 | 0.0001 | V | mV | 20 | 50 | kHz | 0.7 | 1.2 | 0.69357 | 1.1547
| ACV0.0013.320 | ACV | 3.3 | 33 | 0.001 | V | mV | 20 | 50 | kHz | 0.73558 | 1.1536 | 0.73402 | 1.1536
| ACV0.013.320 | ACV | 3.3 | 33 | 0.01 | V | mV | 20 | 50 | kHz | 3.8054 | 1.0735 | 3.8004 | 1.0736
| ACV0.13.320 | ACV | 3.3 | 33 | 0.1 | V | mV | 20 | 50 | kHz | 56.887 | 0.39084 | 56.885 | 0.39077
| ACV13.320 | ACV | 3.3 | 33 | 1 | V | mV | 20 | 50 | kHz | 579.88 | 0.043096 | 579.88 | 0.043055
| ACV0.00013.350 | ACV | 3.3 | 33 | 0.0001 | V | mV | 50 | 100 | kHz | 2.3 | 2.7 | 2.3098 | 2.6558
| ACV0.0013.350 | ACV | 3.3 | 33 | 0.001 | V | mV | 50 | 100 | kHz | 2.3287 | 2.6553 | 2.3261 | 2.6554
| ACV0.013.350 | ACV | 3.3 | 33 | 0.01 | V | mV | 50 | 100 | kHz | 3.8845 | 2.6138 | 3.8742 | 2.6141
| ACV0.13.350 | ACV | 3.3 | 33 | 0.1 | V | mV | 50 | 100 | kHz | 53.724 | 1.6155 | 53.718 | 1.6155
| ACV13.350 | ACV | 3.3 | 33 | 1 | V | mV | 50 | 100 | kHz | 579.35 | 0.22997 | 579.35 | 0.2299
| ACV0.00013.3100 | ACV | 3.3 | 33 | 0.0001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3098 | 2.6558
| ACV0.0013.3100 | ACV | 3.3 | 33 | 0.001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3261 | 2.6554
| ACV0.013.3100 | ACV | 3.3 | 33 | 0.01 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 3.8742 | 2.6141
| ACV0.13.3100 | ACV | 3.3 | 33 | 0.1 | V | mV | 100 | 500 | kHz | 47.305 | 3.5605 | 53.718 | 1.6155
| ACV13.3100 | ACV | 3.3 | 33 | 1 | V | mV | 100 | 500 | kHz | 579.35 | 0.22997 | 579.35 | 0.2299
| ACV0.0013310 | ACV | 33 | 330 | 0.001 | V | mV | 10 | 45 | kHz | 3.5 | 0.58 | 3.4804 | 0.57731
| ACV0.013310 | ACV | 33 | 330 | 0.01 | V | mV | 10 | 45 | kHz | 4.3066 | 0.57509 | 4.279 | 0.57515
| ACV0.13310 | ACV | 33 | 330 | 0.1 | V | mV | 10 | 45 | kHz | 46.676 | 0.47222 | 46.637 | 0.47224
| ACV13310 | ACV | 33 | 330 | 1 | V | mV | 10 | 45 | kHz | 576.99 | 0.10496 | 576.98 | 0.10486
| ACV103310 | ACV | 33 | 330 | 10 | V | mV | 10 | 45 | kHz | 5799.7 | 0.010813 | 5799.7 | 0.010773
| ACV0.001330.045 | ACV | 33 | 330 | 0.001 | V | mV | 0.045 | 10 | kHz | 12 | 0.92 | 10.401 | 0.92374
| ACV0.01330.045 | ACV | 33 | 330 | 0.01 | V | mV | 0.045 | 10 | kHz | 12 | 0.92246 | 10.846 | 0.92255
| ACV0.1330.045 | ACV | 33 | 330 | 0.1 | V | mV | 0.045 | 10 | kHz | 43.33 | 0.84014 | 43.23 | 0.84029
| ACV1330.045 | ACV | 33 | 330 | 1 | V | mV | 0.045 | 10 | kHz | 572.74 | 0.26513 | 572.7 | 0.26496
| ACV10330.045 | ACV | 33 | 330 | 10 | V | mV | 0.045 | 10 | kHz | 5799.2 | 0.028425 | 5799.2 | 0.028338
| ACV0.0013310 | ACV | 33 | 330 | 0.001 | V | mV | 10 | 20 | kHz | 24 | 1 | 10.403 | 1.0392
| ACV0.013310 | ACV | 33 | 330 | 0.01 | V | mV | 10 | 20 | kHz | 24 | 1.0324 | 10.814 | 1.0381
| ACV0.13310 | ACV | 33 | 330 | 0.1 | V | mV | 10 | 20 | kHz | 41.668 | 0.95893 | 41.572 | 0.95908
| ACV13310 | ACV | 33 | 330 | 1 | V | mV | 10 | 20 | kHz | 570.93 | 0.32823 | 570.89 | 0.32806
| ACV103310 | ACV | 33 | 330 | 10 | V | mV | 10 | 20 | kHz | 5799 | 0.035713 | 5799 | 0.035625
| ACV0.0013320 | ACV | 33 | 330 | 0.001 | V | mV | 20 | 50 | kHz | 22 | 1.4 | 10.406 | 1.3856
| ACV0.013320 | ACV | 33 | 330 | 0.01 | V | mV | 20 | 50 | kHz | 22 | 1.4 | 10.74 | 1.3847
| ACV0.13320 | ACV | 33 | 330 | 0.1 | V | mV | 20 | 50 | kHz | 35.99 | 1.3576 | 37.32 | 1.3149
| ACV13320 | ACV | 33 | 330 | 1 | V | mV | 20 | 50 | kHz | 564.71 | 0.54675 | 564.67 | 0.54661
| ACV103320 | ACV | 33 | 330 | 10 | V | mV | 20 | 50 | kHz | 5798.2 | 0.062551 | 5798.2 | 0.062463
| ACV0.0013350 | ACV | 33 | 330 | 0.001 | V | mV | 50 | 100 | kHz | 93 | 2.8 | 92.377 | 2.7713
| ACV0.013350 | ACV | 33 | 330 | 0.01 | V | mV | 50 | 100 | kHz | 93 | 2.771 | 92.486 | 2.771
| ACV0.13350 | ACV | 33 | 330 | 0.1 | V | mV | 50 | 100 | kHz | 102.21 | 2.7467 | 102.13 | 2.7468
| ACV13350 | ACV | 33 | 330 | 1 | V | mV | 50 | 100 | kHz | 547.03 | 1.864 | 546.93 | 1.8639
| ACV103350 | ACV | 33 | 330 | 10 | V | mV | 50 | 100 | kHz | 5793.6 | 0.28252 | 5793.6 | 0.28229
| ACV0.0013300.045 | ACV | 330 | 1020 | 0.001 | V | mV | 0.045 | 1 | kHz | 24 | 0.58 | 23.302 | 0.5772
| ACV0.013300.045 | ACV | 330 | 1020 | 0.01 | V | mV | 0.045 | 1 | kHz | 33.348 | 0.5702 | 33.313 | 0.57005
| ACV0.13300.045 | ACV | 330 | 1020 | 0.1 | V | mV | 0.045 | 1 | kHz | 510.39 | 0.32681 | 510.48 | 0.32619
| ACV13300.045 | ACV | 330 | 1020 | 1 | V | mV | 0.045 | 1 | kHz | 5790.4 | 0.041348 | 5790.4 | 0.040965
| ACV103300.045 | ACV | 330 | 1020 | 10 | V | mV | 0.045 | 1 | kHz | 57999 | 0.0042634 | 57999 | 0.0041091
| ACV0.0013301 | ACV | 330 | 1020 | 0.001 | V | mV | 1 | 5 | kHz | 27 | 0.92 | 23.231 | 0.92366
| ACV0.013301 | ACV | 330 | 1020 | 0.01 | V | mV | 1 | 5 | kHz | 29.878 | 0.91895 | 29.852 | 0.91886
| ACV0.13301 | ACV | 330 | 1020 | 0.1 | V | mV | 1 | 5 | kHz | 446.35 | 0.66691 | 446.38 | 0.66648
| ACV13301 | ACV | 330 | 1020 | 1 | V | mV | 1 | 5 | kHz | 5775.5 | 0.10258 | 5775.6 | 0.1022
| ACV103301 | ACV | 330 | 1020 | 10 | V | mV | 1 | 5 | kHz | 57998 | 0.010452 | 57998 | 0.010298
| ACV0.0013305 | ACV | 330 | 1020 | 0.001 | V | mV | 5 | 10 | kHz | 64 | 1 | 23.258 | 1.0391
| ACV0.013305 | ACV | 330 | 1020 | 0.01 | V | mV | 5 | 10 | kHz | 64 | 1.0349 | 29.181 | 1.0348
| ACV0.13305 | ACV | 330 | 1020 | 0.1 | V | mV | 5 | 10 | kHz | 426.27 | 0.78701 | 426.28 | 0.78663
| ACV13305 | ACV | 330 | 1020 | 1 | V | mV | 5 | 10 | kHz | 5769 | 0.12897 | 5769.1 | 0.12859
| ACV103305 | ACV | 330 | 1020 | 10 | V | mV | 5 | 10 | kHz | 57997 | 0.013136 | 57997 | 0.012982
| ACA0.000000010.0000290.01 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 0.01 | 0.02 | kHz | 0.25 | 4600 | 0.11567 | 2308.9
| ACA0.00000010.0000290.01 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 0.01 | 0.02 | kHz | 0.25 | 4600 | 0.12526 | 2285.6
| ACA0.0000010.0000290.01 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 0.01 | 0.02 | kHz | 0.49657 | 3852.8 | 0.56529 | 1474.7
| ACA0.000010.0000290.01 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 0.01 | 0.02 | kHz | 5.7969 | 210.33 | 5.7968 | 209.79
| ACA0.00010.0000290.01 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 0.01 | 0.02 | kHz | 58 | 21.321 | 58 | 21.102
| ACA0.000000010.0000290.02 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 0.26 | 1700 | 0.1157 | 1731.5
| ACA0.00000010.0000290.02 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 0.26 | 1700 | 0.12615 | 1707.2
| ACA0.0000010.0000290.02 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 0.57496 | 983.71 | 0.57472 | 983.11
| ACA0.000010.0000290.02 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.7988 | 127.4 | 5.7987 | 126.86
| ACA0.00010.0000290.02 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 58 | 12.951 | 58 | 12.732
| ACA0.000000010.0000290.045 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.23 | 1100 | 0.1157 | 1442.9
| ACA0.00000010.0000290.045 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.23 | 1144.6 | 0.12666 | 1418.2
| ACA0.0000010.0000290.045 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.57929 | 757.04 | 0.57906 | 756.32
| ACA0.000010.0000290.045 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.7995 | 93.495 | 5.7995 | 92.953
| ACA0.00010.0000290.045 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 0.045 | 1 | kHz | 58 | 9.5395 | 58 | 9.3209
| ACA0.000000010.0000291 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 1 | 5 | kHz | 0.18 | 3500 | 0.17334 | 3463.8
| ACA0.00000010.0000291 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 1 | 5 | kHz | 0.18 | 3447.8 | 0.17982 | 3448
| ACA0.0000010.0000291 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 1 | 5 | kHz | 0.56479 | 2648.4 | 0.56455 | 2648.3
| ACA0.000010.0000291 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 1 | 5 | kHz | 5.793 | 469.14 | 5.7929 | 468.61
| ACA0.00010.0000291 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 1 | 5 | kHz | 57.999 | 47.695 | 57.999 | 47.477
| ACA0.000000010.0000295 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 5 | 10 | kHz | 0.25 | 9200 | 0.23102 | 9237.4
| ACA0.00000010.0000295 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 5 | 10 | kHz | 0.25 | 9227.6 | 0.23463 | 9228
| ACA0.0000010.0000295 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 5 | 10 | kHz | 0.51843 | 8521.5 | 0.5179 | 8522.3
| ACA0.000010.0000295 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 5 | 10 | kHz | 5.7406 | 2796.5 | 5.7404 | 2795.6
| ACA0.00010.0000295 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 5 | 10 | kHz | 57.994 | 301.1 | 57.994 | 300.63
| ACA0.000000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46195 | 18475
| ACA0.00000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 10 | 30 | kHz | 0.62 | 18470 | 0.46386 | 18470
| ACA0.0000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 10 | 30 | kHz | 0.63124 | 18040 | 0.63083 | 18041
| ACA0.000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 10 | 30 | kHz | 5.6089 | 9536 | 5.6087 | 9535.6
| ACA0.00010.00002910 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 10 | 30 | kHz | 57.974 | 1200.1 | 57.974 | 1199.6
| ACA0.000000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 0.01 | 0.02 | kHz | 2.5 | 4600 | 0.17324 | 2309.4
| ACA0.00000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 0.01 | 0.02 | kHz | 2.5 | 4600 | 0.17519 | 2308.9
| ACA0.0000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 0.01 | 0.02 | kHz | 2.5 | 4600 | 0.35449 | 2261
| ACA0.000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 0.01 | 0.02 | kHz | 4.5637 | 3974.7 | 5.4482 | 1293.1
| ACA0.00010.000330.01 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 0.01 | 0.02 | kHz | 57.951 | 173.04 | 57.951 | 173.01
| ACA0.000000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 2.6 | 1700 | 0.17325 | 1443.4
| ACA0.00000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 2.6 | 1700 | 0.17611 | 1442.6
| ACA0.0000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 2.6 | 1700 | 0.41532 | 1380.3
| ACA0.000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.5729 | 799.14 | 5.6385 | 599.33
| ACA0.00010.000330.02 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 57.981 | 69.408 | 57.981 | 69.377
| ACA0.000000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 0.045 | 1 | kHz | 2.3 | 530 | 0.17327 | 1154.7
| ACA0.00000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 0.045 | 1 | kHz | 2.3 | 528.18 | 0.17661 | 1153.8
| ACA0.0000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 0.045 | 1 | kHz | 2.3 | 542.23 | 0.44424 | 1085.3
| ACA0.000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.6923 | 407.42 | 5.692 | 407.38
| ACA0.00010.000330.045 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.988 | 45.149 | 57.988 | 45.119
| ACA0.000000010.000331 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 1 | 5 | kHz | 2 | 1800 | 0.23097 | 2309.4
| ACA0.00000010.000331 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 1 | 5 | kHz | 2 | 1830.5 | 0.23282 | 2308.9
| ACA0.0000010.000331 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 1 | 5 | kHz | 2 | 1837.7 | 0.40315 | 2263.7
| ACA0.000010.000331 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 1 | 5 | kHz | 5.4541 | 1305.5 | 5.4537 | 1305.5
| ACA0.00010.000331 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 1 | 5 | kHz | 57.951 | 175.31 | 57.951 | 175.28
| ACA0.000000010.000335 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 5 | 10 | kHz | 0.66 | 5700 | 0.34643 | 5773.5
| ACA0.00000010.000335 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 5 | 10 | kHz | 0.66 | 5679.5 | 0.34724 | 5773.3
| ACA0.0000010.000335 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 5 | 10 | kHz | 0.66 | 5682.4 | 0.42712 | 5751.7
| ACA0.000010.000335 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 5 | 10 | kHz | 4.664 | 4722.4 | 4.6633 | 4722.5
| ACA0.00010.000335 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 5 | 10 | kHz | 57.698 | 1048.7 | 57.698 | 1048.6
| ACA0.000000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 10 | 30 | kHz | 2.8 | 11000 | 0.69284 | 11547
| ACA0.00000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 10 | 30 | kHz | 2.8 | 10936 | 0.6932 | 11547
| ACA0.0000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 10 | 30 | kHz | 2.8 | 10937 | 0.73366 | 11536
| ACA0.000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 10 | 30 | kHz | 3.8012 | 10736 | 3.8003 | 10736
| ACA0.00010.0003310 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 10 | 30 | kHz | 56.886 | 3907.8 | 56.885 | 3907.7
| ACA0.000000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 0.01 | 0.02 | kHz | 25 | 4600 | 2.3096 | 2078.5
| ACA0.00000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 0.01 | 0.02 | kHz | 25 | 4600 | 2.3304 | 2077.9
| ACA0.0000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 0.01 | 0.02 | kHz | 25 | 4600 | 4.1521 | 2029.8
| ACA0.000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 0.01 | 0.02 | kHz | 45.602 | 3975.7 | 55.067 | 1107.1
| ACA0.00010.00330.01 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 0.01 | 0.02 | kHz | 579.6 | 142.98 | 579.6 | 142.92
| ACA0.000000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 26 | 1700 | 2.3105 | 1039.2
| ACA0.00000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 26 | 1700 | 2.3417 | 1038.4
| ACA0.0000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 26 | 1700 | 4.9485 | 973.1
| ACA0.000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 55.639 | 801.88 | 57.139 | 346.85
| ACA0.00010.00330.02 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 579.9 | 37.954 | 579.9 | 37.896
| ACA0.000000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 0.045 | 1 | kHz | 23 | 350 | 2.3103 | 461.86
| ACA0.00000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 0.045 | 1 | kHz | 23 | 350 | 2.3574 | 460.72
| ACA0.0000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 0.045 | 1 | kHz | 23 | 350 | 5.6713 | 388.3
| ACA0.000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 0.045 | 1 | kHz | 57.853 | 83.334 | 57.852 | 83.196
| ACA0.00010.00330.045 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 0.045 | 1 | kHz | 579.99 | 8.5712 | 579.99 | 8.513
| ACA0.000000010.00331 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 1 | 5 | kHz | 23 | 350 | 2.3102 | 923.74
| ACA0.00000010.00331 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 1 | 5 | kHz | 23 | 350 | 2.344 | 922.87
| ACA0.0000010.00331 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 1 | 5 | kHz | 23 | 350 | 5.0734 | 855.43
| ACA0.000010.00331 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 1 | 5 | kHz | 57.317 | 282.34 | 57.316 | 282.22
| ACA0.00010.00331 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 1 | 5 | kHz | 579.93 | 30.415 | 579.93 | 30.357
| ACA0.000000010.00335 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 5 | 10 | kHz | 21 | 1800 | 3.4644 | 2309.4
| ACA0.00000010.00335 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 5 | 10 | kHz | 21 | 1830.5 | 3.4808 | 2309
| ACA0.0000010.00335 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 5 | 10 | kHz | 21 | 1837.6 | 5.0249 | 2268.5
| ACA0.000010.00335 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 5 | 10 | kHz | 54.668 | 1330 | 54.661 | 1330
| ACA0.00010.00335 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 5 | 10 | kHz | 579.52 | 179.93 | 579.51 | 179.83
| ACA0.000000010.003310 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 10 | 30 | kHz | 48 | 4600 | 4.619 | 4618.8
| ACA0.00000010.003310 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 10 | 30 | kHz | 48 | 4600 | 4.6281 | 4618.6
| ACA0.0000010.003310 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 10 | 30 | kHz | 48 | 4600 | 5.5287 | 4594.5
| ACA0.000010.003310 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 10 | 30 | kHz | 48 | 4600 | 49.519 | 3572.4
| ACA0.00010.003310 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 10 | 30 | kHz | 578.06 | 691.95 | 578.06 | 691.72
| ACA0.0000010.0330.01 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 0.01 | 0.02 | kHz | 0.25 | 4.6 | 0.023096 | 2.0785
| ACA0.000010.0330.01 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 0.01 | 0.02 | kHz | 0.25 | 4.6 | 0.023304 | 2.0779
| ACA0.00010.0330.01 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 0.01 | 0.02 | kHz | 0.25 | 4.6 | 0.041521 | 2.0298
| ACA0.0010.0330.01 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 0.01 | 0.02 | kHz | 0.45601 | 3.9757 | 0.55067 | 1.1071
| ACA0.010.0330.01 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 0.01 | 0.02 | kHz | 5.796 | 0.14295 | 5.796 | 0.14292
| ACA0.0000010.0330.02 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.023105 | 1.0392
| ACA0.000010.0330.02 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.023417 | 1.0384
| ACA0.00010.0330.02 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.049485 | 0.9731
| ACA0.0010.0330.02 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 0.02 | 0.045 | kHz | 0.55049 | 0.98035 | 0.57139 | 0.34685
| ACA0.010.0330.02 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 0.02 | 0.045 | kHz | 5.799 | 0.037922 | 5.799 | 0.037896
| ACA0.0000010.0330.045 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 0.045 | 1 | kHz | 0.23 | 1.2 | 0.023103 | 0.46186
| ACA0.000010.0330.045 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.23 | 1.2 | 0.023574 | 0.46072
| ACA0.00010.0330.045 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.23 | 1.2 | 0.056713 | 0.3883
| ACA0.0010.0330.045 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.5763 | 0.15061 | 0.57852 | 0.083196
| ACA0.010.0330.045 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7999 | 0.0085387 | 5.7999 | 0.008513
| ACA0.0000010.0331 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 1 | 5 | kHz | 0.23 | 1.2 | 0.057739 | 1.1547
| ACA0.000010.0331 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 1 | 5 | kHz | 0.23 | 1.2 | 0.057939 | 1.1542
| ACA0.00010.0331 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 1 | 5 | kHz | 0.23 | 1.2 | 0.075293 | 1.1131
| ACA0.0010.0331 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 1 | 5 | kHz | 0.5724 | 0.46948 | 0.57238 | 0.46941
| ACA0.010.0331 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 1 | 5 | kHz | 5.7991 | 0.053184 | 5.799 | 0.053139
| ACA0.0000010.0335 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 5 | 10 | kHz | 0.23 | 3.5 | 0.11547 | 2.3094
| ACA0.000010.0335 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 5 | 10 | kHz | 0.23 | 3.5 | 0.11557 | 2.3092
| ACA0.00010.0335 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 5 | 10 | kHz | 0.23 | 3.5 | 0.12482 | 2.2869
| ACA0.0010.0335 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 5 | 10 | kHz | 0.5249 | 2.6064 | 0.56185 | 1.4851
| ACA0.010.0335 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 5 | 10 | kHz | 5.7962 | 0.21171 | 5.7962 | 0.21161
| ACA0.0000010.03310 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 10 | 30 | kHz | 0.42 | 12 | 0.23094 | 4.6188
| ACA0.000010.03310 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 10 | 30 | kHz | 0.42 | 12 | 0.23099 | 4.6187
| ACA0.00010.03310 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 10 | 30 | kHz | 0.42 | 12 | 0.2357 | 4.6073
| ACA0.0010.03310 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 10 | 30 | kHz | 0.42 | 12 | 0.56711 | 3.883
| ACA0.010.03310 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 10 | 30 | kHz | 5.7853 | 0.8322 | 5.7852 | 0.83196
| ACA0.000010.330.01 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.01 | 0.045 | kHz | 0.22 | 2 | 0.11555 | 2.0784
| ACA0.00010.330.01 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.01 | 0.045 | kHz | 0.22 | 1.9851 | 0.11823 | 2.0766
| ACA0.0010.330.01 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 0.35436 | 1.9242 | 0.35432 | 1.9241
| ACA0.010.330.01 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 5.674 | 0.54897 | 5.674 | 0.54873
| ACA0.10.330.01 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 57.987 | 0.057472 | 57.987 | 0.057366
| ACA0.000010.330.045 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.23 | 1.2 | 0.11561 | 0.57727
| ACA0.00010.330.045 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.23 | 1.2 | 0.12237 | 0.57314
| ACA0.0010.330.045 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.27277 | 1.1611 | 0.53032 | 0.38022
| ACA0.010.330.045 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7908 | 0.052593 | 5.7908 | 0.052332
| ACA0.10.330.045 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.999 | 0.0053641 | 57.999 | 0.0052584
| ACA0.000010.331 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 1 | 5 | kHz | 1.2 | 6.9 | 1.1547 | 6.9282
| ACA0.00010.331 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 1 | 5 | kHz | 1.2 | 6.9278 | 1.1554 | 6.9278
| ACA0.0010.331 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 1 | 5 | kHz | 1.2167 | 6.89 | 1.2159 | 6.89
| ACA0.010.331 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 1 | 5 | kHz | 5.1273 | 4.9042 | 5.126 | 4.903
| ACA0.10.331 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 1 | 5 | kHz | 57.863 | 0.72635 | 57.863 | 0.7249
| ACA0.000010.335 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7733 | 28.868
| ACA0.00010.335 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 28.867 | 5.7742 | 28.867
| ACA0.0010.335 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 28.859 | 5.7874 | 28.859
| ACA0.010.335 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 5 | 10 | kHz | 7.1025 | 28.066 | 7.1008 | 28.066
| ACA0.10.335 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 5 | 10 | kHz | 56.086 | 11.818 | 56.085 | 11.816
| ACA0.000011.10.01 | ACA | 1.1 | 3 | 1e-05 | A | mA | 0.01 | 0.045 | kHz | 0.12 | 2.1 | 0.11544 | 2.0785
| ACA0.00011.10.01 | ACA | 1.1 | 3 | 0.0001 | A | mA | 0.01 | 0.045 | kHz | 0.12 | 2.0782 | 0.11653 | 2.0782
| ACA0.0011.10.01 | ACA | 1.1 | 3 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 0.21221 | 2.0559 | 0.20929 | 2.056
| ACA0.011.10.01 | ACA | 1.1 | 3 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 4.9354 | 1.2243 | 4.9327 | 1.2227
| ACA0.11.10.01 | ACA | 1.1 | 3 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 57.878 | 0.15761 | 57.878 | 0.15629
| ACA0.000011.10.045 | ACA | 1.1 | 3 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11552 | 0.69281
| ACA0.00011.10.045 | ACA | 1.1 | 3 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69213 | 0.11808 | 0.69221
| ACA0.0011.10.045 | ACA | 1.1 | 3 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.35354 | 0.64097 | 0.34715 | 0.64072
| ACA0.011.10.045 | ACA | 1.1 | 3 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.6736 | 0.17933 | 5.672 | 0.17635
| ACA0.11.10.045 | ACA | 1.1 | 3 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.987 | 0.019669 | 57.987 | 0.018337
| ACA0.000011.11 | ACA | 1.1 | 3 | 1e-05 | A | mA | 1 | 5 | kHz | 1.3 | 6.9 | 1.1547 | 6.9282
| ACA0.00011.11 | ACA | 1.1 | 3 | 0.0001 | A | mA | 1 | 5 | kHz | 1.3 | 6.9281 | 1.1551 | 6.9281
| ACA0.0011.11 | ACA | 1.1 | 3 | 0.001 | A | mA | 1 | 5 | kHz | 1.3 | 6.9221 | 1.1805 | 6.9222
| ACA0.011.11 | ACA | 1.1 | 3 | 0.01 | A | mA | 1 | 5 | kHz | 3.4738 | 6.4074 | 3.4712 | 6.4073
| ACA0.11.11 | ACA | 1.1 | 3 | 0.1 | A | mA | 1 | 5 | kHz | 56.721 | 1.7647 | 56.72 | 1.7635
| ACA0.000011.15 | ACA | 1.1 | 3 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.773 | 28.868
| ACA0.00011.15 | ACA | 1.1 | 3 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 28.868 | 5.7751 | 28.867
| ACA0.0011.15 | ACA | 1.1 | 3 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 28.866 | 5.7799 | 28.866
| ACA0.011.15 | ACA | 1.1 | 3 | 0.01 | A | mA | 5 | 10 | kHz | 6.3734 | 28.729 | 6.3715 | 28.729
| ACA0.11.15 | ACA | 1.1 | 3 | 0.1 | A | mA | 5 | 10 | kHz | 45.931 | 21.05 | 45.928 | 21.049
| ACA0.000130.045 | ACA | 3 | 11 | 0.0001 | A | mA | 0.045 | 0.1 | kHz | 1.4 | 2.9 | 1.3514 | 2.9169
| ACA0.00130.045 | ACA | 3 | 11 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 1.4 | 2.9059 | 1.3976 | 2.9058
| ACA0.0130.045 | ACA | 3 | 11 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 4.8276 | 2.2255 | 4.8256 | 2.2249
| ACA0.130.045 | ACA | 3 | 11 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 57.774 | 0.35784 | 57.773 | 0.35697
| ACA130.045 | ACA | 3 | 11 | 1 | A | mA | 0.045 | 0.1 | kHz | 579.98 | 0.036361 | 579.98 | 0.036004
| ACA0.000130.1 | ACA | 3 | 11 | 0.0001 | A | mA | 0.1 | 1 | kHz | 0.72 | 4.9 | 0.71179 | 4.8617
| ACA0.00130.1 | ACA | 3 | 11 | 0.001 | A | mA | 0.1 | 1 | kHz | 0.74791 | 4.8524 | 0.74718 | 4.8524
| ACA0.0130.1 | ACA | 3 | 11 | 0.01 | A | mA | 0.1 | 1 | kHz | 3.6609 | 4.1627 | 3.6587 | 4.1625
| ACA0.130.1 | ACA | 3 | 11 | 0.1 | A | mA | 0.1 | 1 | kHz | 57.347 | 0.8558 | 57.346 | 0.85495
| ACA130.1 | ACA | 3 | 11 | 1 | A | mA | 0.1 | 1 | kHz | 579.93 | 0.087455 | 579.93 | 0.087098
| ACA0.000131 | ACA | 3 | 11 | 0.0001 | A | mA | 1 | 5 | kHz | 2.4 | 35 | 2.3096 | 34.641
| ACA0.00131 | ACA | 3 | 11 | 0.001 | A | mA | 1 | 5 | kHz | 2.4 | 34.641 | 2.3111 | 34.641
| ACA0.0131 | ACA | 3 | 11 | 0.01 | A | mA | 1 | 5 | kHz | 2.5151 | 34.627 | 2.5097 | 34.627
| ACA0.131 | ACA | 3 | 11 | 0.1 | A | mA | 1 | 5 | kHz | 21.042 | 33.336 | 21.028 | 33.336
| ACA131 | ACA | 3 | 11 | 1 | A | mA | 1 | 5 | kHz | 550.06 | 13.201 | 550.05 | 13.2
| ACA0.0001110.045 | ACA | 11 | 20.5 | 0.0001 | A | mA | 0.045 | 0.1 | kHz | 5.8 | 1.4 | 5.7744 | 1.3856
| ACA0.001110.045 | ACA | 11 | 20.5 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 5.8 | 1.3858 | 5.7855 | 1.3853
| ACA0.01110.045 | ACA | 11 | 20.5 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 6.9251 | 1.3563 | 6.9029 | 1.3544
| ACA0.1110.045 | ACA | 11 | 20.5 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 55.154 | 0.59968 | 55.169 | 0.59282
| ACA1110.045 | ACA | 11 | 20.5 | 1 | A | mA | 0.045 | 0.1 | kHz | 579.65 | 0.069283 | 579.66 | 0.065853
| ACA0.0001110.1 | ACA | 11 | 20.5 | 0.0001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7746 | 1.732
| ACA0.001110.1 | ACA | 11 | 20.5 | 0.001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7321 | 5.7832 | 1.7318
| ACA0.01110.1 | ACA | 11 | 20.5 | 0.01 | A | mA | 0.1 | 1 | kHz | 6.768 | 1.7058 | 6.7463 | 1.7044
| ACA0.1110.1 | ACA | 11 | 20.5 | 0.1 | A | mA | 0.1 | 1 | kHz | 53.701 | 0.85898 | 53.71 | 0.8527
| ACA1110.1 | ACA | 11 | 20.5 | 1 | A | mA | 0.1 | 1 | kHz | 579.44 | 0.10196 | 579.45 | 0.098538
| ACA0.0001111 | ACA | 11 | 20.5 | 0.0001 | A | mA | 1 | 5 | kHz | 6.5 | 1.7 | 5.7746 | 1.732
| ACA0.001111 | ACA | 11 | 20.5 | 0.001 | A | mA | 1 | 5 | kHz | 6.5 | 1.7321 | 5.7832 | 1.7318
| ACA0.01111 | ACA | 11 | 20.5 | 0.01 | A | mA | 1 | 5 | kHz | 6.768 | 1.7058 | 6.7463 | 1.7044
| ACA0.1111 | ACA | 11 | 20.5 | 0.1 | A | mA | 1 | 5 | kHz | 53.701 | 0.85898 | 53.71 | 0.8527
| ACA1111 | ACA | 11 | 20.5 | 1 | A | mA | 1 | 5 | kHz | 579.44 | 0.10196 | 579.45 | 0.098538
| ACA 10 turn0.000110.01 | ACA 10 turn | 1 | 11 | 0.0001 | A | mA | 0.01 | 0.045 | kHz | 3.6 | 2.3 | 0.51064 | 1.9786
| ACA 10 turn0.00110.01 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 4.0309 | 2.2608 | 4.6123 | 1.6726
| ACA 10 turn0.0110.01 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 57.645 | 0.40888 | 57.643 | 0.40772
| ACA 10 turn0.110.01 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 579.96 | 0.042733 | 579.96 | 0.042225
| ACA 10 turn110.01 | ACA 10 turn | 1 | 11 | 1 | A | mA | 0.01 | 0.045 | kHz | 5800 | 0.0044275 | 5800 | 0.004224
| ACA 10 turn1010.01 | ACA 10 turn | 1 | 11 | 10 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.00050379 | 58000 | 0.0004224
| ACA 10 turn0.00110.045 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.45 | 2 | 0.44338 | 1.967
| ACA 10 turn0.0110.045 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 0.045 | 1 | kHz | 4.5988 | 1.6572 | 4.5973 | 1.6568
| ACA 10 turn0.110.045 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.646 | 0.40189 | 57.646 | 0.4011
| ACA 10 turn110.045 | ACA 10 turn | 1 | 11 | 1 | A | mA | 0.045 | 1 | kHz | 579.96 | 0.041857 | 579.96 | 0.041514
| ACA 10 turn1010.045 | ACA 10 turn | 1 | 11 | 10 | A | mA | 0.045 | 1 | kHz | 5800 | 0.0042902 | 5800 | 0.0041529
| ACA 10 turn0.00111 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 1 | 5 | kHz | 11 | 1.7 | 10.739 | 1.6769
| ACA 10 turn0.0111 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 1 | 5 | kHz | 12.102 | 1.6065 | 12.087 | 1.6058
| ACA 10 turn0.111 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 1 | 5 | kHz | 58.762 | 0.56421 | 58.75 | 0.56151
| ACA 10 turn111 | ACA 10 turn | 1 | 11 | 1 | A | mA | 1 | 5 | kHz | 580.08 | 0.061363 | 580.07 | 0.060093
| ACA 10 turn1011 | ACA 10 turn | 1 | 11 | 10 | A | mA | 1 | 5 | kHz | 5800 | 0.0065224 | 5800 | 0.0060138
| ACA 10 turn0.00115 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 3.2 | 57.508 | 3.1498
| ACA 10 turn0.0115 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 3.1412 | 57.791 | 3.1405
| ACA 10 turn0.115 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 5 | 10 | kHz | 81.456 | 2.4995 | 81.427 | 2.4964
| ACA 10 turn115 | ACA 10 turn | 1 | 11 | 1 | A | mA | 5 | 10 | kHz | 582.77 | 0.41409 | 582.75 | 0.41124
| ACA 10 turn1015 | ACA 10 turn | 1 | 11 | 10 | A | mA | 5 | 10 | kHz | 5800.3 | 0.04265 | 5800.3 | 0.041491
| ACA 10 turn0.001110.01 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 7.6 | 3.2 | 0.17121 | 2.0095
| ACA 10 turn0.01110.01 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 7.6 | 3.2 | 1.1723 | 1.9853
| ACA 10 turn0.1110.01 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 38.311 | 2.173 | 49.593 | 1.1395
| ACA 10 turn1110.01 | ACA 10 turn | 11 | 30 | 1 | A | mA | 0.01 | 0.045 | kHz | 578.88 | 0.14434 | 578.86 | 0.14288
| ACA 10 turn10110.01 | ACA 10 turn | 11 | 30 | 10 | A | mA | 0.01 | 0.045 | kHz | 5799.9 | 0.014921 | 5799.9 | 0.014332
| ACA 10 turn0.001110.045 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.099 | 2 | 0.092875 | 1.9999
| ACA 10 turn0.01110.045 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 0.045 | 1 | kHz | 1.1295 | 1.9754 | 1.1019 | 1.9755
| ACA 10 turn0.1110.045 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 0.045 | 1 | kHz | 49.662 | 1.1308 | 49.643 | 1.129
| ACA 10 turn1110.045 | ACA 10 turn | 11 | 30 | 1 | A | mA | 0.045 | 1 | kHz | 578.87 | 0.14259 | 578.87 | 0.14126
| ACA 10 turn10110.045 | ACA 10 turn | 11 | 30 | 10 | A | mA | 0.045 | 1 | kHz | 5799.9 | 0.014704 | 5799.9 | 0.014168
| ACA 10 turn0.001111 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 1 | 5 | kHz | 6.7 | 2.1 | 6.6205 | 2.0513
| ACA 10 turn0.01111 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 1 | 5 | kHz | 7.4052 | 2.0342 | 7.3715 | 2.0344
| ACA 10 turn0.1111 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 1 | 5 | kHz | 50.743 | 1.2946 | 50.702 | 1.2931
| ACA 10 turn1111 | ACA 10 turn | 11 | 30 | 1 | A | mA | 1 | 5 | kHz | 578.86 | 0.17288 | 578.85 | 0.17144
| ACA 10 turn10111 | ACA 10 turn | 11 | 30 | 10 | A | mA | 1 | 5 | kHz | 5799.9 | 0.017795 | 5799.9 | 0.017213
| ACA 10 turn0.001115 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 5 | 10 | kHz | 56 | 3.4 | 55.012 | 3.3767
| ACA 10 turn0.01115 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 5 | 10 | kHz | 56 | 3.3729 | 55.236 | 3.3728
| ACA 10 turn0.1115 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 5 | 10 | kHz | 75.465 | 3.0449 | 75.403 | 3.0441
| ACA 10 turn1115 | ACA 10 turn | 11 | 30 | 1 | A | mA | 5 | 10 | kHz | 579.54 | 0.70852 | 579.51 | 0.70625
| ACA 10 turn10115 | ACA 10 turn | 11 | 30 | 10 | A | mA | 5 | 10 | kHz | 5800 | 0.073274 | 5799.9 | 0.07231
| ACA 10 turn0.01300.045 | ACA 10 turn | 30 | 110 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 91 | 3.3 | 6.5002 | 1.961
| ACA 10 turn0.1300.045 | ACA 10 turn | 30 | 110 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 91 | 3.3 | 33.761 | 1.7802
| ACA 10 turn1300.045 | ACA 10 turn | 30 | 110 | 1 | A | mA | 0.045 | 0.1 | kHz | 569.58 | 0.46952 | 569.6 | 0.46805
| ACA 10 turn10300.045 | ACA 10 turn | 30 | 110 | 10 | A | mA | 0.045 | 0.1 | kHz | 5798.9 | 0.049226 | 5798.9 | 0.048591
| ACA 10 turn100300.045 | ACA 10 turn | 30 | 110 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0051153 | 58000 | 0.004861
| ACA 10 turn0.01300.1 | ACA 10 turn | 30 | 110 | 0.01 | A | mA | 0.1 | 1 | kHz | 7 | 2 | 6.9624 | 1.9637
| ACA 10 turn0.1300.1 | ACA 10 turn | 30 | 110 | 0.1 | A | mA | 0.1 | 1 | kHz | 34.08 | 1.7849 | 34.055 | 1.7842
| ACA 10 turn1300.1 | ACA 10 turn | 30 | 110 | 1 | A | mA | 0.1 | 1 | kHz | 569.57 | 0.47215 | 569.58 | 0.47069
| ACA 10 turn10300.1 | ACA 10 turn | 30 | 110 | 10 | A | mA | 0.1 | 1 | kHz | 5798.9 | 0.049513 | 5798.9 | 0.048877
| ACA 10 turn100300.1 | ACA 10 turn | 30 | 110 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.005144 | 58000 | 0.0048897
| ACA 10 turn0.01301 | ACA 10 turn | 30 | 110 | 0.01 | A | mA | 1 | 5 | kHz | 21 | 4 | 20.754 | 3.9948
| ACA 10 turn0.1301 | ACA 10 turn | 30 | 110 | 0.1 | A | mA | 1 | 5 | kHz | 35.187 | 3.8981 | 35.052 | 3.8976
| ACA 10 turn1301 | ACA 10 turn | 30 | 110 | 1 | A | mA | 1 | 5 | kHz | 542.96 | 1.7989 | 542.92 | 1.7949
| ACA 10 turn10301 | ACA 10 turn | 30 | 110 | 10 | A | mA | 1 | 5 | kHz | 5795.5 | 0.20876 | 5795.5 | 0.20654
| ACA 10 turn100301 | ACA 10 turn | 30 | 110 | 100 | A | mA | 1 | 5 | kHz | 58000 | 0.02158 | 58000 | 0.020689
| ACA 10 turn0.011100.045 | ACA 10 turn | 110 | 205 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 91 | 3.3 | 15.183 | 1.9698
| ACA 10 turn0.11100.045 | ACA 10 turn | 110 | 205 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 91 | 3.3 | 25.866 | 1.937
| ACA 10 turn11100.045 | ACA 10 turn | 110 | 205 | 1 | A | mA | 0.045 | 0.1 | kHz | 460.09 | 1.4996 | 519.36 | 0.9567
| ACA 10 turn101100.045 | ACA 10 turn | 110 | 205 | 10 | A | mA | 0.045 | 0.1 | kHz | 5792.4 | 0.113 | 5792.5 | 0.11033
| ACA 10 turn1001100.045 | ACA 10 turn | 110 | 205 | 100 | A | mA | 0.045 | 0.1 | kHz | 57999 | 0.012123 | 57999 | 0.011052
| ACA 10 turn0.011100.1 | ACA 10 turn | 110 | 205 | 0.01 | A | mA | 0.1 | 1 | kHz | 17 | 2 | 16.085 | 1.9728
| ACA 10 turn0.11100.1 | ACA 10 turn | 110 | 205 | 0.1 | A | mA | 0.1 | 1 | kHz | 26.884 | 1.9412 | 26.694 | 1.9403
| ACA 10 turn11100.1 | ACA 10 turn | 110 | 205 | 1 | A | mA | 0.1 | 1 | kHz | 519.23 | 0.96608 | 519.32 | 0.96125
| ACA 10 turn101100.1 | ACA 10 turn | 110 | 205 | 10 | A | mA | 0.1 | 1 | kHz | 5792.4 | 0.11364 | 5792.5 | 0.11097
| ACA 10 turn1001100.1 | ACA 10 turn | 110 | 205 | 100 | A | mA | 0.1 | 1 | kHz | 57999 | 0.012187 | 57999 | 0.011116
| ACA 10 turn0.011101 | ACA 10 turn | 110 | 205 | 0.01 | A | mA | 1 | 5 | kHz | 52 | 4 | 51.339 | 3.9959
| ACA 10 turn0.11101 | ACA 10 turn | 110 | 205 | 0.1 | A | mA | 1 | 5 | kHz | 56.51 | 3.9809 | 56.41 | 3.9805
| ACA 10 turn11101 | ACA 10 turn | 110 | 205 | 1 | A | mA | 1 | 5 | kHz | 428.53 | 3.0151 | 428.45 | 3.0125
| ACA 10 turn101101 | ACA 10 turn | 110 | 205 | 10 | A | mA | 1 | 5 | kHz | 5769.4 | 0.46816 | 5769.5 | 0.46553
| ACA 10 turn1001101 | ACA 10 turn | 110 | 205 | 100 | A | mA | 1 | 5 | kHz | 57997 | 0.047964 | 57997 | 0.046893
| ACA 20 turn0.001220.01 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 91 | 3.3 | -5.1487 | 2.1491
| ACA 20 turn0.01220.01 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 91 | 3.3 | -4.666 | 2.1421
| ACA 20 turn0.1220.01 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 91 | 3.3 | 33.745 | 1.7049
| ACA 20 turn1220.01 | ACA 20 turn | 22 | 60 | 1 | A | mA | 0.01 | 0.045 | kHz | 574.86 | 0.30182 | 574.86 | 0.30094
| ACA 20 turn10220.01 | ACA 20 turn | 22 | 60 | 10 | A | mA | 0.01 | 0.045 | kHz | 5799.5 | 0.030827 | 5799.5 | 0.030463
| ACA 20 turn0.001220.045 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 0.045 | 1 | kHz | -5.1 | 2.1 | -5.3489 | 2.1091
| ACA 20 turn0.01220.045 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 0.045 | 1 | kHz | -4.8311 | 2.1018 | -4.8537 | 2.1019
| ACA 20 turn0.1220.045 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 0.045 | 1 | kHz | 34.168 | 1.66 | 34.129 | 1.6596
| ACA 20 turn1220.045 | ACA 20 turn | 22 | 60 | 1 | A | mA | 0.045 | 1 | kHz | 575.05 | 0.28981 | 575.04 | 0.28893
| ACA 20 turn10220.045 | ACA 20 turn | 22 | 60 | 10 | A | mA | 0.045 | 1 | kHz | 5799.5 | 0.029596 | 5799.5 | 0.029232
| ACA 20 turn0.001221 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 1 | 5 | kHz | 11 | 2.5 | 9.8177 | 2.5158
| ACA 20 turn0.01221 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 1 | 5 | kHz | 11 | 2.5118 | 10.129 | 2.5118
| ACA 20 turn0.1221 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 1 | 5 | kHz | 38.512 | 2.1996 | 38.481 | 2.1993
| ACA 20 turn1221 | ACA 20 turn | 22 | 60 | 1 | A | mA | 1 | 5 | kHz | 573.1 | 0.47828 | 573.09 | 0.47742
| ACA 20 turn10221 | ACA 20 turn | 22 | 60 | 10 | A | mA | 1 | 5 | kHz | 5799.3 | 0.049137 | 5799.3 | 0.048774
| ACA 20 turn0.001225 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 5 | 10 | kHz | 96 | 6.4 | 95.759 | 6.3904
| ACA 20 turn0.01225 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 5 | 10 | kHz | 96 | 6.3894 | 95.847 | 6.3894
| ACA 20 turn0.1225 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 5 | 10 | kHz | 104.35 | 6.2966 | 104.27 | 6.2968
| ACA 20 turn1225 | ACA 20 turn | 22 | 60 | 1 | A | mA | 5 | 10 | kHz | 552.98 | 3.3048 | 552.9 | 3.3032
| ACA 20 turn10225 | ACA 20 turn | 22 | 60 | 10 | A | mA | 5 | 10 | kHz | 5796.2 | 0.39228 | 5796.1 | 0.3912
| ACA 20 turn0.01600.045 | ACA 20 turn | 60 | 220 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 91 | 3.3 | 13.907 | 1.9671
| ACA 20 turn0.1600.045 | ACA 20 turn | 60 | 220 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 91 | 3.3 | 29.106 | 1.9149
| ACA 20 turn1600.045 | ACA 20 turn | 60 | 220 | 1 | A | mA | 0.045 | 0.1 | kHz | 511.53 | 1.3885 | 543.32 | 0.85782
| ACA 20 turn10600.045 | ACA 20 turn | 60 | 220 | 10 | A | mA | 0.045 | 0.1 | kHz | 5795.6 | 0.098278 | 5795.6 | 0.097962
| ACA 20 turn100600.045 | ACA 20 turn | 60 | 220 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0099389 | 58000 | 0.0098118
| ACA 20 turn0.01600.1 | ACA 20 turn | 60 | 220 | 0.01 | A | mA | 0.1 | 1 | kHz | 16 | 2 | 15.707 | 1.9768
| ACA 20 turn0.1600.1 | ACA 20 turn | 60 | 220 | 0.1 | A | mA | 0.1 | 1 | kHz | 30.71 | 1.9256 | 30.652 | 1.9256
| ACA 20 turn1600.1 | ACA 20 turn | 60 | 220 | 1 | A | mA | 0.1 | 1 | kHz | 543.17 | 0.8701 | 543.15 | 0.86953
| ACA 20 turn10600.1 | ACA 20 turn | 60 | 220 | 10 | A | mA | 0.1 | 1 | kHz | 5795.6 | 0.099832 | 5795.6 | 0.099516
| ACA 20 turn100600.1 | ACA 20 turn | 60 | 220 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.010095 | 58000 | 0.0099677
| ACA 20 turn0.01601 | ACA 20 turn | 60 | 220 | 0.01 | A | mA | 1 | 5 | kHz | 47 | 7.2 | 44.639 | 7.2102
| ACA 20 turn0.1601 | ACA 20 turn | 60 | 220 | 0.1 | A | mA | 1 | 5 | kHz | 49.097 | 7.1948 | 49.045 | 7.1948
| ACA 20 turn1601 | ACA 20 turn | 60 | 220 | 1 | A | mA | 1 | 5 | kHz | 383.77 | 6.1244 | 383.66 | 6.124
| ACA 20 turn10601 | ACA 20 turn | 60 | 220 | 10 | A | mA | 1 | 5 | kHz | 5742.6 | 1.2843 | 5742.6 | 1.2833
| ACA 20 turn100601 | ACA 20 turn | 60 | 220 | 100 | A | mA | 1 | 5 | kHz | 57994 | 0.13141 | 57994 | 0.13101
| ACA 20 turn0.012200.045 | ACA 20 turn | 220 | 410 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 38 | 2 | 37.202 | 1.9859
| ACA 20 turn0.12200.045 | ACA 20 turn | 220 | 410 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 42.616 | 1.978 | 42.478 | 1.9778
| ACA 20 turn12200.045 | ACA 20 turn | 220 | 410 | 1 | A | mA | 0.045 | 0.1 | kHz | 423.75 | 1.4809 | 423.6 | 1.4795
| ACA 20 turn102200.045 | ACA 20 turn | 220 | 410 | 10 | A | mA | 0.045 | 0.1 | kHz | 5769.7 | 0.22683 | 5769.8 | 0.22535
| ACA 20 turn1002200.045 | ACA 20 turn | 220 | 410 | 100 | A | mA | 0.045 | 0.1 | kHz | 57997 | 0.023295 | 57997 | 0.022692
| ACA 20 turn0.012200.1 | ACA 20 turn | 220 | 410 | 0.01 | A | mA | 0.1 | 1 | kHz | 41 | 2 | 40.525 | 1.9975
| ACA 20 turn0.12200.1 | ACA 20 turn | 220 | 410 | 0.1 | A | mA | 0.1 | 1 | kHz | 45.878 | 1.9897 | 45.743 | 1.9895
| ACA 20 turn12200.1 | ACA 20 turn | 220 | 410 | 1 | A | mA | 0.1 | 1 | kHz | 424.05 | 1.4965 | 423.93 | 1.495
| ACA 20 turn102200.1 | ACA 20 turn | 220 | 410 | 10 | A | mA | 0.1 | 1 | kHz | 5769.4 | 0.2305 | 5769.5 | 0.22901
| ACA 20 turn1002200.1 | ACA 20 turn | 220 | 410 | 100 | A | mA | 0.1 | 1 | kHz | 57997 | 0.023667 | 57997 | 0.023064
| ACA 20 turn0.012201 | ACA 20 turn | 220 | 410 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 7.2 | 111.41 | 7.2104
| ACA 20 turn0.12201 | ACA 20 turn | 220 | 410 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 7.2081 | 112.9 | 7.2081
| ACA 20 turn12201 | ACA 20 turn | 220 | 410 | 1 | A | mA | 1 | 5 | kHz | 256.97 | 6.9898 | 256.36 | 6.9894
| ACA 20 turn102201 | ACA 20 turn | 220 | 410 | 10 | A | mA | 1 | 5 | kHz | 5443.7 | 2.7293 | 5443.6 | 2.726
| ACA 20 turn1002201 | ACA 20 turn | 220 | 410 | 100 | A | mA | 1 | 5 | kHz | 57960 | 0.29758 | 57960 | 0.29594
| ACA 50 turn0.001550.01 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 35 | 3.2 | 2.8048 | 2.2532
| ACA 50 turn0.01550.01 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 35 | 3.2 | 2.9878 | 2.2523
| ACA 50 turn0.1550.01 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 35 | 3.2 | 19.929 | 2.1717
| ACA 50 turn1550.01 | ACA 50 turn | 55 | 150 | 1 | A | mA | 0.01 | 0.045 | kHz | 547.95 | 0.83256 | 547.93 | 0.83182
| ACA 50 turn10550.01 | ACA 50 turn | 55 | 150 | 10 | A | mA | 0.01 | 0.045 | kHz | 5796.4 | 0.091096 | 5796.4 | 0.090722
| ACA 50 turn100550.01 | ACA 50 turn | 55 | 150 | 100 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.0092312 | 57995 | 0.009081
| ACA 50 turn0.001550.045 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 5.6 | 2 | 1.1856 | 2.0288
| ACA 50 turn0.01550.045 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 5.6 | 2.0278 | 1.3902 | 2.0278
| ACA 50 turn0.1550.045 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 20.258 | 1.9385 | 20.191 | 1.9385
| ACA 50 turn1550.045 | ACA 50 turn | 55 | 150 | 1 | A | mA | 0.045 | 0.1 | kHz | 553.43 | 0.68162 | 553.42 | 0.68084
| ACA 50 turn10550.045 | ACA 50 turn | 55 | 150 | 10 | A | mA | 0.045 | 0.1 | kHz | 5797.1 | 0.073473 | 5797.1 | 0.073098
| ACA 50 turn100550.045 | ACA 50 turn | 55 | 150 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0074657 | 57996 | 0.0073155
| ACA 50 turn0.001550.1 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 0.1 | 1 | kHz | 5.6 | 2 | 1.1856 | 2.0288
| ACA 50 turn0.01550.1 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 0.1 | 1 | kHz | 5.6 | 2.0278 | 1.3902 | 2.0278
| ACA 50 turn0.1550.1 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 0.1 | 1 | kHz | 20.258 | 1.9385 | 20.191 | 1.9385
| ACA 50 turn1550.1 | ACA 50 turn | 55 | 150 | 1 | A | mA | 0.1 | 1 | kHz | 553.43 | 0.68162 | 553.42 | 0.68084
| ACA 50 turn10550.1 | ACA 50 turn | 55 | 150 | 10 | A | mA | 0.1 | 1 | kHz | 5797.1 | 0.073473 | 5797.1 | 0.073098
| ACA 50 turn100550.1 | ACA 50 turn | 55 | 150 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0074657 | 57996 | 0.0073155
| ACA 50 turn0.001551 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 1 | 5 | kHz | 53 | 4 | 52.066 | 3.9905
| ACA 50 turn0.01551 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 1 | 5 | kHz | 53 | 3.9901 | 52.14 | 3.9902
| ACA 50 turn0.1551 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 1 | 5 | kHz | 60.324 | 3.9531 | 60.241 | 3.9532
| ACA 50 turn1551 | ACA 50 turn | 55 | 150 | 1 | A | mA | 1 | 5 | kHz | 506.66 | 2.4345 | 506.56 | 2.4338
| ACA 50 turn10551 | ACA 50 turn | 55 | 150 | 10 | A | mA | 1 | 5 | kHz | 5789 | 0.31673 | 5789 | 0.31608
| ACA 50 turn100551 | ACA 50 turn | 55 | 150 | 100 | A | mA | 1 | 5 | kHz | 57999 | 0.031987 | 57983 | 0.031723
| ACA 50 turn0.001555 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 5 | 10 | kHz | 290 | 15 | 286.96 | 14.567
| ACA 50 turn0.01555 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 5 | 10 | kHz | 290 | 14.568 | 286.97 | 14.567
| ACA 50 turn0.1555 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 5 | 10 | kHz | 290 | 14.559 | 289 | 14.558
| ACA 50 turn1555 | ACA 50 turn | 55 | 150 | 1 | A | mA | 5 | 10 | kHz | 476.92 | 13.752 | 476.94 | 13.748
| ACA 50 turn10555 | ACA 50 turn | 55 | 150 | 10 | A | mA | 5 | 10 | kHz | 5667.3 | 4.2563 | 5667.5 | 4.2489
| ACA 50 turn100555 | ACA 50 turn | 55 | 150 | 100 | A | mA | 5 | 10 | kHz | 57986 | 0.45012 | 57777 | 0.44684
| ACA 50 turn0.011500.045 | ACA 50 turn | 150 | 550 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 45 | 2 | 44.316 | 1.9954
| ACA 50 turn0.11500.045 | ACA 50 turn | 150 | 550 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 50.562 | 1.987 | 50.394 | 1.987
| ACA 50 turn11500.045 | ACA 50 turn | 150 | 550 | 1 | A | mA | 0.045 | 0.1 | kHz | 446.98 | 1.5167 | 446.68 | 1.5163
| ACA 50 turn101500.045 | ACA 50 turn | 150 | 550 | 10 | A | mA | 0.045 | 0.1 | kHz | 5772.3 | 0.25355 | 5772.2 | 0.25287
| ACA 50 turn1001500.045 | ACA 50 turn | 150 | 550 | 100 | A | mA | 0.045 | 0.1 | kHz | 57997 | 0.025829 | 57997 | 0.025549
| ACA 50 turn0.011500.1 | ACA 50 turn | 150 | 550 | 0.01 | A | mA | 0.1 | 1 | kHz | 55 | 2.1 | 54.127 | 2.051
| ACA 50 turn0.11500.1 | ACA 50 turn | 150 | 550 | 0.1 | A | mA | 0.1 | 1 | kHz | 60.071 | 2.043 | 59.88 | 2.0431
| ACA 50 turn11500.1 | ACA 50 turn | 150 | 550 | 1 | A | mA | 0.1 | 1 | kHz | 446.4 | 1.5834 | 446.09 | 1.5831
| ACA 50 turn101500.1 | ACA 50 turn | 150 | 550 | 10 | A | mA | 0.1 | 1 | kHz | 5770.8 | 0.27063 | 5770.8 | 0.26995
| ACA 50 turn1001500.1 | ACA 50 turn | 150 | 550 | 100 | A | mA | 0.1 | 1 | kHz | 57997 | 0.027576 | 57997 | 0.027296
| ACA 50 turn0.011501 | ACA 50 turn | 150 | 550 | 0.01 | A | mA | 1 | 5 | kHz | 360 | 17 | 114.69 | 17.436
| ACA 50 turn0.11501 | ACA 50 turn | 150 | 550 | 0.1 | A | mA | 1 | 5 | kHz | 360 | 17.434 | 115.77 | 17.434
| ACA 50 turn11501 | ACA 50 turn | 150 | 550 | 1 | A | mA | 1 | 5 | kHz | 360 | 17.326 | 191.97 | 17.327
| ACA 50 turn101501 | ACA 50 turn | 150 | 550 | 10 | A | mA | 1 | 5 | kHz | 4575.2 | 12.238 | 4574.7 | 12.238
| ACA 50 turn1001501 | ACA 50 turn | 150 | 550 | 100 | A | mA | 1 | 5 | kHz | 57786 | 1.8555 | 57786 | 1.8551
| ACA 50 turn0.015500.045 | ACA 50 turn | 550 | 1025 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 140 | 2.1 | 138.65 | 2.0895
| ACA 50 turn0.15500.045 | ACA 50 turn | 550 | 1025 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 140.64 | 2.0887 | 140.53 | 2.0884
| ACA 50 turn15500.045 | ACA 50 turn | 550 | 1025 | 1 | A | mA | 0.045 | 0.1 | kHz | 323.54 | 1.9815 | 323.39 | 1.9801
| ACA 50 turn105500.045 | ACA 50 turn | 550 | 1025 | 10 | A | mA | 0.045 | 0.1 | kHz | 5603.6 | 0.61564 | 5604.4 | 0.61248
| ACA 50 turn1005500.045 | ACA 50 turn | 550 | 1025 | 100 | A | mA | 0.045 | 0.1 | kHz | 57979 | 0.065636 | 57979 | 0.064244
| ACA 50 turn0.015500.1 | ACA 50 turn | 550 | 1025 | 0.01 | A | mA | 0.1 | 1 | kHz | 160 | 2.2 | 154.57 | 2.1551
| ACA 50 turn0.15500.1 | ACA 50 turn | 550 | 1025 | 0.1 | A | mA | 0.1 | 1 | kHz | 160 | 2.1543 | 156.42 | 2.154
| ACA 50 turn15500.1 | ACA 50 turn | 550 | 1025 | 1 | A | mA | 0.1 | 1 | kHz | 332.82 | 2.0511 | 332.62 | 2.0498
| ACA 50 turn105500.1 | ACA 50 turn | 550 | 1025 | 10 | A | mA | 0.1 | 1 | kHz | 5592.6 | 0.65649 | 5593.4 | 0.65336
| ACA 50 turn1005500.1 | ACA 50 turn | 550 | 1025 | 100 | A | mA | 0.1 | 1 | kHz | 57977 | 0.070159 | 57978 | 0.068767
| ACA 50 turn0.015501 | ACA 50 turn | 550 | 1025 | 0.01 | A | mA | 1 | 5 | kHz | 740 | 17 | 286.6 | 17.436
| ACA 50 turn0.15501 | ACA 50 turn | 550 | 1025 | 0.1 | A | mA | 1 | 5 | kHz | 740 | 17.435 | 287.51 | 17.435
| ACA 50 turn15501 | ACA 50 turn | 550 | 1025 | 1 | A | mA | 1 | 5 | kHz | 740 | 17.419 | 313.08 | 17.419
| ACA 50 turn105501 | ACA 50 turn | 550 | 1025 | 10 | A | mA | 1 | 5 | kHz | 2645.4 | 16.018 | 2643.7 | 16.018
| ACA 50 turn1005501 | ACA 50 turn | 550 | 1025 | 100 | A | mA | 1 | 5 | kHz | 56588 | 4.0886 | 56587 | 4.0866
| Capacitance0.00000000010.00000022 | Capacitance | 2.2e-07 | 3.3e-06 | 1e-10 | mF | nF |  |  |  | 0.019 | 3800 | 0.011547 | 5773.5
| Capacitance0.0000000010.00000022 | Capacitance | 2.2e-07 | 3.3e-06 | 1e-09 | mF | nF |  |  |  | 0.019 | 3843.2 | 0.011561 | 5771
| Capacitance0.000000010.00000022 | Capacitance | 2.2e-07 | 3.3e-06 | 1e-08 | mF | nF |  |  |  | 0.019 | 4020.2 | 0.012849 | 5544.1
| Capacitance0.00000010.00000022 | Capacitance | 2.2e-07 | 3.3e-06 | 1e-07 | mF | nF |  |  |  | 0.058965 | 2008.1 | 0.058958 | 2005.7
| Capacitance0.0000010.00000022 | Capacitance | 2.2e-07 | 3.3e-06 | 1e-06 | mF | nF |  |  |  | 0.5801 | 217.1 | 0.5801 | 215.92
| Capacitance0.0000000010.0000033 | Capacitance | 3.3e-06 | 1.1e-05 | 1e-09 | mF | nF |  |  |  | 0.012 | 2900 | 0.011566 | 2885.7
| Capacitance0.000000010.0000033 | Capacitance | 3.3e-06 | 1.1e-05 | 1e-08 | mF | nF |  |  |  | 0.012549 | 2835.8 | 0.012511 | 2834.7
| Capacitance0.00000010.0000033 | Capacitance | 3.3e-06 | 1.1e-05 | 1e-07 | mF | nF |  |  |  | 0.057173 | 1393.9 | 0.05714 | 1385.8
| Capacitance0.0000010.0000033 | Capacitance | 3.3e-06 | 1.1e-05 | 1e-06 | mF | nF |  |  |  | 0.57987 | 164.53 | 0.57986 | 159.93
| Capacitance0.0000000010.000011 | Capacitance | 1.1e-05 | 3.3e-05 | 1e-09 | mF | nF |  |  |  | 0.12 | 2900 | 0.11547 | 2886.7
| Capacitance0.000000010.000011 | Capacitance | 1.1e-05 | 3.3e-05 | 1e-08 | mF | nF |  |  |  | 0.12 | 2885.5 | 0.11561 | 2885.2
| Capacitance0.00000010.000011 | Capacitance | 1.1e-05 | 3.3e-05 | 1e-07 | mF | nF |  |  |  | 0.12812 | 2743.9 | 0.12808 | 2742.3
| Capacitance0.0000010.000011 | Capacitance | 1.1e-05 | 3.3e-05 | 1e-06 | mF | nF |  |  |  | 0.58907 | 853.21 | 0.58904 | 850.14
| Capacitance0.000010.000011 | Capacitance | 1.1e-05 | 3.3e-05 | 1e-05 | mF | nF |  |  |  | 5.8009 | 853.27 | 5.8009 | 89.037
| Capacitance0.000000010.000033 | Capacitance | 3.3e-05 | 0.00011 | 1e-08 | mF | nF |  |  |  | 0.19 | 3500 | 0.11566 | 2885.7
| Capacitance0.00000010.000033 | Capacitance | 3.3e-05 | 0.00011 | 1e-07 | mF | nF |  |  |  | 0.19 | 3500 | 0.12511 | 2834.7
| Capacitance0.0000010.000033 | Capacitance | 3.3e-05 | 0.00011 | 1e-06 | mF | nF |  |  |  | 0.57163 | 1393 | 0.5714 | 1385.8
| Capacitance0.000010.000033 | Capacitance | 3.3e-05 | 0.00011 | 1e-05 | mF | nF |  |  |  | 5.7986 | 163.92 | 5.7986 | 159.93
| Capacitance0.000000010.00011 | Capacitance | 0.00011 | 0.00033 | 1e-08 | mF | nF |  |  |  | 0.33 | 3100 | 0.34649 | 2886.6
| Capacitance0.00000010.00011 | Capacitance | 0.00011 | 0.00033 | 1e-07 | mF | nF |  |  |  | 0.33 | 3083.6 | 0.34959 | 2881.1
| Capacitance0.0000010.00011 | Capacitance | 0.00011 | 0.00033 | 1e-06 | mF | nF |  |  |  | 0.61135 | 2460.1 | 0.61112 | 2459.2
| Capacitance0.000010.00011 | Capacitance | 0.00011 | 0.00033 | 1e-05 | mF | nF |  |  |  | 5.785 | 482.33 | 5.785 | 480.98
| Capacitance0.00000010.00033 | Capacitance | 0.00033 | 0.0011 | 1e-07 | mF | nF |  |  |  | 2 | 4600 | 1.1566 | 2885.7
| Capacitance0.0000010.00033 | Capacitance | 0.00033 | 0.0011 | 1e-06 | mF | nF |  |  |  | 2 | 4600 | 1.2511 | 2834.7
| Capacitance0.000010.00033 | Capacitance | 0.00033 | 0.0011 | 1e-05 | mF | nF |  |  |  | 5.7156 | 1388.9 | 5.714 | 1385.8
| Capacitance0.00010.00033 | Capacitance | 0.00033 | 0.0011 | 0.0001 | mF | nF |  |  |  | 57.986 | 161.7 | 57.986 | 159.93
| Capacitance0.00000010.0011 | Capacitance | 0.0011 | 0.0033 | 1e-07 | mF | nF |  |  |  | 2 | 4600 | 3.4649 | 2886.6
| Capacitance0.0000010.0011 | Capacitance | 0.0011 | 0.0033 | 1e-06 | mF | nF |  |  |  | 2 | 4600 | 3.4959 | 2881.1
| Capacitance0.000010.0011 | Capacitance | 0.0011 | 0.0033 | 1e-05 | mF | nF |  |  |  | 4.6406 | 3799.8 | 6.1112 | 2459.2
| Capacitance0.00010.0011 | Capacitance | 0.0011 | 0.0033 | 0.0001 | mF | nF |  |  |  | 57.851 | 482.52 | 57.85 | 480.98
| Capacitance0.0000010.0033 | Capacitance | 0.0033 | 0.011 | 1e-06 | mF | nF |  |  |  | 19 | 5800 | 3.4973 | 2884.6
| Capacitance0.000010.0033 | Capacitance | 0.0033 | 0.011 | 1e-05 | mF | nF |  |  |  | 19 | 5800 | 5.0416 | 2786.9
| Capacitance0.00010.0033 | Capacitance | 0.0033 | 0.011 | 0.0001 | mF | nF |  |  |  | 49.454 | 3031.5 | 55.833 | 1093.2
| Capacitance0.0010.0033 | Capacitance | 0.0033 | 0.011 | 0.001 | mF | nF |  |  |  | 579.75 | 121.69 | 579.75 | 119.85
| Capacitance0.0000010.011 | Capacitance | 0.011 | 0.033 | 1e-06 | mF | nF |  |  |  | 23 | 5700 | 34.646 | 4618.7
| Capacitance0.000010.011 | Capacitance | 0.011 | 0.033 | 1e-05 | mF | nF |  |  |  | 23.275 | 5670.5 | 34.894 | 4613.9
| Capacitance0.00010.011 | Capacitance | 0.011 | 0.033 | 0.0001 | mF | nF |  |  |  | 49.728 | 4868.9 | 56.988 | 4207.9
| Capacitance0.0010.011 | Capacitance | 0.011 | 0.033 | 0.001 | mF | nF |  |  |  | 574.68 | 1053.3 | 574.68 | 1052.7
| Capacitance0.000010.033 | Capacitance | 0.033 | 0.11 | 1e-05 | mF | nF |  |  |  | 190 | 5800 | 115.62 | 5195.3
| Capacitance0.00010.033 | Capacitance | 0.033 | 0.11 | 0.0001 | mF | nF |  |  |  | 190 | 5800 | 122.79 | 5152.1
| Capacitance0.0010.033 | Capacitance | 0.033 | 0.11 | 0.001 | mF | nF |  |  |  | 539.3 | 3275.3 | 539.12 | 3273
| Capacitance0.010.033 | Capacitance | 0.033 | 0.11 | 0.01 | mF | nF |  |  |  | 5792.8 | 436.47 | 5792.8 | 434.51
| Capacitance0.000010.11 | Capacitance | 0.11 | 0.33 | 1e-05 | mF | µF |  |  |  | 0.19 | 5.8 | 0.0002561 | 1.5583
| Capacitance0.00010.11 | Capacitance | 0.11 | 0.33 | 0.0001 | mF | µF |  |  |  | 0.19 | 5.8 | 0.012811 | 1.53
| Capacitance0.0010.11 | Capacitance | 0.11 | 0.33 | 0.001 | mF | µF |  |  |  | 0.19 | 5.8 | 0.51963 | 0.77472
| Capacitance0.010.11 | Capacitance | 0.11 | 0.33 | 0.01 | mF | µF |  |  |  | 5.7925 | 0.092661 | 5.7924 | 0.091972
| Capacitance0.00010.33 | Capacitance | 0.33 | 1.1 | 0.0001 | mF | µF |  |  |  | 1.9 | 5.8 | 0.002539 | 5.1944
| Capacitance0.0010.33 | Capacitance | 0.33 | 1.1 | 0.001 | mF | µF |  |  |  | 1.9 | 5.8 | 0.12497 | 5.1095
| Capacitance0.010.33 | Capacitance | 0.33 | 1.1 | 0.01 | mF | µF |  |  |  | 5.096 | 2.8946 | 5.1507 | 2.7206
| Capacitance0.10.33 | Capacitance | 0.33 | 1.1 | 0.1 | mF | µF |  |  |  | 57.916 | 0.33394 | 57.916 | 0.33197
| Capacitance0.00011.1 | Capacitance | 1.1 | 3.3 | 0.0001 | mF | µF |  |  |  | 3.3 | 5.4 | 3.4645 | 5.1961
| Capacitance0.0011.1 | Capacitance | 1.1 | 3.3 | 0.001 | mF | µF |  |  |  | 3.3 | 5.3807 | 3.4878 | 5.1915
| Capacitance0.011.1 | Capacitance | 1.1 | 3.3 | 0.01 | mF | µF |  |  |  | 5.595 | 4.7979 | 5.5823 | 4.7969
| Capacitance0.11.1 | Capacitance | 1.1 | 3.3 | 0.1 | mF | µF |  |  |  | 57.311 | 1.2925 | 57.306 | 1.2871
| Capacitance0.0013.3 | Capacitance | 3.3 | 11 | 0.001 | mF | µF |  |  |  | 68 | 12 | 12.08 | 5.1953
| Capacitance0.013.3 | Capacitance | 3.3 | 11 | 0.01 | mF | µF |  |  |  | 68 | 12 | 12.723 | 5.1521
| Capacitance0.13.3 | Capacitance | 3.3 | 11 | 0.1 | mF | µF |  |  |  | 68 | 12 | 53.912 | 3.273
| Capacitance13.3 | Capacitance | 3.3 | 11 | 1 | mF | µF |  |  |  | 579.28 | 0.43758 | 579.28 | 0.43451
| Capacitance0.00111 | Capacitance | 11 | 33 | 0.001 | mF | µF |  |  |  | 68 | 12 | 34.644 | 8.6602
| Capacitance0.0111 | Capacitance | 11 | 33 | 0.01 | mF | µF |  |  |  | 68 | 12 | 34.812 | 8.6567
| Capacitance0.111 | Capacitance | 11 | 33 | 0.1 | mF | µF |  |  |  | 68 | 12 | 50.579 | 8.3351
| Capacitance111 | Capacitance | 11 | 33 | 1 | mF | µF |  |  |  | 560.27 | 3.1061 | 560.24 | 3.1027
| Capacitance0.0133 | Capacitance | 33 | 110 | 0.01 | mF | µF* |  |  |  | 120 | 13 | 115.57 | 12.701
| Capacitance0.133 | Capacitance | 33 | 110 | 0.1 | mF | µF* |  |  |  | 120 | 12.674 | 119.54 | 12.675
| Capacitance133 | Capacitance | 33 | 110 | 1 | mF | µF* |  |  |  | 433.38 | 10.795 | 432.61 | 10.795
| Capacitance1033 | Capacitance | 33 | 110 | 10 | mF | µF* |  |  |  | 5752.2 | 2.2058 | 5752 | 2.2003
| Frequency0.000010.001 | Frequency | 0.001 | 0.12 | 1e-05 | kHz | Hz |  |  |  | 0.006 | 0.013 | 0.0059069 | 0.012562
| Frequency0.00010.001 | Frequency | 0.001 | 0.12 | 0.0001 | kHz | Hz |  |  |  | 0.058029 | 0.0019067 | 0.058011 | 0.0014415
| Frequency0.0010.001 | Frequency | 0.001 | 0.12 | 0.001 | kHz | Hz |  |  |  | 0.58001 | 0.00033138 | 0.58 | 0.0001444
| Frequency0.010.001 | Frequency | 0.001 | 0.12 | 0.01 | kHz | Hz |  |  |  | 1.2926 | 8.924e-05 | 5.8 | 1.444e-05
| Frequency0.00010.12 | Frequency | 0.12 | 1.2 | 0.0001 | kHz | Hz |  |  |  | 0.058 | 0.0095 | 0.057075 | 0.0092344
| Frequency0.0010.12 | Frequency | 0.12 | 1.2 | 0.001 | kHz | Hz |  |  |  | 0.58007 | 0.0011513 | 0.5799 | 0.0010048
| Frequency0.010.12 | Frequency | 0.12 | 1.2 | 0.01 | kHz | Hz |  |  |  | 5.8001 | 0.00015944 | 5.8 | 0.00010057
| Frequency0.10.12 | Frequency | 0.12 | 1.2 | 0.1 | kHz | Hz |  |  |  | 58 | 3.3606e-05 | 58 | 1.0058e-05
| Frequency0.0011.2 | Frequency | 1.2 | 12 | 0.001 | kHz | Hz |  |  |  | 0.58 | 0.0091 | 0.57054 | 0.0088022
| Frequency0.011.2 | Frequency | 1.2 | 12 | 0.01 | kHz | Hz |  |  |  | 5.8007 | 0.0010997 | 5.799 | 0.00095317
| Frequency0.11.2 | Frequency | 1.2 | 12 | 0.1 | kHz | Hz |  |  |  | 5.9215 | 0.00015427 | 58 | 9.5402e-05
| Frequency11.2 | Frequency | 1.2 | 12 | 1 | kHz | Hz |  |  |  | 5.9216 | 3.3088e-05 | 580 | 9.5403e-06
| Frequency0.0112 | Frequency | 12 | 120 | 0.01 | kHz | Hz |  |  |  | 19 | 0.028 | 18.263 | 0.02763
| Frequency0.112 | Frequency | 12 | 120 | 0.1 | kHz | Hz |  |  |  | 60.591 | 0.0096232 | 60.522 | 0.0090804
| Frequency112 | Frequency | 12 | 120 | 1 | kHz | Hz |  |  |  | 580.29 | 0.0011916 | 580.26 | 0.00095637
| Frequency1012 | Frequency | 12 | 120 | 10 | kHz | Hz |  |  |  | 5800 | 0.00018988 | 5800 | 9.569e-05
| Frequency0.1120 | Frequency | 120 | 1200 | 0.1 | kHz | Hz |  |  |  | 61 | 0.016 | 59.805 | 0.015056
| Frequency1120 | Frequency | 120 | 1200 | 1 | kHz | Hz |  |  |  | 580.28 | 0.0023182 | 580.16 | 0.0018062
| Frequency10120 | Frequency | 120 | 1200 | 10 | kHz | Hz |  |  |  | 5800.1 | 0.00038706 | 5800 | 0.00018103
| Frequency100120 | Frequency | 120 | 1200 | 100 | kHz | Hz |  |  |  | 58000 | 0.00010052 | 58000 | 1.8104e-05
| Frequency11200 | Frequency | 1200 | 2000 | 1 | kHz | kHz |  |  |  | 0.44 | 0.13 | 0.42887 | 0.12787
| Frequency101200 | Frequency | 1200 | 2000 | 10 | kHz | kHz |  |  |  | 5.783 | 0.015929 | 5.7835 | 0.013952
| Frequency1001200 | Frequency | 1200 | 2000 | 100 | kHz | kHz |  |  |  | 57.998 | 0.0021913 | 57.998 | 0.0013965
| Frequency10001200 | Frequency | 1200 | 2000 | 1000 | kHz | kHz |  |  |  | 580 | 0.00045756 | 580 | 0.00013966

## Standard2
Updated: | 2019-05-10 | 3458A
|  | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCVmeas0.00000010 | DCVmeas | 0 | 0.1 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.61124 | 4.4552
| DCVmeas0.0000010 | DCVmeas | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 0.84074 | 3.6486 | 0.84024 | 3.6383
| DCVmeas0.000010 | DCVmeas | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 5.832 | 0.64501 | 5.8317 | 0.63535
| DCVmeas0.00010 | DCVmeas | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.003 | 0.068165 | 58.003 | 0.064226
| DCVmeas0.0010 | DCVmeas | 0 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.0079989 | 580 | 0.0064233
| DCVmeas0.010 | DCVmeas | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5800 | 0.0012726 | 5800 | 0.00064233
| DCVmeas0.00000010.1 | DCVmeas | 0.1 | 1 | 1e-07 | V | µV |  |  |  | 0.5 | 0 | 0.48404 | 0
| DCVmeas0.0000010.1 | DCVmeas | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 0.66084 | 0 | 0.6591 | 0
| DCVmeas0.000010.1 | DCVmeas | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 5.6669 | 0 | 5.6658 | 0
| DCVmeas0.00010.1 | DCVmeas | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 57.983 | 0 | 57.983 | 0
| DCVmeas0.0010.1 | DCVmeas | 0.1 | 1 | 0.001 | V | µV |  |  |  | 580 | 0 | 580 | 0
| DCVmeas0.010.1 | DCVmeas | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5800 | 0 | 5800 | 0
| DCVmeas0.00000011 | DCVmeas | 1 | 10 | 1e-07 | V | µV |  |  |  | 1.1 | 4.6 | 0.79598 | 4.6216
| DCVmeas0.0000011 | DCVmeas | 1 | 10 | 1e-06 | V | µV |  |  |  | 1.1 | 4.6 | 0.82953 | 4.6186
| DCVmeas0.000011 | DCVmeas | 1 | 10 | 1e-05 | V | µV |  |  |  | 3.5575 | 4.3812 | 3.5554 | 4.3813
| DCVmeas0.00011 | DCVmeas | 1 | 10 | 0.0001 | V | µV |  |  |  | 56.431 | 1.8232 | 56.429 | 1.8231
| DCVmeas0.0011 | DCVmeas | 1 | 10 | 0.001 | V | µV |  |  |  | 579.82 | 0.20862 | 579.82 | 0.20854
| DCVmeas0.011 | DCVmeas | 1 | 10 | 0.01 | V | µV |  |  |  | 5800 | 0.02092 | 5799.8 | 0.020888
| DCVmeas0.0000110 | DCVmeas | 10 | 100 | 1e-05 | V | µV |  |  |  | 36 | 7 | 35.28 | 6.9511
| DCVmeas0.000110 | DCVmeas | 10 | 100 | 0.0001 | V | µV |  |  |  | 50.925 | 6.8508 | 51.257 | 6.8111
| DCVmeas0.00110 | DCVmeas | 10 | 100 | 0.001 | V | µV |  |  |  | 551.23 | 3.8135 | 551.19 | 3.8123
| DCVmeas0.0110 | DCVmeas | 10 | 100 | 0.01 | V | µV |  |  |  | 5796 | 0.49898 | 5796 | 0.49811
| DCVmeas0.110 | DCVmeas | 10 | 100 | 0.1 | V | µV |  |  |  | 58000 | 0.050363 | 58000 | 0.050009
| DCVmeas0.0001100 | DCVmeas | 100 | 1000 | 0.0001 | V | mV |  |  |  | 0.16 | 0.0071 | 0.15283 | 0.0070612
| DCVmeas0.001100 | DCVmeas | 100 | 1000 | 0.001 | V | mV |  |  |  | 0.32538 | 0.0069346 | 0.32929 | 0.0068876
| DCVmeas0.01100 | DCVmeas | 100 | 1000 | 0.01 | V | mV |  |  |  | 5.4847 | 0.0037586 | 5.4844 | 0.0037561
| DCVmeas0.1100 | DCVmeas | 100 | 1000 | 0.1 | V | mV |  |  |  | 57.957 | 0.00048878 | 57.957 | 0.00048707
| DCVmeas1100 | DCVmeas | 100 | 1000 | 1 | V | mV |  |  |  | 580 | 4.9587e-05 | 580 | 4.8894e-05
| DCVmeas10100 | DCVmeas | 100 | 1000 | 10 | V | mV |  |  |  | 5800 | 5.1669e-06 | 5800 | 4.8896e-06
| DCV0.00000010 | DCV | 0 | 0.1 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.61128 | 4.4558
| DCV0.0000010 | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 0.84076 | 3.6493 | 0.84026 | 3.639
| DCV0.000010 | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 5.832 | 0.64516 | 5.8317 | 0.6355
| DCV0.00010 | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.003 | 0.06818 | 58.003 | 0.064242
| DCV0.0010 | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.0080005 | 580 | 0.0064249
| DCV0.010 | DCV | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5800 | 0.0012727 | 5800 | 0.00064249
| DCV0.00000010.1 | DCV | 0.1 | 1 | 1e-07 | V | µV |  |  |  | 0.5 | 4.7 | 0.48411 | 4.7118
| DCV0.0000010.1 | DCV | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 0.66097 | 4.5678 | 0.65925 | 4.5686
| DCV0.000010.1 | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 5.6669 | 2.1224 | 5.6658 | 2.121
| DCV0.00010.1 | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 57.983 | 0.25024 | 57.983 | 0.24934
| DCV0.0010.1 | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 580 | 0.025347 | 580 | 0.024985
| DCV0.010.1 | DCV | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5800 | 0.0026432 | 5800 | 0.0024986
| DCV0.00000011 | DCV | 1 | 10 | 1e-07 | V | µV |  |  |  | 1.1 | 4.6 | 0.796 | 4.6216
| DCV0.0000011 | DCV | 1 | 10 | 1e-06 | V | µV |  |  |  | 1.1 | 4.6 | 0.82955 | 4.6186
| DCV0.000011 | DCV | 1 | 10 | 1e-05 | V | µV |  |  |  | 3.5576 | 4.3812 | 3.5554 | 4.3813
| DCV0.00011 | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 56.431 | 1.8232 | 56.429 | 1.8231
| DCV0.0011 | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 579.82 | 0.20862 | 579.82 | 0.20854
| DCV0.011 | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5800 | 0.02092 | 5799.8 | 0.020888
| DCV0.0000110 | DCV | 10 | 100 | 1e-05 | V | µV |  |  |  | 36 | 7 | 35.283 | 6.9511
| DCV0.000110 | DCV | 10 | 100 | 0.0001 | V | µV |  |  |  | 50.93 | 6.8507 | 51.26 | 6.8111
| DCV0.00110 | DCV | 10 | 100 | 0.001 | V | µV |  |  |  | 551.23 | 3.8135 | 551.19 | 3.8123
| DCV0.0110 | DCV | 10 | 100 | 0.01 | V | µV |  |  |  | 5796 | 0.49899 | 5796 | 0.49811
| DCV0.110 | DCV | 10 | 100 | 0.1 | V | µV |  |  |  | 58000 | 0.050364 | 58000 | 0.050009
| DCV0.0001100 | DCV | 100 | 1000 | 0.0001 | V | mV |  |  |  | 0.16 | 0.0071 | 0.15289 | 0.0070612
| DCV0.001100 | DCV | 100 | 1000 | 0.001 | V | mV |  |  |  | 0.32543 | 0.0069346 | 0.32934 | 0.0068876
| DCV0.01100 | DCV | 100 | 1000 | 0.01 | V | mV |  |  |  | 5.4848 | 0.0037586 | 5.4845 | 0.0037561
| DCV0.1100 | DCV | 100 | 1000 | 0.1 | V | mV |  |  |  | 57.957 | 0.00048879 | 57.957 | 0.00048707
| DCV1100 | DCV | 100 | 1000 | 1 | V | mV |  |  |  | 580 | 4.9588e-05 | 580 | 4.8894e-05
| DCV10100 | DCV | 100 | 1000 | 10 | V | mV |  |  |  | 5800 | 5.1669e-06 | 5800 | 4.8896e-06
| DCAmeas0.000000000010 | DCAmeas | 0 | 1e-07 | 1e-11 | A | nA |  |  |  | 0.048 | 0.034 | 0.047325 | 0.034
| DCAmeas0.00000000010 | DCAmeas | 0 | 1e-07 | 1e-10 | A | nA |  |  |  | 0.074493 | 0.022232 | 0.074426 | 0.022232
| DCAmeas0.0000000010 | DCAmeas | 0 | 1e-07 | 1e-09 | A | nA |  |  |  | 0.5793 | 0.0030937 | 0.57926 | 0.0030937
| DCAmeas0.000000010 | DCAmeas | 0 | 1e-07 | 1e-08 | A | nA |  |  |  | 5.7738 | 0.0003879 | 5.7737 | 0.0003879
| DCAmeas0.00000010 | DCAmeas | 0 | 1e-07 | 1e-07 | A | nA |  |  |  | 57.736 | 6.9759e-05 | 57.736 | 6.9759e-05
| DCAmeas0.000000000010.0000001 | DCAmeas | 1e-07 | 1e-06 | 1e-11 | A | nA |  |  |  | 0.049 | 23000 | 0.047281 | 22626
| DCAmeas0.00000000010.0000001 | DCAmeas | 1e-07 | 1e-06 | 1e-10 | A | nA |  |  |  | 0.074342 | 16229 | 0.074234 | 16248
| DCAmeas0.0000000010.0000001 | DCAmeas | 1e-07 | 1e-06 | 1e-09 | A | nA |  |  |  | 0.57927 | 2327.7 | 0.57922 | 2327.9
| DCAmeas0.000000010.0000001 | DCAmeas | 1e-07 | 1e-06 | 1e-08 | A | nA |  |  |  | 5.7738 | 234.04 | 5.7737 | 234.05
| DCAmeas0.00000010.0000001 | DCAmeas | 1e-07 | 1e-06 | 1e-07 | A | nA |  |  |  | 57.736 | 23.406 | 57.736 | 23.406
| DCAmeas0.0000010.0000001 | DCAmeas | 1e-07 | 1e-06 | 1e-06 | A | nA |  |  |  | 577.36 | 2.3406 | 577.36 | 2.3406
| DCAmeas0.000000000010.000001 | DCAmeas | 1e-06 | 1e-05 | 1e-11 | A | nA |  |  |  | 0.12 | 24000 | 0.11535 | 23529
| DCAmeas0.00000000010.000001 | DCAmeas | 1e-06 | 1e-05 | 1e-10 | A | nA |  |  |  | 0.12705 | 23295 | 0.12751 | 22781
| DCAmeas0.0000000010.000001 | DCAmeas | 1e-06 | 1e-05 | 1e-09 | A | nA |  |  |  | 0.58478 | 9074 | 0.58472 | 9074.7
| DCAmeas0.000000010.000001 | DCAmeas | 1e-06 | 1e-05 | 1e-08 | A | nA |  |  |  | 5.7742 | 996.43 | 5.7742 | 996.43
| DCAmeas0.00000010.000001 | DCAmeas | 1e-06 | 1e-05 | 1e-07 | A | nA |  |  |  | 57.736 | 99.748 | 57.736 | 99.748
| DCAmeas0.0000010.000001 | DCAmeas | 1e-06 | 1e-05 | 1e-06 | A | nA |  |  |  | 577.36 | 9.9749 | 577.36 | 9.9749
| DCAmeas0.00000000010.00001 | DCAmeas | 1e-05 | 0.0001 | 1e-10 | A | nA |  |  |  | 0.12 | 23000 | 0.10201 | 23088
| DCAmeas0.0000000010.00001 | DCAmeas | 1e-05 | 0.0001 | 1e-09 | A | nA |  |  |  | 0.46304 | 20158 | 0.46239 | 20159
| DCAmeas0.000000010.00001 | DCAmeas | 1e-05 | 0.0001 | 1e-08 | A | nA |  |  |  | 5.7305 | 5267.9 | 5.7302 | 5261.6
| DCAmeas0.00000010.00001 | DCAmeas | 1e-05 | 0.0001 | 1e-07 | A | nA |  |  |  | 57.731 | 551.18 | 57.731 | 548.34
| DCAmeas0.0000010.00001 | DCAmeas | 1e-05 | 0.0001 | 1e-06 | A | nA |  |  |  | 577.35 | 55.994 | 577.35 | 54.858
| DCAmeas0.000010.00001 | DCAmeas | 1e-05 | 0.0001 | 1e-05 | A | nA |  |  |  | 5773.6 | 5.9401 | 5773.6 | 5.4859e-06
| DCAmeas0.000000010.0001 | DCAmeas | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 6 | 23000 | 5.7953 | 23112
| DCAmeas0.00000010.0001 | DCAmeas | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 7.7654 | 21710 | 7.7646 | 21708
| DCAmeas0.0000010.0001 | DCAmeas | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 57.603 | 6967.4 | 57.603 | 6962.7
| DCAmeas0.000010.0001 | DCAmeas | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 577.34 | 742.48 | 577.34 | 740.36
| DCAmeas0.00010.0001 | DCAmeas | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 5773.6 | 74.937 | 5773.6 | 74.085
| DCAmeas0.0010.0001 | DCAmeas | 0.0001 | 0.001 | 0.001 | A | nA |  |  |  | 57736 | 7.7492 | 57736 | 7.4086
| DCAmeas0.000000010.001 | DCAmeas | 0.001 | 0.01 | 1e-08 | A | µA |  |  |  | 0.059 | 23 | 0.057957 | 23.095
| DCAmeas0.00000010.001 | DCAmeas | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 0.077661 | 21.693 | 0.077653 | 21.691
| DCAmeas0.0000010.001 | DCAmeas | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 0.57604 | 6.9594 | 0.57603 | 6.9547
| DCAmeas0.000010.001 | DCAmeas | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.7734 | 0.74158 | 5.7734 | 0.73945
| DCAmeas0.00010.001 | DCAmeas | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.736 | 0.074846 | 57.736 | 0.073995
| DCAmeas0.0010.001 | DCAmeas | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 0.0077402 | 577.36 | 0.0073995
| DCAmeas0.00000010.01 | DCAmeas | 0.01 | 0.1 | 1e-07 | A | µA |  |  |  | 0.63 | 40 | 0.57903 | 40.423
| DCAmeas0.0000010.01 | DCAmeas | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 0.74996 | 39.081 | 0.74788 | 39.09
| DCAmeas0.000010.01 | DCAmeas | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 5.6869 | 17.112 | 5.6855 | 17.096
| DCAmeas0.00010.01 | DCAmeas | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 57.725 | 1.9687 | 57.724 | 1.9587
| DCAmeas0.0010.01 | DCAmeas | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.35 | 0.20021 | 577.35 | 0.19619
| DCAmeas0.010.01 | DCAmeas | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 0.021225 | 5773.6 | 0.01962
| DCAmeas0.0000010.1 | DCAmeas | 0.1 | 1 | 1e-06 | A | µA |  |  |  | 12 | 130 | 11.556 | 127.04
| DCAmeas0.000010.1 | DCAmeas | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 12 | 130 | 12.285 | 126.43
| DCAmeas0.00010.1 | DCAmeas | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 52.913 | 97.227 | 52.897 | 97.243
| DCAmeas0.0010.1 | DCAmeas | 0.1 | 1 | 0.001 | A | µA |  |  |  | 576.11 | 17.652 | 576.1 | 17.659
| DCAmeas0.010.1 | DCAmeas | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5773.4 | 1.7889 | 5773.4 | 1.7915
| DCAmeas0.10.1 | DCAmeas | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 0.17812 | 57735 | 0.17917
| DCAmeas0.000011 | DCAmeas | 1 | 20 | 1e-05 | A | mA |  |  |  | 0.008 | 0.061 | 0 | 0.061993
| DCAmeas0.00011 | DCAmeas | 1 | 20 | 0.0001 | A | mA |  |  |  | 0.011185 | 0.060547 | 0.011118 | 0.060547
| DCAmeas0.0011 | DCAmeas | 1 | 20 | 0.001 | A | mA |  |  |  | 0.53833 | 0.040613 | 0.53831 | 0.040602
| DCAmeas0.011 | DCAmeas | 1 | 20 | 0.01 | A | mA |  |  |  | 5.767 | 0.0067206 | 5.767 | 0.0067094
| DCAmeas0.11 | DCAmeas | 1 | 20 | 0.1 | A | mA |  |  |  | 57.735 | 0.00068288 | 57.735 | 0.00067829
| DCAmeas11 | DCAmeas | 1 | 20 | 1 | A | mA |  |  |  | 577.36 | 6.9672e-05 | 577.35 | 6.7836e-05
| DCA0.000000000010 | DCA | 0 | 1e-07 | 1e-11 | A | nA |  |  |  | 0.048 | 34000 | 0.047332 | 33576
| DCA0.00000000010 | DCA | 0 | 1e-07 | 1e-10 | A | nA |  |  |  | 0.074498 | 22261 | 0.07443 | 21792
| DCA0.0000000010 | DCA | 0 | 1e-07 | 1e-09 | A | nA |  |  |  | 0.5793 | 3097.7 | 0.57926 | 2840.8
| DCA0.000000010 | DCA | 0 | 1e-07 | 1e-08 | A | nA |  |  |  | 5.7738 | 388.3 | 5.7737 | 285.08
| DCA0.00000010 | DCA | 0 | 1e-07 | 1e-07 | A | nA |  |  |  | 57.736 | 69.799 | 57.736 | 28.509
| DCA0.000000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-11 | A | nA |  |  |  | 0.048 | 23000 | 0.047291 | 22623
| DCA0.00000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-10 | A | nA |  |  |  | 0.074348 | 16228 | 0.074241 | 16247
| DCA0.0000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-09 | A | nA |  |  |  | 0.57927 | 2327.7 | 0.57922 | 2327.9
| DCA0.000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-08 | A | nA |  |  |  | 5.7738 | 234.04 | 5.7737 | 234.05
| DCA0.00000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-07 | A | nA |  |  |  | 57.736 | 23.406 | 57.736 | 23.406
| DCA0.0000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-06 | A | nA |  |  |  | 577.36 | 2.3406 | 577.36 | 2.3406
| DCA0.000000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-11 | A | nA |  |  |  | 0.12 | 24000 | 0.11535 | 23529
| DCA0.00000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-10 | A | nA |  |  |  | 0.12706 | 23295 | 0.12752 | 22781
| DCA0.0000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-09 | A | nA |  |  |  | 0.58478 | 9074 | 0.58472 | 9074.7
| DCA0.000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-08 | A | nA |  |  |  | 5.7742 | 996.43 | 5.7742 | 996.43
| DCA0.00000010.000001 | DCA | 1e-06 | 1e-05 | 1e-07 | A | nA |  |  |  | 57.736 | 99.748 | 57.736 | 99.748
| DCA0.0000010.000001 | DCA | 1e-06 | 1e-05 | 1e-06 | A | nA |  |  |  | 577.36 | 9.9749 | 577.36 | 9.9749
| DCA0.000000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-11 | A | nA |  |  |  | 0.12 | 23000 | 0.10212 | 23000
| DCA0.00000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-10 | A | nA |  |  |  | 0.46307 | 20158 | 0.46242 | 20158
| DCA0.0000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-09 | A | nA |  |  |  | 5.7305 | 5268 | 5.7302 | 5268
| DCA0.000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-08 | A | nA |  |  |  | 57.731 | 551.19 | 57.731 | 551.19
| DCA0.00000010.00001 | DCA | 1e-05 | 0.0001 | 1e-07 | A | nA |  |  |  | 577.35 | 55.995 | 577.35 | 55.995
| DCA0.0000010.00001 | DCA | 1e-05 | 0.0001 | 1e-06 | A | nA |  |  |  | 5773.6 | 5.9402 | 5773.6 | 5.9402
| DCA0.000000010.0001 | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 6 | 23000 | 5.7954 | 23112
| DCA0.00000010.0001 | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 7.7655 | 21710 | 7.7647 | 21708
| DCA0.0000010.0001 | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 57.603 | 6967.5 | 57.603 | 6962.8
| DCA0.000010.0001 | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 577.34 | 742.49 | 577.34 | 740.36
| DCA0.00010.0001 | DCA | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 5773.6 | 74.938 | 5773.6 | 74.086
| DCA0.0010.0001 | DCA | 0.0001 | 0.001 | 0.001 | A | nA |  |  |  | 57736 | 7.7493 | 57736 | 7.4086
| DCA0.000000010.001 | DCA | 0.001 | 0.01 | 1e-08 | A | µA |  |  |  | 0.059 | 23 | 0.057958 | 23.095
| DCA0.00000010.001 | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 0.077662 | 21.693 | 0.077655 | 21.691
| DCA0.0000010.001 | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 0.57604 | 6.9594 | 0.57603 | 6.9548
| DCA0.000010.001 | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.7734 | 0.74159 | 5.7734 | 0.73946
| DCA0.00010.001 | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.736 | 0.074847 | 57.736 | 0.073996
| DCA0.0010.001 | DCA | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 0.0077402 | 577.36 | 0.0073996
| DCA0.00000010.01 | DCA | 0.01 | 0.1 | 1e-07 | A | µA |  |  |  | 0.63 | 40 | 0.57918 | 40.422
| DCA0.0000010.01 | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 0.75012 | 39.08 | 0.74803 | 39.089
| DCA0.000010.01 | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 5.6868 | 17.113 | 5.6855 | 17.097
| DCA0.00010.01 | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 57.725 | 1.9687 | 57.724 | 1.9587
| DCA0.0010.01 | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.35 | 0.20021 | 577.35 | 0.1962
| DCA0.010.01 | DCA | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 0.021225 | 5773.6 | 0.01962
| DCA0.0000010.1 | DCA | 0.1 | 1 | 1e-06 | A | µA |  |  |  | 12 | 130 | 11.556 | 127.04
| DCA0.000010.1 | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 12 | 130 | 12.285 | 126.43
| DCA0.00010.1 | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 52.913 | 97.227 | 52.898 | 97.242
| DCA0.0010.1 | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 576.11 | 17.652 | 576.1 | 17.659
| DCA0.010.1 | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5773.4 | 1.7889 | 5773.4 | 1.7915
| DCA0.10.1 | DCA | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 0.17812 | 57735 | 0.17917
| DCA0.000011 | DCA | 1 | 20 | 1e-05 | A | mA |  |  |  | 0.0086 | 0.06 | 0.0085 | 0.059958
| DCA0.00011 | DCA | 1 | 20 | 0.0001 | A | mA |  |  |  | 0.030481 | 0.05893 | 0.030432 | 0.058929
| DCA0.0011 | DCA | 1 | 20 | 0.001 | A | mA |  |  |  | 0.54154 | 0.039863 | 0.54152 | 0.039852
| DCA0.011 | DCA | 1 | 20 | 0.01 | A | mA |  |  |  | 5.7674 | 0.0065662 | 5.7674 | 0.006555
| DCA0.11 | DCA | 1 | 20 | 0.1 | A | mA |  |  |  | 57.735 | 0.00066713 | 57.735 | 0.00066254
| DCA11 | DCA | 1 | 20 | 1 | A | mA |  |  |  | 577.36 | 6.8097e-05 | 577.35 | 6.6261e-05
| Resistance0.000000010 | Resistance | 0 | 0.01 | 1e-08 | kΩ | µΩ |  |  |  | 62 | 17000 | 59.451 | 17218
| Resistance0.00000010 | Resistance | 0 | 0.01 | 1e-07 | kΩ | µΩ |  |  |  | 80.01 | 15875 | 79.93 | 15872
| Resistance0.0000010 | Resistance | 0 | 0.01 | 1e-06 | kΩ | µΩ |  |  |  | 578.03 | 4420.1 | 577.99 | 4407.3
| Resistance0.000010 | Resistance | 0 | 0.01 | 1e-05 | kΩ | µΩ |  |  |  | 5773.6 | 465.18 | 5773.6 | 459.51
| Resistance0.00010 | Resistance | 0 | 0.01 | 0.0001 | kΩ | µΩ |  |  |  | 57736 | 48.242 | 57736 | 45.971
| Resistance0.0010 | Resistance | 0 | 0.01 | 0.001 | kΩ | µΩ |  |  |  | 577360 | 5.5056 | 577360 | 4.5972
| Resistance0.000000010.01 | Resistance | 0.01 | 0.1 | 1e-08 | kΩ | µΩ |  |  |  | 580 | 14000 | 576.96 | 13918
| Resistance0.00000010.01 | Resistance | 0.01 | 0.1 | 1e-07 | kΩ | µΩ |  |  |  | 580 | 13901 | 579.48 | 13901
| Resistance0.0000010.01 | Resistance | 0.01 | 0.1 | 1e-06 | kΩ | µΩ |  |  |  | 794.26 | 12575 | 794.14 | 12575
| Resistance0.000010.01 | Resistance | 0.01 | 0.1 | 1e-05 | kΩ | µΩ |  |  |  | 5786.5 | 3136.4 | 5786.4 | 3135.4
| Resistance0.00010.01 | Resistance | 0.01 | 0.1 | 0.0001 | kΩ | µΩ |  |  |  | 57737 | 323.91 | 57737 | 323.5
| Resistance0.0010.01 | Resistance | 0.01 | 0.1 | 0.001 | kΩ | µΩ |  |  |  | 577360 | 32.524 | 577350 | 32.361
| Resistance0.00000010.1 | Resistance | 0.1 | 1 | 1e-07 | kΩ | µΩ |  |  |  | 590 | 12000 | 585.52 | 11553
| Resistance0.0000010.1 | Resistance | 0.1 | 1 | 1e-06 | kΩ | µΩ |  |  |  | 638.43 | 11952 | 686.25 | 11466
| Resistance0.000010.1 | Resistance | 0.1 | 1 | 1e-05 | kΩ | µΩ |  |  |  | 5207.3 | 8235.9 | 5206.4 | 8235.1
| Resistance0.00010.1 | Resistance | 0.1 | 1 | 0.0001 | kΩ | µΩ |  |  |  | 57625 | 1374.6 | 57624 | 1373.3
| Resistance0.0010.1 | Resistance | 0.1 | 1 | 0.001 | kΩ | µΩ |  |  |  | 577340 | 139.35 | 577340 | 138.85
| Resistance0.010.1 | Resistance | 0.1 | 1 | 0.01 | kΩ | µΩ |  |  |  | 5773600 | 14.086 | 5773500 | 13.887
| Resistance0.0000011 | Resistance | 1 | 10 | 1e-06 | kΩ | mΩ |  |  |  | 5.9 | 12 | 5.8027 | 11.551
| Resistance0.000011 | Resistance | 1 | 10 | 1e-05 | kΩ | mΩ |  |  |  | 6.3322 | 11.957 | 6.8126 | 11.464
| Resistance0.00011 | Resistance | 1 | 10 | 0.0001 | kΩ | mΩ |  |  |  | 52.068 | 8.2304 | 52.054 | 8.2296
| Resistance0.0011 | Resistance | 1 | 10 | 0.001 | kΩ | mΩ |  |  |  | 576.25 | 1.3734 | 576.24 | 1.3719
| Resistance0.011 | Resistance | 1 | 10 | 0.01 | kΩ | mΩ |  |  |  | 5773.4 | 0.13935 | 5773.4 | 0.1387
| Resistance0.11 | Resistance | 1 | 10 | 0.1 | kΩ | mΩ |  |  |  | 57736 | 0.014132 | 57735 | 0.013872
| Resistance0.0000110 | Resistance | 10 | 100 | 1e-05 | kΩ | mΩ |  |  |  | 58 | 12 | 57.628 | 11.584
| Resistance0.000110 | Resistance | 10 | 100 | 0.0001 | kΩ | mΩ |  |  |  | 63.382 | 11.946 | 67.785 | 11.496
| Resistance0.00110 | Resistance | 10 | 100 | 0.001 | kΩ | mΩ |  |  |  | 520.37 | 8.2597 | 520.23 | 8.2589
| Resistance0.0110 | Resistance | 10 | 100 | 0.01 | kΩ | mΩ |  |  |  | 5762.4 | 1.3801 | 5762.4 | 1.3785
| Resistance0.110 | Resistance | 10 | 100 | 0.1 | kΩ | mΩ |  |  |  | 57734 | 0.14003 | 57734 | 0.13938
| Resistance110 | Resistance | 10 | 100 | 1 | kΩ | mΩ |  |  |  | 577360 | 0.0142 | 577350 | 0.01394
| Resistance0.0001100 | Resistance | 100 | 1000 | 0.0001 | kΩ | Ω |  |  |  | 2.3 | 0.018 | 2.2863 | 0.017572
| Resistance0.001100 | Resistance | 100 | 1000 | 0.001 | kΩ | Ω |  |  |  | 2.3 | 0.018 | 2.3305 | 0.017536
| Resistance0.01100 | Resistance | 100 | 1000 | 0.01 | kΩ | Ω |  |  |  | 5.5381 | 0.015146 | 5.5337 | 0.015147
| Resistance0.1100 | Resistance | 100 | 1000 | 0.1 | kΩ | Ω |  |  |  | 57.526 | 0.0035348 | 57.524 | 0.0035314
| Resistance1100 | Resistance | 100 | 1000 | 1 | kΩ | Ω |  |  |  | 577.33 | 0.00036512 | 577.33 | 0.00036361
| Resistance10100 | Resistance | 100 | 1000 | 10 | kΩ | Ω |  |  |  | 5773.6 | 3.6975e-05 | 5773.5 | 3.6372e-05
| Resistance0.011000 | Resistance | 1000 | 10000 | 0.01 | kΩ | Ω |  |  |  | 120 | 0.059 | 114.65 | 0.058688
| Resistance0.11000 | Resistance | 1000 | 10000 | 0.1 | kΩ | Ω |  |  |  | 124.09 | 0.058591 | 124.69 | 0.057919
| Resistance11000 | Resistance | 1000 | 10000 | 1 | kΩ | Ω |  |  |  | 568.89 | 0.033992 | 568.81 | 0.033973
| Resistance101000 | Resistance | 1000 | 10000 | 10 | kΩ | Ω |  |  |  | 5771.7 | 0.0044433 | 5771.7 | 0.0044292
| Resistance1001000 | Resistance | 1000 | 10000 | 100 | kΩ | Ω |  |  |  | 57735 | 0.00045031 | 57735 | 0.00044464
| Resistance10001000 | Resistance | 1000 | 10000 | 1000 | kΩ | Ω |  |  |  | 577360 | 4.6736e-05 | 577350 | 4.4465e-05
| Resistance0.110000 | Resistance | 10000 | 100000 | 0.1 | kΩ | kΩ |  |  |  | 1.2 | 0.00058 | 1.1491 | 0.00057803
| Resistance110000 | Resistance | 10000 | 100000 | 1 | kΩ | kΩ |  |  |  | 1.2 | 0.00058 | 1.1755 | 0.00057779
| Resistance1010000 | Resistance | 10000 | 100000 | 10 | kΩ | kΩ |  |  |  | 3.4427 | 0.00055793 | 3.4399 | 0.00055794
| Resistance10010000 | Resistance | 10000 | 100000 | 100 | kΩ | kΩ |  |  |  | 55.444 | 0.00027077 | 55.443 | 0.00027072
| Resistance1000010000 | Resistance | 10000 | 100000 | 10000 | kΩ | kΩ |  |  |  | 5773.5 | 3.3092e-06 | 5773.5 | 3.2978e-06
| Resistance10000010000 | Resistance | 10000 | 100000 | 100000 | kΩ | kΩ |  |  |  | 57736 | 3.3093e-07 | 57735 | 3.2979e-07
| Resistance1100000 | Resistance | 100000 | 1000000 | 1 | kΩ | kΩ |  |  |  | 12 | 0.0058 | 11.203 | 0.0057772
| Resistance10100000 | Resistance | 100000 | 1000000 | 10 | kΩ | kΩ |  |  |  | 12 | 0.0058 | 11.204 | 0.0057772
| Resistance100100000 | Resistance | 100000 | 1000000 | 100 | kΩ | kΩ |  |  |  | 12 | 0.0058 | 11.212 | 0.0055471
| Resistance10000100000 | Resistance | 100000 | 1000000 | 10000 | kΩ | kΩ |  |  |  | 270.04 | 0.0055471 | 270.03 | 0.0055471
| Resistance100000100000 | Resistance | 100000 | 1000000 | 100000 | kΩ | kΩ |  |  |  | 5540 | 0.0026357 | 5540 | 0.0026356
| Resistance1000000100000 | Resistance | 100000 | 1000000 | 1000000 | kΩ | kΩ |  |  |  | 57707 | 0.00031828 | 57475 | 0.00031826
| ACVmeas0.000000100.001 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 3.6 | 350 | 3.5004 | 345.17
| ACVmeas0.00000100.001 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 3.6 | 350 | 3.5471 | 342.88
| ACVmeas0.0000100.001 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 6.7477 | 229.21 | 6.7466 | 229.01
| ACVmeas0.000100.001 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 57.841 | 31.346 | 57.84 | 31.206
| ACVmeas0.00100.001 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 577.37 | 3.1911 | 577.37 | 3.1347
| ACVmeas0.0100.001 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.6 | 0.33605 | 5773.6 | 0.31348
| ACVmeas0.000000100.04 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 0.04 | 1 | kHz | 1.4 | 230 | 1.3669 | 226.12
| ACVmeas0.00000100.04 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 0.04 | 1 | kHz | 1.4827 | 221.74 | 1.4817 | 219.16
| ACVmeas0.0000100.04 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 0.04 | 1 | kHz | 5.9304 | 89.221 | 5.9292 | 88.939
| ACVmeas0.000100.04 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 0.04 | 1 | kHz | 57.752 | 9.9552 | 57.751 | 9.8145
| ACVmeas0.00100.04 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.36 | 1.039 | 577.36 | 0.98255
| ACVmeas0.0100.04 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.6 | 0.12083 | 5773.6 | 0.098256
| ACVmeas0.000000101 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 1 | 20 | kHz | 1.6 | 330 | 1.5445 | 331.03
| ACVmeas0.00000101 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 1 | 20 | kHz | 1.6476 | 325.25 | 1.6466 | 324.22
| ACVmeas0.0000101 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 1 | 20 | kHz | 5.9704 | 157.65 | 5.9692 | 157.4
| ACVmeas0.000101 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 1 | 20 | kHz | 57.756 | 18.545 | 57.755 | 18.405
| ACVmeas0.00101 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 1 | 20 | kHz | 577.36 | 1.9004 | 577.36 | 1.844
| ACVmeas0.0101 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 1 | 20 | kHz | 5773.6 | 0.20698 | 5773.6 | 0.18441
| ACVmeas0.0000001020 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 20 | 50 | kHz | 1.9 | 1100 | 1.5841 | 1129.8
| ACVmeas0.000001020 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 20 | 50 | kHz | 1.9 | 1100 | 1.6792 | 1121.6
| ACVmeas0.00001020 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 20 | 50 | kHz | 5.9388 | 818.09 | 5.9356 | 818.11
| ACVmeas0.0001020 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 20 | 50 | kHz | 57.748 | 141.04 | 57.746 | 140.89
| ACVmeas0.001020 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 20 | 50 | kHz | 577.36 | 14.328 | 577.36 | 14.263
| ACVmeas0.01020 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 20 | 50 | kHz | 5773.6 | 1.4524 | 5773.6 | 1.4265
| ACVmeas0.0000001050 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 50 | 100 | kHz | 2.1 | 5700 | 1.6 | 5742.9
| ACVmeas0.000001050 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 50 | 100 | kHz | 2.1 | 5700 | 1.6756 | 5735.6
| ACVmeas0.00001050 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 50 | 100 | kHz | 5.6393 | 5367.4 | 5.6325 | 5367.8
| ACVmeas0.0001050 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 50 | 100 | kHz | 57.529 | 2504.9 | 57.526 | 2504.4
| ACVmeas0.001050 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 50 | 100 | kHz | 577.33 | 303.94 | 577.33 | 303.6
| ACVmeas0.01050 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 50 | 100 | kHz | 5773.6 | 30.576 | 5773.5 | 30.438
| ACVmeas0.00000010100 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.5802 | 46162
| ACVmeas0.0000010100 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.6033 | 46160
| ACVmeas0.000010100 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 100 | 300 | kHz | 4.648 | 45959 | 4.6346 | 45960
| ACVmeas0.00010100 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 100 | 300 | kHz | 54.056 | 41373 | 54.046 | 41373
| ACVmeas0.0010100 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 100 | 300 | kHz | 575.75 | 16509 | 575.75 | 16508
| ACVmeas0.010100 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 100 | 300 | kHz | 5773.4 | 1882 | 5771.9 | 1881.5
| ACVmeas0.00000010300 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 300 | 1000 | kHz | 7.8 | 46000 | 5.9623 | 46174
| ACVmeas0.0000010300 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 300 | 1000 | kHz | 7.8 | 46000 | 5.9767 | 46173
| ACVmeas0.000010300 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 300 | 1000 | kHz | 7.8 | 46000 | 7.4484 | 46029
| ACVmeas0.00010300 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 300 | 1000 | kHz | 54.541 | 41672 | 54.533 | 41672
| ACVmeas0.0010300 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 300 | 1000 | kHz | 575.78 | 16726 | 575.78 | 16725
| ACVmeas0.010300 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 300 | 1000 | kHz | 5773.4 | 1910 | 5771.9 | 1909.4
| ACVmeas0.000000101000 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 1000 | 4000 | kHz | 9.7 | 81000 | 9.678 | 80706
| ACVmeas0.00000101000 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 1000 | 4000 | kHz | 9.7 | 81000 | 9.6882 | 80705
| ACVmeas0.0000101000 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 1000 | 4000 | kHz | 10.605 | 80910 | 10.599 | 80616
| ACVmeas0.000101000 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 1000 | 4000 | kHz | 52.775 | 76693 | 52.74 | 76604
| ACVmeas0.00101000 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 1000 | 4000 | kHz | 573.38 | 42685 | 573.36 | 42684
| ACVmeas0.0101000 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 1000 | 4000 | kHz | 5773 | 5804.8 | 5769.3 | 5803.6
| ACVmeas0.000000104000 | ACVmeas | 0 | 0.01 | 1e-07 | V | µV | 4000 | 8000 | kHz | 20 | 230000 | 12.166 | 230700
| ACVmeas0.00000104000 | ACVmeas | 0 | 0.01 | 1e-06 | V | µV | 4000 | 8000 | kHz | 20 | 230000 | 12.171 | 230700
| ACVmeas0.0000104000 | ACVmeas | 0 | 0.01 | 1e-05 | V | µV | 4000 | 8000 | kHz | 20 | 230000 | 12.643 | 230650
| ACVmeas0.000104000 | ACVmeas | 0 | 0.01 | 0.0001 | V | µV | 4000 | 8000 | kHz | 44.989 | 227500 | 44.889 | 227500
| ACVmeas0.00104000 | ACVmeas | 0 | 0.01 | 0.001 | V | µV | 4000 | 8000 | kHz | 560.18 | 182980 | 560.13 | 182980
| ACVmeas0.0104000 | ACVmeas | 0 | 0.01 | 0.01 | V | µV | 4000 | 8000 | kHz | 5769.2 | 45281 | 5755.4 | 45279
| ACVmeas0.00000010.010.001 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 4.7 | 83 | 4.6245 | 83.307
| ACVmeas0.0000010.010.001 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 4.7 | 83 | 4.6567 | 83.113
| ACVmeas0.000010.010.001 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 7.2521 | 69.332 | 7.2514 | 69.32
| ACVmeas0.00010.010.001 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 57.862 | 13.107 | 57.862 | 13.092
| ACVmeas0.0010.010.001 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 577.37 | 1.3346 | 577.37 | 1.3282
| ACVmeas0.010.010.001 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.6 | 0.13538 | 5773.5 | 0.13284
| ACVmeas0.00000010.010.04 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 0.04 | 1 | kHz | 2.4 | 84 | 2.3369 | 83.638
| ACVmeas0.0000010.010.04 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 0.04 | 1 | kHz | 2.4 | 84 | 2.3925 | 83.236
| ACVmeas0.000010.010.04 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 0.04 | 1 | kHz | 5.9698 | 61.912 | 5.9689 | 61.898
| ACVmeas0.00010.010.04 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 0.04 | 1 | kHz | 57.723 | 9.9724 | 57.723 | 9.9569
| ACVmeas0.0010.010.04 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.35 | 1.0112 | 577.35 | 1.0048
| ACVmeas0.010.010.04 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.6 | 0.10303 | 5773.5 | 0.10049
| ACVmeas0.00000010.011 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 1 | 20 | kHz | 2.8 | 160 | 2.4423 | 162.82
| ACVmeas0.0000010.011 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 1 | 20 | kHz | 2.8 | 160 | 2.4862 | 162.47
| ACVmeas0.000010.011 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 1 | 20 | kHz | 5.6727 | 139.23 | 5.6718 | 139.22
| ACVmeas0.00010.011 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 1 | 20 | kHz | 57.566 | 31.314 | 57.566 | 31.3
| ACVmeas0.0010.011 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 1 | 20 | kHz | 577.34 | 3.2196 | 577.34 | 3.2132
| ACVmeas0.010.011 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.6 | 0.32395 | 5773.5 | 0.32141
| ACVmeas0.00000010.0120 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 20 | 50 | kHz | 2.5 | 350 | 2.4389 | 347.76
| ACVmeas0.0000010.0120 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 20 | 50 | kHz | 2.5 | 350 | 2.4694 | 347.5
| ACVmeas0.000010.0120 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 20 | 50 | kHz | 5.0039 | 326.59 | 5.0004 | 326.6
| ACVmeas0.00010.0120 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 20 | 50 | kHz | 56.856 | 118.41 | 56.854 | 118.36
| ACVmeas0.0010.0120 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 20 | 50 | kHz | 577.26 | 12.999 | 577.26 | 12.976
| ACVmeas0.010.0120 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 20 | 50 | kHz | 5773.5 | 1.3082 | 5773.4 | 1.299
| ACVmeas0.00000010.0150 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 50 | 100 | kHz | 2.5 | 930 | 2.4165 | 925.07
| ACVmeas0.0000010.0150 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 50 | 100 | kHz | 2.5 | 930 | 2.4316 | 924.94
| ACVmeas0.000010.0150 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 50 | 100 | kHz | 3.8611 | 916.39 | 3.897 | 912.02
| ACVmeas0.00010.0150 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 50 | 100 | kHz | 53.109 | 580.09 | 53.102 | 580.01
| ACVmeas0.0010.0150 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 50 | 100 | kHz | 576.63 | 84.904 | 576.62 | 84.816
| ACVmeas0.010.0150 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 50 | 100 | kHz | 5773.5 | 8.5746 | 5772.7 | 8.5388
| ACVmeas0.00000010.01100 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.626 | 3465.3
| ACVmeas0.0000010.01100 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.633 | 3465.2
| ACVmeas0.000010.01100 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 100 | 300 | kHz | 12 | 3500 | 12.02 | 3461.8
| ACVmeas0.00010.01100 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 100 | 300 | kHz | 41.934 | 3208.6 | 41.907 | 3208.7
| ACVmeas0.0010.01100 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 100 | 300 | kHz | 568.08 | 1113.7 | 568.07 | 1113.5
| ACVmeas0.010.01100 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 100 | 300 | kHz | 5772.5 | 121.36 | 5762.6 | 121.25
| ACVmeas0.00000010.01300 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.64 | 11557
| ACVmeas0.0000010.01300 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.641 | 11557
| ACVmeas0.000010.01300 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.767 | 11556
| ACVmeas0.00010.01300 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 300 | 1000 | kHz | 20.584 | 11914 | 25.368 | 11434
| ACVmeas0.0010.01300 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 300 | 1000 | kHz | 512.23 | 7901.3 | 512.19 | 7901.3
| ACVmeas0.010.01300 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 300 | 1000 | kHz | 5762.1 | 1282.8 | 5695.9 | 1282.5
| ACVmeas0.00000010.011000 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 48 | 17000 | 11.629 | 17356
| ACVmeas0.0000010.011000 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 48 | 17000 | 11.63 | 17356
| ACVmeas0.000010.011000 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 48 | 17000 | 11.729 | 17355
| ACVmeas0.00010.011000 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 48 | 17002 | 21.29 | 17269
| ACVmeas0.0010.011000 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 1000 | 2000 | kHz | 469.31 | 13709 | 469.24 | 13709
| ACVmeas0.010.011000 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 1000 | 2000 | kHz | 5748.1 | 2840.4 | 5639.4 | 2840.1
| ACVmeas0.00000010.012000 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.866 | 46201
| ACVmeas0.0000010.012000 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.866 | 46201
| ACVmeas0.000010.012000 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.879 | 46201
| ACVmeas0.00010.012000 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 84.221 | 46171
| ACVmeas0.0010.012000 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 2000 | 4000 | kHz | 354.38 | 43820 | 354.29 | 43820
| ACVmeas0.010.012000 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 2000 | 4000 | kHz | 5616.1 | 18294 | 5360.8 | 18293
| ACVmeas0.00000010.014000 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.339 | 46247
| ACVmeas0.0000010.014000 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.339 | 46247
| ACVmeas0.000010.014000 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.411 | 46246
| ACVmeas0.00010.014000 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 95.64 | 46217
| ACVmeas0.0010.014000 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 4000 | 8000 | kHz | 362.15 | 43903 | 361.69 | 43905
| ACVmeas0.010.014000 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 4000 | 8000 | kHz | 5616.4 | 18396 | 5361.1 | 18392
| ACVmeas0.00000010.018000 | ACVmeas | 0.01 | 0.1 | 1e-07 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 116.3 | 173250
| ACVmeas0.0000010.018000 | ACVmeas | 0.01 | 0.1 | 1e-06 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 116.31 | 173250
| ACVmeas0.000010.018000 | ACVmeas | 0.01 | 0.1 | 1e-05 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 116.31 | 173250
| ACVmeas0.00010.018000 | ACVmeas | 0.01 | 0.1 | 0.0001 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 117.4 | 173240
| ACVmeas0.0010.018000 | ACVmeas | 0.01 | 0.1 | 0.001 | V | µV | 8000 | 10000 | kHz | 450.06 | 170010 | 212.86 | 172380
| ACVmeas0.010.018000 | ACVmeas | 0.01 | 0.1 | 0.01 | V | µV | 8000 | 10000 | kHz | 4695.2 | 136780 | 4338.4 | 136780
| ACVmeas0.0000010.10.001 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 46.159 | 81.751
| ACVmeas0.000010.10.001 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 46.481 | 81.558
| ACVmeas0.00010.10.001 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 72.504 | 67.846 | 72.496 | 67.84
| ACVmeas0.0010.10.001 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 578.64 | 12.73 | 578.63 | 12.72
| ACVmeas0.010.10.001 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.7 | 1.294 | 5773.7 | 1.2901
| ACVmeas0.10.10.001 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 0.001 | 0.04 | kHz | 57736 | 0.1306 | 57735 | 0.12903
| ACVmeas0.0000010.10.04 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 23.095 | 81.918
| ACVmeas0.000010.10.04 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 23.659 | 81.511
| ACVmeas0.00010.10.04 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 0.04 | 1 | kHz | 59.657 | 60.196 | 59.646 | 60.19
| ACVmeas0.0010.10.04 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.25 | 9.5934 | 577.24 | 9.5838
| ACVmeas0.010.10.04 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.5 | 0.97079 | 5773.5 | 0.96685
| ACVmeas0.10.10.04 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 0.04 | 1 | kHz | 57736 | 0.098269 | 57735 | 0.096694
| ACVmeas0.0000010.11 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.156 | 162.3
| ACVmeas0.000010.11 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.61 | 161.93
| ACVmeas0.00010.11 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 1 | 20 | kHz | 56.098 | 138.17 | 56.07 | 138.16
| ACVmeas0.0010.11 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 1 | 20 | kHz | 575.63 | 30.824 | 575.62 | 30.791
| ACVmeas0.010.11 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.4 | 3.1737 | 5773.4 | 3.1593
| ACVmeas0.10.11 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 1 | 20 | kHz | 57736 | 0.3218 | 57735 | 0.31601
| ACVmeas0.0000010.120 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.229 | 346.82
| ACVmeas0.000010.120 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.536 | 346.56
| ACVmeas0.00010.120 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 20 | 50 | kHz | 49.328 | 325.25 | 49.245 | 325.28
| ACVmeas0.0010.120 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 20 | 50 | kHz | 568.57 | 117.3 | 568.53 | 117.24
| ACVmeas0.010.120 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 20 | 50 | kHz | 5772.6 | 12.875 | 5772.6 | 12.84
| ACVmeas0.10.120 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 20 | 50 | kHz | 57735 | 1.2992 | 57734 | 1.2854
| ACVmeas0.0000010.150 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 50 | 100 | kHz | 28 | 920 | 23.289 | 923.8
| ACVmeas0.000010.150 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 50 | 100 | kHz | 28 | 920 | 23.447 | 923.66
| ACVmeas0.00010.150 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 50 | 100 | kHz | 38.36 | 910.53 | 38.211 | 910.64
| ACVmeas0.0010.150 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 50 | 100 | kHz | 531.13 | 578.22 | 531.01 | 578.19
| ACVmeas0.010.150 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 50 | 100 | kHz | 5766.3 | 84.519 | 5766.3 | 84.451
| ACVmeas0.10.150 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 50 | 100 | kHz | 57735 | 8.5295 | 57727 | 8.5018
| ACVmeas0.0000010.1100 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 100 | 300 | kHz | 120 | 3500 | 115.61 | 3464.2
| ACVmeas0.000010.1100 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 100 | 300 | kHz | 120 | 3500 | 115.68 | 3464.1
| ACVmeas0.00010.1100 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 100 | 300 | kHz | 120 | 3500 | 119.56 | 3460.7
| ACVmeas0.0010.1100 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 100 | 300 | kHz | 419.06 | 3207.1 | 418.74 | 3207.3
| ACVmeas0.010.1100 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 100 | 300 | kHz | 5680.9 | 1112.6 | 5680.8 | 1112.5
| ACVmeas0.10.1100 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 100 | 300 | kHz | 57725 | 121.19 | 57626 | 121.14
| ACVmeas0.0000010.1300 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 116.27 | 11549
| ACVmeas0.000010.1300 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 116.29 | 11549
| ACVmeas0.00010.1300 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 117.69 | 11548
| ACVmeas0.0010.1300 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 300 | 1000 | kHz | 204.93 | 11915 | 253.55 | 11426
| ACVmeas0.010.1300 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 300 | 1000 | kHz | 5122.8 | 7893.3 | 5122.5 | 7893.4
| ACVmeas0.10.1300 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 300 | 1000 | kHz | 57622 | 1280.9 | 56960 | 1280.8
| ACVmeas0.0000010.11000 | ACVmeas | 0.1 | 1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 470 | 17000 | 116.68 | 17345
| ACVmeas0.000010.11000 | ACVmeas | 0.1 | 1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 470 | 17000 | 116.68 | 17345
| ACVmeas0.00010.11000 | ACVmeas | 0.1 | 1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 470 | 17000 | 117.78 | 17344
| ACVmeas0.0010.11000 | ACVmeas | 0.1 | 1 | 0.001 | V | µV | 1000 | 2000 | kHz | 470.28 | 17001 | 213.22 | 17258
| ACVmeas0.010.11000 | ACVmeas | 0.1 | 1 | 0.01 | V | µV | 1000 | 2000 | kHz | 4693.7 | 13698 | 4693.4 | 13698
| ACVmeas0.10.11000 | ACVmeas | 0.1 | 1 | 0.1 | V | µV | 1000 | 2000 | kHz | 57482 | 2836.9 | 56395 | 2836.9
| ACVmeas0.0000010.12000 | ACVmeas | 0.1 | 1 | 1e-06 | V | mV | 2000 | 4000 | kHz | 1.1 | 46 | 0.80743 | 46.247
| ACVmeas0.000010.12000 | ACVmeas | 0.1 | 1 | 1e-05 | V | mV | 2000 | 4000 | kHz | 1.1 | 46 | 0.80744 | 46.247
| ACVmeas0.00010.12000 | ACVmeas | 0.1 | 1 | 0.0001 | V | mV | 2000 | 4000 | kHz | 1.1 | 46 | 0.80808 | 46.246
| ACVmeas0.0010.12000 | ACVmeas | 0.1 | 1 | 0.001 | V | mV | 2000 | 4000 | kHz | 1.1 | 46 | 0.84158 | 46.216
| ACVmeas0.010.12000 | ACVmeas | 0.1 | 1 | 0.01 | V | mV | 2000 | 4000 | kHz | 3.5417 | 43.866 | 3.541 | 43.866
| ACVmeas0.10.12000 | ACVmeas | 0.1 | 1 | 0.1 | V | mV | 2000 | 4000 | kHz | 56.159 | 18.324 | 53.604 | 18.323
| ACVmeas0.0000010.14000 | ACVmeas | 0.1 | 1 | 1e-06 | V | mV | 4000 | 8000 | kHz | 1.2 | 46 | 0.9228 | 46.247
| ACVmeas0.000010.14000 | ACVmeas | 0.1 | 1 | 1e-05 | V | mV | 4000 | 8000 | kHz | 1.2 | 46 | 0.92281 | 46.247
| ACVmeas0.00010.14000 | ACVmeas | 0.1 | 1 | 0.0001 | V | mV | 4000 | 8000 | kHz | 1.2 | 46 | 0.9234 | 46.246
| ACVmeas0.0010.14000 | ACVmeas | 0.1 | 1 | 0.001 | V | mV | 4000 | 8000 | kHz | 1.2 | 46 | 0.9559 | 46.217
| ACVmeas0.010.14000 | ACVmeas | 0.1 | 1 | 0.01 | V | mV | 4000 | 8000 | kHz | 3.6172 | 43.905 | 3.6164 | 43.905
| ACVmeas0.10.14000 | ACVmeas | 0.1 | 1 | 0.1 | V | mV | 4000 | 8000 | kHz | 56.163 | 18.393 | 53.611 | 18.392
| ACVmeas0.0000010.18000 | ACVmeas | 0.1 | 1 | 1e-06 | V | mV | 8000 | 10000 | kHz | 4.5 | 170 | 1.155 | 173.26
| ACVmeas0.000010.18000 | ACVmeas | 0.1 | 1 | 1e-05 | V | mV | 8000 | 10000 | kHz | 4.5 | 170 | 1.1551 | 173.26
| ACVmeas0.00010.18000 | ACVmeas | 0.1 | 1 | 0.0001 | V | mV | 8000 | 10000 | kHz | 4.5 | 170 | 1.1551 | 173.26
| ACVmeas0.0010.18000 | ACVmeas | 0.1 | 1 | 0.001 | V | mV | 8000 | 10000 | kHz | 4.5 | 170 | 1.166 | 173.25
| ACVmeas0.010.18000 | ACVmeas | 0.1 | 1 | 0.01 | V | mV | 8000 | 10000 | kHz | 4.503 | 170.01 | 2.1222 | 172.39
| ACVmeas0.10.18000 | ACVmeas | 0.1 | 1 | 0.1 | V | mV | 8000 | 10000 | kHz | 46.972 | 136.76 | 43.382 | 136.78
| ACVmeas0.0000110.001 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 0.001 | 0.04 | kHz | 0.47 | 0.083 | 0.45985 | 0.0831
| ACVmeas0.000110.001 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 0.47 | 0.083 | 0.46308 | 0.082905
| ACVmeas0.00110.001 | ACVmeas | 1 | 10 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 0.72377 | 0.0691 | 0.72347 | 0.06906
| ACVmeas0.0110.001 | ACVmeas | 1 | 10 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 5.7861 | 0.013061 | 5.786 | 0.013008
| ACVmeas0.110.001 | ACVmeas | 1 | 10 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.737 | 0.0013416 | 57.737 | 0.0013195
| ACVmeas110.001 | ACVmeas | 1 | 10 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.36 | 0.00014081 | 577.35 | 0.00013197
| ACVmeas0.0000110.04 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 0.04 | 1 | kHz | 0.24 | 0.082 | 0.23104 | 0.081982
| ACVmeas0.000110.04 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 0.04 | 1 | kHz | 0.24 | 0.082 | 0.23589 | 0.081574
| ACVmeas0.00110.04 | ACVmeas | 1 | 10 | 0.001 | V | mV | 0.04 | 1 | kHz | 0.59647 | 0.060263 | 0.59612 | 0.060216
| ACVmeas0.0110.04 | ACVmeas | 1 | 10 | 0.01 | V | mV | 0.04 | 1 | kHz | 5.7725 | 0.00964 | 5.7724 | 0.0095859
| ACVmeas0.110.04 | ACVmeas | 1 | 10 | 0.1 | V | mV | 0.04 | 1 | kHz | 57.735 | 0.00098916 | 57.735 | 0.00096707
| ACVmeas110.04 | ACVmeas | 1 | 10 | 1 | V | mV | 0.04 | 1 | kHz | 577.36 | 0.00010556 | 577.35 | 9.6715e-05
| ACVmeas0.0000111 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 1 | 20 | kHz | 0.26 | 0.16 | 0.23132 | 0.16229
| ACVmeas0.000111 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 1 | 20 | kHz | 0.26 | 0.16 | 0.23522 | 0.16193
| ACVmeas0.00111 | ACVmeas | 1 | 10 | 0.001 | V | mV | 1 | 20 | kHz | 0.56168 | 0.13811 | 0.56026 | 0.13812
| ACVmeas0.0111 | ACVmeas | 1 | 10 | 0.01 | V | mV | 1 | 20 | kHz | 5.7567 | 0.030878 | 5.7561 | 0.030767
| ACVmeas0.111 | ACVmeas | 1 | 10 | 0.1 | V | mV | 1 | 20 | kHz | 57.734 | 0.0032046 | 57.734 | 0.0031567
| ACVmeas111 | ACVmeas | 1 | 10 | 1 | V | mV | 1 | 20 | kHz | 577.36 | 0.00033493 | 577.35 | 0.00031576
| ACVmeas0.00001120 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.2306 | 0.34716
| ACVmeas0.0001120 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.23368 | 0.3469
| ACVmeas0.001120 | ACVmeas | 1 | 10 | 0.001 | V | mV | 20 | 50 | kHz | 0.49448 | 0.32537 | 0.49117 | 0.32558
| ACVmeas0.01120 | ACVmeas | 1 | 10 | 0.01 | V | mV | 20 | 50 | kHz | 5.6867 | 0.11746 | 5.685 | 0.11736
| ACVmeas0.1120 | ACVmeas | 1 | 10 | 0.1 | V | mV | 20 | 50 | kHz | 57.726 | 0.012918 | 57.726 | 0.012854
| ACVmeas1120 | ACVmeas | 1 | 10 | 1 | V | mV | 20 | 50 | kHz | 577.35 | 0.0013125 | 577.34 | 0.0012868
| ACVmeas0.00001150 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 50 | 100 | kHz | 0.28 | 0.92 | 0.23089 | 0.92405
| ACVmeas0.0001150 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 50 | 100 | kHz | 0.28 | 0.92 | 0.23243 | 0.92391
| ACVmeas0.001150 | ACVmeas | 1 | 10 | 0.001 | V | mV | 50 | 100 | kHz | 0.38423 | 0.91055 | 0.38032 | 0.91086
| ACVmeas0.01150 | ACVmeas | 1 | 10 | 0.01 | V | mV | 50 | 100 | kHz | 5.3128 | 0.57823 | 5.3097 | 0.57827
| ACVmeas0.1150 | ACVmeas | 1 | 10 | 0.1 | V | mV | 50 | 100 | kHz | 57.664 | 0.084554 | 57.663 | 0.084464
| ACVmeas1150 | ACVmeas | 1 | 10 | 1 | V | mV | 50 | 100 | kHz | 577.35 | 0.0085405 | 577.27 | 0.0085031
| ACVmeas0.000011100 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1622 | 3.4645
| ACVmeas0.00011100 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1627 | 3.4644
| ACVmeas0.0011100 | ACVmeas | 1 | 10 | 0.001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.194 | 3.461
| ACVmeas0.011100 | ACVmeas | 1 | 10 | 0.01 | V | mV | 100 | 300 | kHz | 4.1916 | 3.2072 | 4.1867 | 3.2075
| ACVmeas0.11100 | ACVmeas | 1 | 10 | 0.1 | V | mV | 100 | 300 | kHz | 56.81 | 1.1127 | 56.808 | 1.1126
| ACVmeas11100 | ACVmeas | 1 | 10 | 1 | V | mV | 100 | 300 | kHz | 577.25 | 0.12122 | 576.26 | 0.12115
| ACVmeas0.000011300 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 300 | 1000 | kHz | 1.2 | 12 | 1.1976 | 11.551
| ACVmeas0.00011300 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 300 | 1000 | kHz | 1.2 | 12 | 1.1977 | 11.551
| ACVmeas0.0011300 | ACVmeas | 1 | 10 | 0.001 | V | mV | 300 | 1000 | kHz | 1.2 | 12 | 1.2052 | 11.55
| ACVmeas0.011300 | ACVmeas | 1 | 10 | 0.01 | V | mV | 300 | 1000 | kHz | 2.0433 | 11.916 | 2.5267 | 11.428
| ACVmeas0.11300 | ACVmeas | 1 | 10 | 0.1 | V | mV | 300 | 1000 | kHz | 51.227 | 7.8943 | 51.222 | 7.8945
| ACVmeas11300 | ACVmeas | 1 | 10 | 1 | V | mV | 300 | 1000 | kHz | 576.22 | 1.2811 | 569.6 | 1.281
| ACVmeas0.0000111000 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 1000 | 2000 | kHz | 4.6 | 17 | 1.5088 | 17.334
| ACVmeas0.000111000 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 1000 | 2000 | kHz | 4.6 | 17 | 1.5089 | 17.334
| ACVmeas0.00111000 | ACVmeas | 1 | 10 | 0.001 | V | mV | 1000 | 2000 | kHz | 4.6 | 17 | 1.515 | 17.333
| ACVmeas0.0111000 | ACVmeas | 1 | 10 | 0.01 | V | mV | 1000 | 2000 | kHz | 4.6017 | 17 | 2.1325 | 17.247
| ACVmeas0.111000 | ACVmeas | 1 | 10 | 0.1 | V | mV | 1000 | 2000 | kHz | 46.946 | 13.687 | 46.94 | 13.687
| ACVmeas111000 | ACVmeas | 1 | 10 | 1 | V | mV | 1000 | 2000 | kHz | 574.82 | 2.8334 | 563.96 | 2.8333
| ACVmeas0.0000112000 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 2000 | 4000 | kHz | 11 | 46 | 8.1106 | 46.211
| ACVmeas0.000112000 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 2000 | 4000 | kHz | 11 | 46 | 8.1107 | 46.211
| ACVmeas0.00112000 | ACVmeas | 1 | 10 | 0.001 | V | mV | 2000 | 4000 | kHz | 11 | 46 | 8.113 | 46.211
| ACVmeas0.0112000 | ACVmeas | 1 | 10 | 0.01 | V | mV | 2000 | 4000 | kHz | 11 | 46 | 8.4461 | 46.181
| ACVmeas0.112000 | ACVmeas | 1 | 10 | 0.1 | V | mV | 2000 | 4000 | kHz | 35.448 | 43.831 | 35.442 | 43.831
| ACVmeas112000 | ACVmeas | 1 | 10 | 1 | V | mV | 2000 | 4000 | kHz | 561.61 | 18.302 | 536.07 | 18.301
| ACVmeas0.0000114000 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 4000 | 8000 | kHz | 12 | 46 | 9.3016 | 46.242
| ACVmeas0.000114000 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 4000 | 8000 | kHz | 12 | 46 | 9.3017 | 46.242
| ACVmeas0.00114000 | ACVmeas | 1 | 10 | 0.001 | V | mV | 4000 | 8000 | kHz | 12 | 46 | 9.302 | 46.242
| ACVmeas0.0114000 | ACVmeas | 1 | 10 | 0.01 | V | mV | 4000 | 8000 | kHz | 12 | 46 | 9.627 | 46.213
| ACVmeas0.114000 | ACVmeas | 1 | 10 | 0.1 | V | mV | 4000 | 8000 | kHz | 36.218 | 43.903 | 36.212 | 43.903
| ACVmeas114000 | ACVmeas | 1 | 10 | 1 | V | mV | 4000 | 8000 | kHz | 561.63 | 18.395 | 536.12 | 18.393
| ACVmeas0.0000118000 | ACVmeas | 1 | 10 | 1e-05 | V | mV | 8000 | 10000 | kHz | 44 | 170 | 11.574 | 173.23
| ACVmeas0.000118000 | ACVmeas | 1 | 10 | 0.0001 | V | mV | 8000 | 10000 | kHz | 44 | 170 | 11.575 | 173.23
| ACVmeas0.00118000 | ACVmeas | 1 | 10 | 0.001 | V | mV | 8000 | 10000 | kHz | 44 | 170 | 11.576 | 173.23
| ACVmeas0.0118000 | ACVmeas | 1 | 10 | 0.01 | V | mV | 8000 | 10000 | kHz | 44 | 170 | 11.675 | 173.22
| ACVmeas0.118000 | ACVmeas | 1 | 10 | 0.1 | V | mV | 8000 | 10000 | kHz | 44 | 170.09 | 21.301 | 172.35
| ACVmeas118000 | ACVmeas | 1 | 10 | 1 | V | mV | 8000 | 10000 | kHz | 469.51 | 136.75 | 433.86 | 136.75
| ACVmeas0.0001100.001 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 4.8 | 0.23 | 4.6192 | 0.23147
| ACVmeas0.001100.001 | ACVmeas | 10 | 100 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 4.8 | 0.23 | 4.6433 | 0.23128
| ACVmeas0.01100.001 | ACVmeas | 10 | 100 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 6.8737 | 0.21486 | 6.8729 | 0.21486
| ACVmeas0.1100.001 | ACVmeas | 10 | 100 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.493 | 0.065725 | 57.493 | 0.065717
| ACVmeas1100.001 | ACVmeas | 10 | 100 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.33 | 0.0069549 | 577.33 | 0.0069512
| ACVmeas10100.001 | ACVmeas | 10 | 100 | 10 | V | mV | 0.001 | 0.04 | kHz | 5773.5 | 0.000697 | 5773.5 | 0.00069555
| ACVmeas0.0001100.04 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 0.04 | 1 | kHz | 2.5 | 0.23 | 2.3097 | 0.23154
| ACVmeas0.001100.04 | ACVmeas | 10 | 100 | 0.001 | V | mV | 0.04 | 1 | kHz | 2.5 | 0.23 | 2.345 | 0.23122
| ACVmeas0.01100.04 | ACVmeas | 10 | 100 | 0.01 | V | mV | 0.04 | 1 | kHz | 5.3174 | 0.2079 | 5.3167 | 0.2079
| ACVmeas0.1100.04 | ACVmeas | 10 | 100 | 0.1 | V | mV | 0.04 | 1 | kHz | 57.345 | 0.057562 | 57.345 | 0.057554
| ACVmeas1100.04 | ACVmeas | 10 | 100 | 1 | V | mV | 0.04 | 1 | kHz | 577.31 | 0.0060326 | 577.31 | 0.006029
| ACVmeas10100.04 | ACVmeas | 10 | 100 | 10 | V | mV | 0.04 | 1 | kHz | 5773.5 | 0.00060465 | 5773.5 | 0.0006032
| ACVmeas0.0001101 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 1 | 20 | kHz | 2.5 | 0.23 | 2.306 | 0.2317
| ACVmeas0.001101 | ACVmeas | 10 | 100 | 0.001 | V | mV | 1 | 20 | kHz | 2.5 | 0.23 | 2.344 | 0.23137
| ACVmeas0.01101 | ACVmeas | 10 | 100 | 0.01 | V | mV | 1 | 20 | kHz | 5.3157 | 0.20806 | 5.3147 | 0.20806
| ACVmeas0.1101 | ACVmeas | 10 | 100 | 0.1 | V | mV | 1 | 20 | kHz | 57.344 | 0.057626 | 57.344 | 0.057618
| ACVmeas1101 | ACVmeas | 10 | 100 | 1 | V | mV | 1 | 20 | kHz | 577.31 | 0.0060396 | 577.31 | 0.0060359
| ACVmeas10101 | ACVmeas | 10 | 100 | 10 | V | mV | 1 | 20 | kHz | 5773.5 | 0.00060535 | 5773.5 | 0.00060389
| ACVmeas0.00011020 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 20 | 50 | kHz | 2.8 | 0.4 | 2.3073 | 0.40484
| ACVmeas0.0011020 | ACVmeas | 10 | 100 | 0.001 | V | mV | 20 | 50 | kHz | 2.8 | 0.4 | 2.3352 | 0.4046
| ACVmeas0.011020 | ACVmeas | 10 | 100 | 0.01 | V | mV | 20 | 50 | kHz | 4.7448 | 0.38436 | 4.7424 | 0.38437
| ACVmeas0.11020 | ACVmeas | 10 | 100 | 0.1 | V | mV | 20 | 50 | kHz | 56.555 | 0.15314 | 56.553 | 0.15311
| ACVmeas11020 | ACVmeas | 10 | 100 | 1 | V | mV | 20 | 50 | kHz | 577.22 | 0.017224 | 577.22 | 0.017207
| ACVmeas101020 | ACVmeas | 10 | 100 | 10 | V | mV | 20 | 50 | kHz | 5773.5 | 0.0017297 | 5773.4 | 0.0017231
| ACVmeas0.00011050 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3108 | 1.386
| ACVmeas0.0011050 | ACVmeas | 10 | 100 | 0.001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.321 | 1.3859
| ACVmeas0.011050 | ACVmeas | 10 | 100 | 0.01 | V | mV | 50 | 100 | kHz | 3.2577 | 1.3914 | 3.4081 | 1.3762
| ACVmeas0.11050 | ACVmeas | 10 | 100 | 0.1 | V | mV | 50 | 100 | kHz | 49.704 | 1.0258 | 49.7 | 1.0258
| ACVmeas11050 | ACVmeas | 10 | 100 | 1 | V | mV | 50 | 100 | kHz | 575.73 | 0.18581 | 575.72 | 0.18578
| ACVmeas101050 | ACVmeas | 10 | 100 | 10 | V | mV | 50 | 100 | kHz | 5773.4 | 0.018864 | 5771.7 | 0.018852
| ACVmeas0.000110100 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 100 | 300 | kHz | 14 | 4.6 | 11.544 | 4.6193
| ACVmeas0.00110100 | ACVmeas | 10 | 100 | 0.001 | V | mV | 100 | 300 | kHz | 14 | 4.6 | 11.551 | 4.6192
| ACVmeas0.0110100 | ACVmeas | 10 | 100 | 0.01 | V | mV | 100 | 300 | kHz | 14 | 4.6 | 11.861 | 4.6165
| ACVmeas0.110100 | ACVmeas | 10 | 100 | 0.1 | V | mV | 100 | 300 | kHz | 37.743 | 4.3924 | 37.728 | 4.3925
| ACVmeas110100 | ACVmeas | 10 | 100 | 1 | V | mV | 100 | 300 | kHz | 561.75 | 1.8493 | 561.74 | 1.8493
| ACVmeas1010100 | ACVmeas | 10 | 100 | 10 | V | mV | 100 | 300 | kHz | 5771.7 | 0.21218 | 5755.3 | 0.21214
| ACVmeas0.000110300 | ACVmeas | 10 | 100 | 0.0001 | V | mV | 300 | 1000 | kHz | 44 | 17 | 11.591 | 17.321
| ACVmeas0.00110300 | ACVmeas | 10 | 100 | 0.001 | V | mV | 300 | 1000 | kHz | 44 | 17 | 11.591 | 17.321
| ACVmeas0.0110300 | ACVmeas | 10 | 100 | 0.01 | V | mV | 300 | 1000 | kHz | 44 | 17 | 11.701 | 17.32
| ACVmeas0.110300 | ACVmeas | 10 | 100 | 0.1 | V | mV | 300 | 1000 | kHz | 44 | 17.007 | 21.247 | 17.234
| ACVmeas110300 | ACVmeas | 10 | 100 | 1 | V | mV | 300 | 1000 | kHz | 469.5 | 13.673 | 469.49 | 13.673
| ACVmeas1010300 | ACVmeas | 10 | 100 | 10 | V | mV | 300 | 1000 | kHz | 5748.2 | 2.829 | 5639.8 | 2.829
| ACVmeas0.0011000.001 | ACVmeas | 100 | 1000 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 49 | 0.46 | 46.143 | 0.46254
| ACVmeas0.011000.001 | ACVmeas | 100 | 1000 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 49 | 0.46 | 46.336 | 0.46238
| ACVmeas0.11000.001 | ACVmeas | 100 | 1000 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 64.188 | 0.44777 | 64.17 | 0.44778
| ACVmeas11000.001 | ACVmeas | 100 | 1000 | 1 | V | mV | 0.001 | 0.04 | kHz | 564.18 | 0.20532 | 564.17 | 0.20531
| ACVmeas101000.001 | ACVmeas | 100 | 1000 | 10 | V | mV | 0.001 | 0.04 | kHz | 5771.9 | 0.024039 | 5771.9 | 0.02403
| ACVmeas0.0011000.04 | ACVmeas | 100 | 1000 | 0.001 | V | mV | 0.04 | 1 | kHz | 26 | 0.46 | 23.049 | 0.46257
| ACVmeas0.011000.04 | ACVmeas | 100 | 1000 | 0.01 | V | mV | 0.04 | 1 | kHz | 26 | 0.46 | 23.313 | 0.46234
| ACVmeas0.11000.04 | ACVmeas | 100 | 1000 | 0.1 | V | mV | 0.04 | 1 | kHz | 45.908 | 0.44314 | 45.889 | 0.44315
| ACVmeas11000.04 | ACVmeas | 100 | 1000 | 1 | V | mV | 0.04 | 1 | kHz | 562.29 | 0.19216 | 562.28 | 0.19215
| ACVmeas101000.04 | ACVmeas | 100 | 1000 | 10 | V | mV | 0.04 | 1 | kHz | 5771.8 | 0.022199 | 5771.7 | 0.02219
| ACVmeas0.0011001 | ACVmeas | 100 | 1000 | 0.001 | V | mV | 1 | 20 | kHz | 27 | 0.69 | 23.064 | 0.69336
| ACVmeas0.011001 | ACVmeas | 100 | 1000 | 0.01 | V | mV | 1 | 20 | kHz | 27 | 0.69 | 23.257 | 0.69319
| ACVmeas0.11001 | ACVmeas | 100 | 1000 | 0.1 | V | mV | 1 | 20 | kHz | 41.228 | 0.67753 | 41.196 | 0.67755
| ACVmeas11001 | ACVmeas | 100 | 1000 | 1 | V | mV | 1 | 20 | kHz | 547.46 | 0.37269 | 547.43 | 0.37268
| ACVmeas101001 | ACVmeas | 100 | 1000 | 10 | V | mV | 1 | 20 | kHz | 5769.5 | 0.048394 | 5769.5 | 0.048379
| ACVmeas0.00110020 | ACVmeas | 100 | 1000 | 0.001 | V | mV | 20 | 50 | kHz | 24 | 1.4 | 23.091 | 1.3863
| ACVmeas0.0110020 | ACVmeas | 100 | 1000 | 0.01 | V | mV | 20 | 50 | kHz | 24 | 1.4 | 23.203 | 1.3862
| ACVmeas0.110020 | ACVmeas | 100 | 1000 | 0.1 | V | mV | 20 | 50 | kHz | 32.582 | 1.3914 | 34.016 | 1.3766
| ACVmeas110020 | ACVmeas | 100 | 1000 | 1 | V | mV | 20 | 50 | kHz | 497.01 | 1.0261 | 496.97 | 1.0261
| ACVmeas1010020 | ACVmeas | 100 | 1000 | 10 | V | mV | 20 | 50 | kHz | 5757.2 | 0.18589 | 5757.2 | 0.18586
| ACVmeas0.00110050 | ACVmeas | 100 | 1000 | 0.001 | V | mV | 50 | 100 | kHz | 24 | 3.5 | 23.078 | 3.4644
| ACVmeas0.0110050 | ACVmeas | 100 | 1000 | 0.01 | V | mV | 50 | 100 | kHz | 24 | 3.5 | 23.166 | 3.4643
| ACVmeas0.110050 | ACVmeas | 100 | 1000 | 0.1 | V | mV | 50 | 100 | kHz | 24.058 | 3.4999 | 28.039 | 3.4599
| ACVmeas110050 | ACVmeas | 100 | 1000 | 1 | V | mV | 50 | 100 | kHz | 368.98 | 3.166 | 368.87 | 3.1661
| ACVmeas1010050 | ACVmeas | 100 | 1000 | 10 | V | mV | 50 | 100 | kHz | 5678.8 | 1.0664 | 5678.7 | 1.0664
| ACVgen0.000000100.001 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 3.6 | 340 | 3.5358 | 343.43
| ACVgen0.00000100.001 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 3.6 | 340 | 3.582 | 341.18
| ACVgen0.0000100.001 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 6.7662 | 228.75 | 6.7651 | 228.55
| ACVgen0.000100.001 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 57.843 | 31.346 | 57.843 | 31.206
| ACVgen0.00100.001 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 577.37 | 3.1912 | 577.37 | 3.1347
| ACVgen0.0100.001 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.6 | 0.33606 | 5773.6 | 0.31349
| ACVgen0.000000100.04 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 0.04 | 1 | kHz | 1.5 | 220 | 1.4547 | 220.78
| ACVgen0.00000100.04 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 0.04 | 1 | kHz | 1.5644 | 214.48 | 1.5632 | 214.41
| ACVgen0.0000100.04 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 0.04 | 1 | kHz | 5.9515 | 88.952 | 5.9503 | 88.67
| ACVgen0.000100.04 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 0.04 | 1 | kHz | 57.754 | 9.9554 | 57.753 | 9.8147
| ACVgen0.00100.04 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.36 | 1.039 | 577.36 | 0.9826
| ACVgen0.0100.04 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.6 | 0.12083 | 5773.6 | 0.098261
| ACVgen0.000000101 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 1 | 20 | kHz | 1.7 | 330 | 1.6224 | 325.82
| ACVgen0.00000101 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 1 | 20 | kHz | 1.7202 | 327.99 | 1.72 | 319.43
| ACVgen0.0000101 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 1 | 20 | kHz | 5.9913 | 157.22 | 5.9901 | 156.97
| ACVgen0.000101 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 1 | 20 | kHz | 57.758 | 18.545 | 57.757 | 18.405
| ACVgen0.00101 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 1 | 20 | kHz | 577.36 | 1.9005 | 577.36 | 1.8441
| ACVgen0.0101 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 1 | 20 | kHz | 5773.6 | 0.20698 | 5773.6 | 0.18441
| ACVgen0.0000001020 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 20 | 50 | kHz | 1.9 | 1100 | 1.6569 | 1123.5
| ACVgen0.000001020 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 20 | 50 | kHz | 1.9 | 1100.6 | 1.7487 | 1115.6
| ACVgen0.00001020 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 20 | 50 | kHz | 5.9597 | 816.89 | 5.9565 | 816.91
| ACVgen0.0001020 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 20 | 50 | kHz | 57.75 | 141.04 | 57.749 | 140.88
| ACVgen0.001020 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 20 | 50 | kHz | 577.36 | 14.328 | 577.36 | 14.263
| ACVgen0.01020 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 20 | 50 | kHz | 5773.6 | 1.4524 | 5773.6 | 1.4265
| ACVgen0.0000001050 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 50 | 100 | kHz | 2.1 | 5700 | 1.6579 | 5737.3
| ACVgen0.000001050 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 50 | 100 | kHz | 2.1 | 5700 | 1.7311 | 5730.3
| ACVgen0.00001050 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 50 | 100 | kHz | 5.6597 | 5365.6 | 5.6529 | 5366
| ACVgen0.0001050 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 50 | 100 | kHz | 57.531 | 2504.9 | 57.529 | 2504.3
| ACVgen0.001050 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 50 | 100 | kHz | 577.33 | 303.94 | 577.33 | 303.6
| ACVgen0.01050 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 50 | 100 | kHz | 5773.6 | 30.576 | 5773.5 | 30.438
| ACVgen0.00000010100 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.599 | 46160
| ACVgen0.0000010100 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.6212 | 46158
| ACVgen0.000010100 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 100 | 300 | kHz | 4.6664 | 45957 | 4.6449 | 45959
| ACVgen0.00010100 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 100 | 300 | kHz | 54.062 | 41372 | 54.047 | 41373
| ACVgen0.0010100 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 100 | 300 | kHz | 575.75 | 16509 | 575.75 | 16508
| ACVgen0.010100 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 100 | 300 | kHz | 5773.4 | 1882 | 5771.9 | 1881.5
| ACVgen0.00000010300 | ACVgen | 0 | 0.01 | 1e-07 | V | µV | 300 | 1000 | kHz | 7.8 | 46000 | 5.9731 | 46173
| ACVgen0.0000010300 | ACVgen | 0 | 0.01 | 1e-06 | V | µV | 300 | 1000 | kHz | 7.8 | 46000 | 5.993 | 46171
| ACVgen0.000010300 | ACVgen | 0 | 0.01 | 1e-05 | V | µV | 300 | 1000 | kHz | 7.8 | 46000 | 7.459 | 46028
| ACVgen0.00010300 | ACVgen | 0 | 0.01 | 0.0001 | V | µV | 300 | 1000 | kHz | 54.543 | 41672 | 54.533 | 41672
| ACVgen0.0010300 | ACVgen | 0 | 0.01 | 0.001 | V | µV | 300 | 1000 | kHz | 575.78 | 16726 | 575.78 | 16725
| ACVgen0.010300 | ACVgen | 0 | 0.01 | 0.01 | V | µV | 300 | 1000 | kHz | 5773.4 | 1910 | 5771.9 | 1909.4
| ACVgen0.00000010.010.001 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 4.7 | 83 | 4.6489 | 83.16
| ACVgen0.0000010.010.001 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 4.7 | 83 | 4.6808 | 82.968
| ACVgen0.000010.010.001 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 7.2686 | 69.256 | 7.2678 | 69.244
| ACVgen0.00010.010.001 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 57.864 | 13.107 | 57.864 | 13.092
| ACVgen0.0010.010.001 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 577.37 | 1.3346 | 577.37 | 1.3282
| ACVgen0.010.010.001 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.6 | 0.13538 | 5773.5 | 0.13284
| ACVgen0.00000010.010.04 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 0.04 | 1 | kHz | 2.5 | 83 | 2.3791 | 83.332
| ACVgen0.0000010.010.04 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 0.04 | 1 | kHz | 2.5 | 83 | 2.4341 | 82.937
| ACVgen0.000010.010.04 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 0.04 | 1 | kHz | 5.9897 | 61.816 | 5.9888 | 61.802
| ACVgen0.00010.010.04 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 0.04 | 1 | kHz | 57.726 | 9.9721 | 57.725 | 9.9566
| ACVgen0.0010.010.04 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.35 | 1.0112 | 577.35 | 1.0048
| ACVgen0.010.010.04 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.6 | 0.10304 | 5773.5 | 0.10049
| ACVgen0.00000010.011 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 1 | 20 | kHz | 2.8 | 160 | 2.4757 | 162.55
| ACVgen0.0000010.011 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 1 | 20 | kHz | 2.8 | 160 | 2.5195 | 162.2
| ACVgen0.000010.011 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 1 | 20 | kHz | 5.6916 | 139.1 | 5.6907 | 139.1
| ACVgen0.00010.011 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 1 | 20 | kHz | 57.568 | 31.313 | 57.568 | 31.299
| ACVgen0.0010.011 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 1 | 20 | kHz | 577.34 | 3.2196 | 577.34 | 3.2132
| ACVgen0.010.011 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.6 | 0.32395 | 5773.5 | 0.32141
| ACVgen0.00000010.0120 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 20 | 50 | kHz | 2.5 | 350 | 2.4623 | 347.56
| ACVgen0.0000010.0120 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 20 | 50 | kHz | 2.5 | 350 | 2.4922 | 347.31
| ACVgen0.000010.0120 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 20 | 50 | kHz | 5.0203 | 326.46 | 5.0169 | 326.47
| ACVgen0.00010.0120 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 20 | 50 | kHz | 56.858 | 118.4 | 56.856 | 118.36
| ACVgen0.0010.0120 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 20 | 50 | kHz | 577.26 | 12.999 | 577.26 | 12.976
| ACVgen0.010.0120 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 20 | 50 | kHz | 5773.5 | 1.3082 | 5773.4 | 1.299
| ACVgen0.00000010.0150 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 50 | 100 | kHz | 3 | 920 | 2.428 | 924.97
| ACVgen0.0000010.0150 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 50 | 100 | kHz | 3 | 920 | 2.4437 | 924.83
| ACVgen0.000010.0150 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 50 | 100 | kHz | 3.9159 | 911.89 | 3.9076 | 911.93
| ACVgen0.00010.0150 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 50 | 100 | kHz | 53.111 | 580.08 | 53.105 | 580
| ACVgen0.0010.0150 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 50 | 100 | kHz | 576.63 | 84.904 | 576.63 | 84.816
| ACVgen0.010.0150 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 50 | 100 | kHz | 5773.5 | 8.5746 | 5772.7 | 8.5388
| ACVgen0.00000010.01100 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.632 | 3465.2
| ACVgen0.0000010.01100 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.634 | 3465.2
| ACVgen0.000010.01100 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 100 | 300 | kHz | 12 | 3500 | 12.023 | 3461.8
| ACVgen0.00010.01100 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 100 | 300 | kHz | 41.936 | 3208.6 | 41.909 | 3208.7
| ACVgen0.0010.01100 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 100 | 300 | kHz | 568.08 | 1113.7 | 568.07 | 1113.5
| ACVgen0.010.01100 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 100 | 300 | kHz | 5772.5 | 121.36 | 5762.6 | 121.25
| ACVgen0.00000010.01300 | ACVgen | 0.01 | 0.1 | 1e-07 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.641 | 11557
| ACVgen0.0000010.01300 | ACVgen | 0.01 | 0.1 | 1e-06 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.641 | 11557
| ACVgen0.000010.01300 | ACVgen | 0.01 | 0.1 | 1e-05 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.768 | 11556
| ACVgen0.00010.01300 | ACVgen | 0.01 | 0.1 | 0.0001 | V | µV | 300 | 1000 | kHz | 20.585 | 11914 | 25.368 | 11434
| ACVgen0.0010.01300 | ACVgen | 0.01 | 0.1 | 0.001 | V | µV | 300 | 1000 | kHz | 512.23 | 7901.3 | 512.19 | 7901.3
| ACVgen0.010.01300 | ACVgen | 0.01 | 0.1 | 0.01 | V | µV | 300 | 1000 | kHz | 5762.1 | 1282.8 | 5695.9 | 1282.5
| ACVgen0.0000010.10.001 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 81.758 | 81.758
| ACVgen0.000010.10.001 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 81.565 | 81.565
| ACVgen0.00010.10.001 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 72.505 | 67.853 | 67.847 | 67.847
| ACVgen0.0010.10.001 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 578.64 | 12.732 | 12.722 | 12.722
| ACVgen0.010.10.001 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.7 | 1.2942 | 1.2903 | 1.2903
| ACVgen0.10.10.001 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 0.001 | 0.04 | kHz | 57736 | 0.13062 | 0.12905 | 0.12905
| ACVgen0.0000010.10.04 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 23.098 | 81.925
| ACVgen0.000010.10.04 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 23.663 | 81.518
| ACVgen0.00010.10.04 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 0.04 | 1 | kHz | 59.658 | 60.205 | 59.647 | 60.199
| ACVgen0.0010.10.04 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.25 | 9.5953 | 577.24 | 9.5857
| ACVgen0.010.10.04 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.5 | 0.97099 | 5773.5 | 0.96705
| ACVgen0.10.10.04 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 0.04 | 1 | kHz | 57736 | 0.098288 | 57735 | 0.096714
| ACVgen0.0000010.11 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.16 | 162.3
| ACVgen0.000010.11 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.609 | 161.94
| ACVgen0.00010.11 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 1 | 20 | kHz | 56.1 | 138.17 | 56.076 | 138.16
| ACVgen0.0010.11 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 1 | 20 | kHz | 575.63 | 30.826 | 575.62 | 30.793
| ACVgen0.010.11 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.4 | 3.1739 | 5773.4 | 3.1595
| ACVgen0.10.11 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 1 | 20 | kHz | 57736 | 0.32182 | 57735 | 0.31603
| ACVgen0.0000010.120 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.232 | 346.82
| ACVgen0.000010.120 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.54 | 346.56
| ACVgen0.00010.120 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 20 | 50 | kHz | 49.33 | 325.25 | 49.249 | 325.28
| ACVgen0.0010.120 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 20 | 50 | kHz | 568.57 | 117.31 | 568.53 | 117.24
| ACVgen0.010.120 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 20 | 50 | kHz | 5772.6 | 12.875 | 5772.6 | 12.841
| ACVgen0.10.120 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 20 | 50 | kHz | 57735 | 1.2992 | 57734 | 1.2854
| ACVgen0.0000010.150 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 50 | 100 | kHz | 28 | 920 | 23.291 | 923.8
| ACVgen0.000010.150 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 50 | 100 | kHz | 28 | 920 | 23.448 | 923.66
| ACVgen0.00010.150 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 50 | 100 | kHz | 38.364 | 910.53 | 38.219 | 910.63
| ACVgen0.0010.150 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 50 | 100 | kHz | 531.13 | 578.22 | 531.01 | 578.19
| ACVgen0.010.150 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 50 | 100 | kHz | 5766.3 | 84.519 | 5766.3 | 84.452
| ACVgen0.10.150 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 50 | 100 | kHz | 57735 | 8.5295 | 57727 | 8.5019
| ACVgen0.0000010.1100 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 100 | 300 | kHz | 120 | 3500 | 115.61 | 3464.2
| ACVgen0.000010.1100 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 100 | 300 | kHz | 120 | 3500 | 115.68 | 3464.1
| ACVgen0.00010.1100 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 100 | 300 | kHz | 120 | 3500 | 119.56 | 3460.7
| ACVgen0.0010.1100 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 100 | 300 | kHz | 419.06 | 3207.1 | 418.74 | 3207.3
| ACVgen0.010.1100 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 100 | 300 | kHz | 5680.9 | 1112.6 | 5680.8 | 1112.5
| ACVgen0.10.1100 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 100 | 300 | kHz | 57725 | 121.19 | 57626 | 121.14
| ACVgen0.0000010.1300 | ACVgen | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 116.28 | 11549
| ACVgen0.000010.1300 | ACVgen | 0.1 | 1 | 1e-05 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 116.29 | 11549
| ACVgen0.00010.1300 | ACVgen | 0.1 | 1 | 0.0001 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 117.69 | 11548
| ACVgen0.0010.1300 | ACVgen | 0.1 | 1 | 0.001 | V | µV | 300 | 1000 | kHz | 204.93 | 11915 | 253.55 | 11426
| ACVgen0.010.1300 | ACVgen | 0.1 | 1 | 0.01 | V | µV | 300 | 1000 | kHz | 5122.8 | 7893.3 | 5122.5 | 7893.4
| ACVgen0.10.1300 | ACVgen | 0.1 | 1 | 0.1 | V | µV | 300 | 1000 | kHz | 57622 | 1280.9 | 56960 | 1280.8
| ACVgen0.0000110.001 | ACVgen | 1 | 10 | 1e-05 | V | mV | 0.001 | 0.04 | kHz | 0.47 | 0.083 | 0.45986 | 0.083101
| ACVgen0.000110.001 | ACVgen | 1 | 10 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 0.47 | 0.083 | 0.46309 | 0.082906
| ACVgen0.00110.001 | ACVgen | 1 | 10 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 0.72377 | 0.069102 | 0.72347 | 0.069062
| ACVgen0.0110.001 | ACVgen | 1 | 10 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 5.7861 | 0.013062 | 5.786 | 0.013008
| ACVgen0.110.001 | ACVgen | 1 | 10 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.737 | 0.0013417 | 57.737 | 0.0013196
| ACVgen110.001 | ACVgen | 1 | 10 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.36 | 0.00014082 | 577.35 | 0.00013198
| ACVgen0.0000110.04 | ACVgen | 1 | 10 | 1e-05 | V | mV | 0.04 | 1 | kHz | 0.24 | 0.082 | 0.23105 | 0.081983
| ACVgen0.000110.04 | ACVgen | 1 | 10 | 0.0001 | V | mV | 0.04 | 1 | kHz | 0.24 | 0.082 | 0.2359 | 0.081575
| ACVgen0.00110.04 | ACVgen | 1 | 10 | 0.001 | V | mV | 0.04 | 1 | kHz | 0.59648 | 0.060264 | 0.59612 | 0.060218
| ACVgen0.0110.04 | ACVgen | 1 | 10 | 0.01 | V | mV | 0.04 | 1 | kHz | 5.7725 | 0.0096404 | 5.7724 | 0.0095863
| ACVgen0.110.04 | ACVgen | 1 | 10 | 0.1 | V | mV | 0.04 | 1 | kHz | 57.735 | 0.0009892 | 57.735 | 0.0009671
| ACVgen110.04 | ACVgen | 1 | 10 | 1 | V | mV | 0.04 | 1 | kHz | 577.36 | 0.00010556 | 577.35 | 9.6719e-05
| ACVgen0.0000111 | ACVgen | 1 | 10 | 1e-05 | V | mV | 1 | 20 | kHz | 0.26 | 0.16 | 0.23135 | 0.16229
| ACVgen0.000111 | ACVgen | 1 | 10 | 0.0001 | V | mV | 1 | 20 | kHz | 0.26 | 0.16 | 0.23526 | 0.16193
| ACVgen0.00111 | ACVgen | 1 | 10 | 0.001 | V | mV | 1 | 20 | kHz | 0.5617 | 0.13811 | 0.56028 | 0.13812
| ACVgen0.0111 | ACVgen | 1 | 10 | 0.01 | V | mV | 1 | 20 | kHz | 5.7567 | 0.030878 | 5.7561 | 0.030768
| ACVgen0.111 | ACVgen | 1 | 10 | 0.1 | V | mV | 1 | 20 | kHz | 57.734 | 0.0032047 | 57.734 | 0.0031568
| ACVgen111 | ACVgen | 1 | 10 | 1 | V | mV | 1 | 20 | kHz | 577.36 | 0.00033493 | 577.35 | 0.00031576
| ACVgen0.00001120 | ACVgen | 1 | 10 | 1e-05 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.23064 | 0.34716
| ACVgen0.0001120 | ACVgen | 1 | 10 | 0.0001 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.23377 | 0.34689
| ACVgen0.001120 | ACVgen | 1 | 10 | 0.001 | V | mV | 20 | 50 | kHz | 0.49453 | 0.32537 | 0.49127 | 0.32557
| ACVgen0.01120 | ACVgen | 1 | 10 | 0.01 | V | mV | 20 | 50 | kHz | 5.6867 | 0.11746 | 5.685 | 0.11736
| ACVgen0.1120 | ACVgen | 1 | 10 | 0.1 | V | mV | 20 | 50 | kHz | 57.726 | 0.012918 | 57.726 | 0.012854
| ACVgen1120 | ACVgen | 1 | 10 | 1 | V | mV | 20 | 50 | kHz | 577.35 | 0.0013125 | 577.34 | 0.0012868
| ACVgen0.00001150 | ACVgen | 1 | 10 | 1e-05 | V | mV | 50 | 100 | kHz | 0.28 | 0.92 | 0.23096 | 0.92404
| ACVgen0.0001150 | ACVgen | 1 | 10 | 0.0001 | V | mV | 50 | 100 | kHz | 0.28 | 0.92 | 0.23253 | 0.9239
| ACVgen0.001150 | ACVgen | 1 | 10 | 0.001 | V | mV | 50 | 100 | kHz | 0.38426 | 0.91055 | 0.38037 | 0.91086
| ACVgen0.01150 | ACVgen | 1 | 10 | 0.01 | V | mV | 50 | 100 | kHz | 5.3129 | 0.57823 | 5.3097 | 0.57827
| ACVgen0.1150 | ACVgen | 1 | 10 | 0.1 | V | mV | 50 | 100 | kHz | 57.664 | 0.084554 | 57.663 | 0.084464
| ACVgen1150 | ACVgen | 1 | 10 | 1 | V | mV | 50 | 100 | kHz | 577.35 | 0.0085405 | 577.27 | 0.0085031
| ACVgen0.000011100 | ACVgen | 1 | 10 | 1e-05 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1622 | 3.4645
| ACVgen0.00011100 | ACVgen | 1 | 10 | 0.0001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1627 | 3.4644
| ACVgen0.0011100 | ACVgen | 1 | 10 | 0.001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.194 | 3.461
| ACVgen0.011100 | ACVgen | 1 | 10 | 0.01 | V | mV | 100 | 300 | kHz | 4.1916 | 3.2072 | 4.1867 | 3.2075
| ACVgen0.11100 | ACVgen | 1 | 10 | 0.1 | V | mV | 100 | 300 | kHz | 56.81 | 1.1127 | 56.808 | 1.1126
| ACVgen11100 | ACVgen | 1 | 10 | 1 | V | mV | 100 | 300 | kHz | 577.25 | 0.12122 | 576.26 | 0.12115
| ACVgen0.000011300 | ACVgen | 1 | 10 | 1e-05 | V | mV | 300 | 1000 | kHz | 0.72 | 12 | 1.1976 | 11.551
| ACVgen0.00011300 | ACVgen | 1 | 10 | 0.0001 | V | mV | 300 | 1000 | kHz | 0.72 | 12 | 1.1977 | 11.551
| ACVgen0.0011300 | ACVgen | 1 | 10 | 0.001 | V | mV | 300 | 1000 | kHz | 0.72 | 11.999 | 1.2052 | 11.55
| ACVgen0.011300 | ACVgen | 1 | 10 | 0.01 | V | mV | 300 | 1000 | kHz | 2.098 | 11.861 | 2.5267 | 11.428
| ACVgen0.11300 | ACVgen | 1 | 10 | 0.1 | V | mV | 300 | 1000 | kHz | 51.227 | 7.8943 | 51.222 | 7.8945
| ACVgen11300 | ACVgen | 1 | 10 | 1 | V | mV | 300 | 1000 | kHz | 576.22 | 1.2811 | 569.6 | 1.281
| ACVgen0.0001100.001 | ACVgen | 10 | 20 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 4.7 | 0.23 | 4.6222 | 0.23257
| ACVgen0.001100.001 | ACVgen | 10 | 20 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 4.7 | 0.23 | 4.6516 | 0.23197
| ACVgen0.01100.001 | ACVgen | 10 | 20 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 7.1323 | 0.189 | 7.1324 | 0.18891
| ACVgen0.1100.001 | ACVgen | 10 | 20 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.827 | 0.032367 | 57.827 | 0.032286
| ACVgen1100.001 | ACVgen | 10 | 20 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.36 | 0.0032932 | 577.36 | 0.0032605
| ACVgen10100.001 | ACVgen | 10 | 20 | 10 | V | mV | 0.001 | 0.04 | kHz | 5773.6 | 0.00033916 | 5773.5 | 0.00032608
| ACVgen0.0001100.04 | ACVgen | 10 | 20 | 0.0001 | V | mV | 0.04 | 1 | kHz | 2.4 | 0.23 | 2.307 | 0.23147
| ACVgen0.001100.04 | ACVgen | 10 | 20 | 0.001 | V | mV | 0.04 | 1 | kHz | 2.4 | 0.23 | 2.3543 | 0.23029
| ACVgen0.01100.04 | ACVgen | 10 | 20 | 0.01 | V | mV | 0.04 | 1 | kHz | 5.7658 | 0.16305 | 5.766 | 0.16293
| ACVgen0.1100.04 | ACVgen | 10 | 20 | 0.1 | V | mV | 0.04 | 1 | kHz | 57.689 | 0.02313 | 57.69 | 0.023049
| ACVgen1100.04 | ACVgen | 10 | 20 | 1 | V | mV | 0.04 | 1 | kHz | 577.35 | 0.0023495 | 577.35 | 0.0023168
| ACVgen10100.04 | ACVgen | 10 | 20 | 10 | V | mV | 0.04 | 1 | kHz | 5773.6 | 0.00024477 | 5773.5 | 0.00023169
| ACVgen0.0001101 | ACVgen | 10 | 20 | 0.0001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.3059 | 0.23157
| ACVgen0.001101 | ACVgen | 10 | 20 | 0.001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.3533 | 0.23039
| ACVgen0.01101 | ACVgen | 10 | 20 | 0.01 | V | mV | 1 | 20 | kHz | 5.7651 | 0.16312 | 5.7652 | 0.16301
| ACVgen0.1101 | ACVgen | 10 | 20 | 0.1 | V | mV | 1 | 20 | kHz | 57.689 | 0.023142 | 57.69 | 0.023061
| ACVgen1101 | ACVgen | 10 | 20 | 1 | V | mV | 1 | 20 | kHz | 577.35 | 0.0023507 | 577.35 | 0.002318
| ACVgen10101 | ACVgen | 10 | 20 | 10 | V | mV | 1 | 20 | kHz | 5773.6 | 0.00024489 | 5773.5 | 0.00023181
| ACVgen0.00011020 | ACVgen | 10 | 20 | 0.0001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.3031 | 0.40507
| ACVgen0.0011020 | ACVgen | 10 | 20 | 0.001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.3383 | 0.40406
| ACVgen0.011020 | ACVgen | 10 | 20 | 0.01 | V | mV | 20 | 50 | kHz | 5.2718 | 0.33148 | 5.2706 | 0.33138
| ACVgen0.11020 | ACVgen | 10 | 20 | 0.1 | V | mV | 20 | 50 | kHz | 57.503 | 0.058262 | 57.503 | 0.058138
| ACVgen11020 | ACVgen | 10 | 20 | 1 | V | mV | 20 | 50 | kHz | 577.33 | 0.0059286 | 577.33 | 0.0058775
| ACVgen101020 | ACVgen | 10 | 20 | 10 | V | mV | 20 | 50 | kHz | 5773.6 | 0.00060825 | 5773.5 | 0.00058782
| ACVgen0.00011050 | ACVgen | 10 | 20 | 0.0001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3076 | 1.386
| ACVgen0.0011050 | ACVgen | 10 | 20 | 0.001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3231 | 1.3855
| ACVgen0.011050 | ACVgen | 10 | 20 | 0.01 | V | mV | 50 | 100 | kHz | 3.7587 | 1.341 | 3.7575 | 1.341
| ACVgen0.11050 | ACVgen | 10 | 20 | 0.1 | V | mV | 50 | 100 | kHz | 54.836 | 0.51218 | 54.836 | 0.51207
| ACVgen11050 | ACVgen | 10 | 20 | 1 | V | mV | 50 | 100 | kHz | 577.03 | 0.055451 | 577.03 | 0.0554
| ACVgen101050 | ACVgen | 10 | 20 | 10 | V | mV | 50 | 100 | kHz | 5773.5 | 0.0055652 | 5773 | 0.0055448
| ACVgen0.000110100 | ACVgen | 10 | 20 | 0.0001 | V | mV | 100 | 300 | kHz | 12 | 4.6 | 11.547 | 4.619
| ACVgen0.00110100 | ACVgen | 10 | 20 | 0.001 | V | mV | 100 | 300 | kHz | 12 | 4.6 | 11.552 | 4.6189
| ACVgen0.0110100 | ACVgen | 10 | 20 | 0.01 | V | mV | 100 | 300 | kHz | 12 | 4.6063 | 11.963 | 4.6063
| ACVgen0.110100 | ACVgen | 10 | 20 | 0.1 | V | mV | 100 | 300 | kHz | 44.416 | 3.7237 | 44.416 | 3.7236
| ACVgen110100 | ACVgen | 10 | 20 | 1 | V | mV | 100 | 300 | kHz | 573.84 | 0.64 | 573.84 | 0.63995
| ACVgen1010100 | ACVgen | 10 | 20 | 10 | V | mV | 100 | 300 | kHz | 5773.2 | 0.064683 | 5767.4 | 0.064662
| ACVgen0.000110300 | ACVgen | 10 | 20 | 0.0001 | V | mV | 300 | 1000 | kHz | 18 | 17 | 11.566 | 17.321
| ACVgen0.00110300 | ACVgen | 10 | 20 | 0.001 | V | mV | 300 | 1000 | kHz | 18 | 17 | 11.567 | 17.321
| ACVgen0.0110300 | ACVgen | 10 | 20 | 0.01 | V | mV | 300 | 1000 | kHz | 18 | 17.003 | 11.709 | 17.316
| ACVgen0.110300 | ACVgen | 10 | 20 | 0.1 | V | mV | 300 | 1000 | kHz | 24.569 | 16.902 | 24.568 | 16.902
| ACVgen110300 | ACVgen | 10 | 20 | 1 | V | mV | 300 | 1000 | kHz | 533.07 | 7.3129 | 533.07 | 7.3129
| ACVgen1010300 | ACVgen | 10 | 20 | 10 | V | mV | 300 | 1000 | kHz | 5768.4 | 0.81317 | 5703.4 | 0.81315
| ACVgen0.0001200.001 | ACVgen | 20 | 100 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 4.8 | 0.23 | 4.614 | 0.23155
| ACVgen0.001200.001 | ACVgen | 20 | 100 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 4.8 | 0.23 | 4.6321 | 0.23141
| ACVgen0.01200.001 | ACVgen | 20 | 100 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 6.5308 | 0.21829 | 6.5301 | 0.21829
| ACVgen0.1200.001 | ACVgen | 20 | 100 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.072 | 0.069939 | 57.072 | 0.06993
| ACVgen1200.001 | ACVgen | 20 | 100 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.28 | 0.0074201 | 577.28 | 0.007416
| ACVgen10200.001 | ACVgen | 20 | 100 | 10 | V | mV | 0.001 | 0.04 | kHz | 5773.5 | 0.00074371 | 5773.4 | 0.00074208
| ACVgen0.0001200.04 | ACVgen | 20 | 100 | 0.0001 | V | mV | 0.04 | 1 | kHz | 2.5 | 0.23 | 2.3037 | 0.23163
| ACVgen0.001200.04 | ACVgen | 20 | 100 | 0.001 | V | mV | 0.04 | 1 | kHz | 2.5 | 0.23 | 2.3278 | 0.23141
| ACVgen0.01200.04 | ACVgen | 20 | 100 | 0.01 | V | mV | 0.04 | 1 | kHz | 4.7489 | 0.21359 | 4.748 | 0.21359
| ACVgen0.1200.04 | ACVgen | 20 | 100 | 0.1 | V | mV | 0.04 | 1 | kHz | 56.913 | 0.061886 | 56.912 | 0.061877
| ACVgen1200.04 | ACVgen | 20 | 100 | 1 | V | mV | 0.04 | 1 | kHz | 577.27 | 0.0064981 | 577.27 | 0.006494
| ACVgen10200.04 | ACVgen | 20 | 100 | 10 | V | mV | 0.04 | 1 | kHz | 5773.5 | 0.00065137 | 5773.4 | 0.00064973
| ACVgen0.0001201 | ACVgen | 20 | 100 | 0.0001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8696 | 0.20605
| ACVgen0.001201 | ACVgen | 20 | 100 | 0.001 | V | mV | 1 | 20 | kHz | 4.9 | 0.20607 | 4.8953 | 0.20586
| ACVgen0.01201 | ACVgen | 20 | 100 | 0.01 | V | mV | 1 | 20 | kHz | 7.1206 | 0.19001 | 7.1197 | 0.19001
| ACVgen0.1201 | ACVgen | 20 | 100 | 0.1 | V | mV | 1 | 20 | kHz | 57.6 | 0.055073 | 57.599 | 0.055065
| ACVgen1201 | ACVgen | 20 | 100 | 1 | V | mV | 1 | 20 | kHz | 577.34 | 0.005783 | 577.34 | 0.0057794
| ACVgen10201 | ACVgen | 20 | 100 | 10 | V | mV | 1 | 20 | kHz | 5773.6 | 0.00057969 | 5773.5 | 0.00057823
| ACVgen0.00012020 | ACVgen | 20 | 100 | 0.0001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36 | 6.7954 | 0.35996
| ACVgen0.0012020 | ACVgen | 20 | 100 | 0.001 | V | mV | 20 | 50 | kHz | 6.8127 | 0.35987 | 6.8127 | 0.35983
| ACVgen0.012020 | ACVgen | 20 | 100 | 0.01 | V | mV | 20 | 50 | kHz | 8.4158 | 0.34765 | 8.4142 | 0.34765
| ACVgen0.12020 | ACVgen | 20 | 100 | 0.1 | V | mV | 20 | 50 | kHz | 57.199 | 0.1467 | 57.197 | 0.14667
| ACVgen12020 | ACVgen | 20 | 100 | 1 | V | mV | 20 | 50 | kHz | 577.28 | 0.016573 | 577.28 | 0.016556
| ACVgen102020 | ACVgen | 20 | 100 | 10 | V | mV | 20 | 50 | kHz | 5773.5 | 0.0016646 | 5773.4 | 0.001658
| ACVgen0.00012050 | ACVgen | 20 | 100 | 0.0001 | V | mV | 50 | 100 | kHz | 21 | 1.2 | 17.702 | 1.2321
| ACVgen0.0012050 | ACVgen | 20 | 100 | 0.001 | V | mV | 50 | 100 | kHz | 21 | 1.2 | 17.712 | 1.232
| ACVgen0.012050 | ACVgen | 20 | 100 | 0.01 | V | mV | 50 | 100 | kHz | 21 | 1.2003 | 18.301 | 1.2273
| ACVgen0.12050 | ACVgen | 20 | 100 | 0.1 | V | mV | 50 | 100 | kHz | 55.39 | 0.96894 | 55.386 | 0.96894
| ACVgen12050 | ACVgen | 20 | 100 | 1 | V | mV | 50 | 100 | kHz | 576.34 | 0.17966 | 576.34 | 0.17963
| ACVgen102050 | ACVgen | 20 | 100 | 10 | V | mV | 50 | 100 | kHz | 5773.4 | 0.018248 | 5771.8 | 0.018236
| ACVgen0.0011000.001 | ACVgen | 100 | 330 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 47 | 0.46 | 46.153 | 0.46241
| ACVgen0.011000.001 | ACVgen | 100 | 330 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 47 | 0.46016 | 46.374 | 0.46199
| ACVgen0.11000.001 | ACVgen | 100 | 330 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 66.343 | 0.42619 | 66.334 | 0.42615
| ACVgen11000.001 | ACVgen | 100 | 330 | 1 | V | mV | 0.001 | 0.04 | kHz | 573.44 | 0.11271 | 573.44 | 0.11263
| ACVgen101000.001 | ACVgen | 100 | 330 | 10 | V | mV | 0.001 | 0.04 | kHz | 5773.1 | 0.01169 | 5773.1 | 0.011655
| ACVgen0.0011000.04 | ACVgen | 100 | 330 | 0.001 | V | mV | 0.04 | 1 | kHz | 24 | 0.46 | 23.06 | 0.46246
| ACVgen0.011000.04 | ACVgen | 100 | 330 | 0.01 | V | mV | 0.04 | 1 | kHz | 24 | 0.46 | 23.36 | 0.46184
| ACVgen0.11000.04 | ACVgen | 100 | 330 | 0.1 | V | mV | 0.04 | 1 | kHz | 49.035 | 0.41183 | 49.022 | 0.4118
| ACVgen11000.04 | ACVgen | 100 | 330 | 1 | V | mV | 0.04 | 1 | kHz | 571.94 | 0.09569 | 571.94 | 0.095607
| ACVgen101000.04 | ACVgen | 100 | 330 | 10 | V | mV | 0.04 | 1 | kHz | 5773 | 0.009844 | 5773 | 0.0098088
| ACVgen0.0011001 | ACVgen | 100 | 330 | 0.001 | V | mV | 1 | 20 | kHz | 25 | 0.69 | 23.07 | 0.69331
| ACVgen0.011001 | ACVgen | 100 | 330 | 0.01 | V | mV | 1 | 20 | kHz | 25 | 0.69 | 23.295 | 0.69282
| ACVgen0.11001 | ACVgen | 100 | 330 | 0.1 | V | mV | 1 | 20 | kHz | 44.008 | 0.6497 | 43.978 | 0.64974
| ACVgen11001 | ACVgen | 100 | 330 | 1 | V | mV | 1 | 20 | kHz | 565.07 | 0.19654 | 565.05 | 0.19651
| ACVgen101001 | ACVgen | 100 | 330 | 10 | V | mV | 1 | 20 | kHz | 5772.2 | 0.020677 | 5772.2 | 0.020659
| ACVgen0.00110020 | ACVgen | 100 | 330 | 0.001 | V | mV | 20 | 50 | kHz | 24 | 1.4 | 23.121 | 1.386
| ACVgen0.0110020 | ACVgen | 100 | 330 | 0.01 | V | mV | 20 | 50 | kHz | 24 | 1.4 | 23.242 | 1.3857
| ACVgen0.110020 | ACVgen | 100 | 330 | 0.1 | V | mV | 20 | 50 | kHz | 35.083 | 1.3664 | 35.934 | 1.3575
| ACVgen110020 | ACVgen | 100 | 330 | 1 | V | mV | 20 | 50 | kHz | 533.73 | 0.65893 | 533.68 | 0.65891
| ACVgen1010020 | ACVgen | 100 | 330 | 10 | V | mV | 20 | 50 | kHz | 5768.1 | 0.076956 | 5768.1 | 0.076926
| ACVgen0.00110050 | ACVgen | 100 | 330 | 0.001 | V | mV | 50 | 100 | kHz | 24 | 3.5 | 23.1 | 3.4643
| ACVgen0.0110050 | ACVgen | 100 | 330 | 0.01 | V | mV | 50 | 100 | kHz | 24 | 3.5 | 23.144 | 3.4642
| ACVgen0.110050 | ACVgen | 100 | 330 | 0.1 | V | mV | 50 | 100 | kHz | 24.069 | 3.4998 | 28.907 | 3.451
| ACVgen110050 | ACVgen | 100 | 330 | 1 | V | mV | 50 | 100 | kHz | 417.82 | 2.6777 | 417.69 | 2.6779
| ACVgen1010050 | ACVgen | 100 | 330 | 10 | V | mV | 50 | 100 | kHz | 5739.8 | 0.45576 | 5739.8 | 0.45571
| ACVgen0.0013300.001 | ACVgen | 330 | 1000 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 49 | 0.46 | 46.143 | 0.46254
| ACVgen0.013300.001 | ACVgen | 330 | 1000 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 49 | 0.46 | 46.336 | 0.46238
| ACVgen0.13300.001 | ACVgen | 330 | 1000 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 64.188 | 0.44777 | 64.17 | 0.44778
| ACVgen13300.001 | ACVgen | 330 | 1000 | 1 | V | mV | 0.001 | 0.04 | kHz | 564.18 | 0.20532 | 564.17 | 0.20531
| ACVgen103300.001 | ACVgen | 330 | 1000 | 10 | V | mV | 0.001 | 0.04 | kHz | 5771.9 | 0.024039 | 5771.9 | 0.02403
| ACVgen0.0013300.04 | ACVgen | 330 | 1000 | 0.001 | V | mV | 0.04 | 1 | kHz | 26 | 0.46 | 23.049 | 0.46257
| ACVgen0.013300.04 | ACVgen | 330 | 1000 | 0.01 | V | mV | 0.04 | 1 | kHz | 26 | 0.46 | 23.313 | 0.46234
| ACVgen0.13300.04 | ACVgen | 330 | 1000 | 0.1 | V | mV | 0.04 | 1 | kHz | 45.908 | 0.44314 | 45.889 | 0.44315
| ACVgen13300.04 | ACVgen | 330 | 1000 | 1 | V | mV | 0.04 | 1 | kHz | 562.29 | 0.19216 | 562.28 | 0.19215
| ACVgen103300.04 | ACVgen | 330 | 1000 | 10 | V | mV | 0.04 | 1 | kHz | 5771.8 | 0.022199 | 5771.7 | 0.02219
| ACVgen0.0013301 | ACVgen | 330 | 1000 | 0.001 | V | mV | 1 | 10 | kHz | 27 | 0.69 | 23.064 | 0.69336
| ACVgen0.013301 | ACVgen | 330 | 1000 | 0.01 | V | mV | 1 | 10 | kHz | 27 | 0.69 | 23.258 | 0.69319
| ACVgen0.13301 | ACVgen | 330 | 1000 | 0.1 | V | mV | 1 | 10 | kHz | 41.228 | 0.67753 | 41.197 | 0.67755
| ACVgen13301 | ACVgen | 330 | 1000 | 1 | V | mV | 1 | 10 | kHz | 547.46 | 0.37269 | 547.43 | 0.37268
| ACVgen103301 | ACVgen | 330 | 1000 | 10 | V | mV | 1 | 10 | kHz | 5769.5 | 0.048394 | 5769.5 | 0.048379
| ACA0.00000000100.001 | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0.001 | 0.02 | kHz | 0.077 | 4200 | 0.076633 | 4198.9
| ACA0.0000000100.001 | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0.001 | 0.02 | kHz | 0.077 | 4201.6 | 0.076839 | 4197.2
| ACA0.000000100.001 | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 0.095301 | 4045.8 | 0.09529 | 4045.8
| ACA0.00000100.001 | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 0.58119 | 1803.4 | 0.58118 | 1803.2
| ACA0.0000100.001 | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 5.7739 | 209.71 | 5.7739 | 209.58
| ACA0.00000000100.02 | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0.02 | 0.045 | kHz | 0.038 | 1700 | 0.034648 | 1732
| ACA0.0000000100.02 | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 0.038 | 1700 | 0.03499 | 1729.4
| ACA0.000000100.02 | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 0.062352 | 1534 | 0.062344 | 1533.8
| ACA0.00000100.02 | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 0.57592 | 377.51 | 0.57592 | 377.19
| ACA0.0000100.02 | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.7734 | 39.132 | 5.7734 | 38.993
| ACA0.00000000100.045 | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0.045 | 5 | kHz | 0.036 | 690 | 0.034903 | 694.65
| ACA0.0000000100.045 | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0.045 | 5 | kHz | 0.036 | 690 | 0.035323 | 692.06
| ACA0.000000100.045 | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0.045 | 5 | kHz | 0.065989 | 533.4 | 0.065979 | 533.04
| ACA0.00000100.045 | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0.045 | 5 | kHz | 0.578 | 87.567 | 0.578 | 87.225
| ACA0.0000100.045 | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0.045 | 5 | kHz | 5.7736 | 8.9693 | 5.7736 | 8.83
| ACA0.0000000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.001 | 0.02 | kHz | 0.25 | 4600 | 0.23097 | 4618.8
| ACA0.000000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.001 | 0.02 | kHz | 0.25003 | 4600.1 | 0.2336 | 4616.5
| ACA0.00000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 0.45954 | 4424.6 | 0.45942 | 4424.6
| ACA0.0000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 5.6234 | 1917.2 | 5.6233 | 1916.9
| ACA0.000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 57.718 | 221.49 | 57.718 | 221.34
| ACA0.0000000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.02 | 0.045 | kHz | 0.27 | 1700 | 0.23101 | 1732
| ACA0.000000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 0.27 | 1700 | 0.23547 | 1728.4
| ACA0.00000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 0.55585 | 1490.7 | 0.55573 | 1490.5
| ACA0.0000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 5.7532 | 345.4 | 5.7532 | 345.04
| ACA0.000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 57.733 | 35.687 | 57.733 | 35.535
| ACA0.0000000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.045 | 0.1 | kHz | 0.24 | 690 | 0.231 | 692.83
| ACA0.000000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.045 | 0.1 | kHz | 0.24 | 690 | 0.23687 | 688.78
| ACA0.00000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.045 | 0.1 | kHz | 0.60214 | 487.92 | 0.60204 | 487.48
| ACA0.0000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.045 | 0.1 | kHz | 5.7741 | 73.402 | 5.7741 | 73.027
| ACA0.000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 0.045 | 0.1 | kHz | 57.736 | 7.5343 | 57.736 | 7.3822
| ACA0.0000000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.1 | 5 | kHz | 0.24 | 350 | 0.23064 | 351.87
| ACA0.000000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.1 | 5 | kHz | 0.24 | 350 | 0.23716 | 348.23
| ACA0.00000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.1 | 5 | kHz | 0.61518 | 205.84 | 0.6151 | 205.2
| ACA0.0000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.1 | 5 | kHz | 5.7771 | 26.243 | 5.7771 | 25.865
| ACA0.000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 0.1 | 5 | kHz | 57.736 | 2.7751 | 57.736 | 2.6229
| ACA0.0000000010.00015 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 5 | 20 | kHz | 0.24 | 700 | 0.23152 | 700.45
| ACA0.000000010.00015 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 5 | 20 | kHz | 0.24 | 700 | 0.23736 | 696.41
| ACA0.00000010.00015 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 5 | 20 | kHz | 0.60202 | 494.94 | 0.60193 | 494.5
| ACA0.0000010.00015 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 5 | 20 | kHz | 5.774 | 74.762 | 5.774 | 74.388
| ACA0.000010.00015 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 5 | 20 | kHz | 57.736 | 7.6722 | 57.736 | 7.52
| ACA0.0000000010.000120 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 20 | 50 | kHz | 0.47 | 4700 | 0.46128 | 4664.4
| ACA0.000000010.000120 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 20 | 50 | kHz | 0.47 | 4700 | 0.46325 | 4662.8
| ACA0.00000010.000120 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 20 | 50 | kHz | 0.64039 | 4529.6 | 0.64112 | 4517.1
| ACA0.0000010.000120 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 20 | 50 | kHz | 5.6399 | 2082.4 | 5.6396 | 2081.3
| ACA0.000010.000120 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 20 | 50 | kHz | 57.719 | 244.83 | 57.719 | 244.19
| ACA0.0000000010.000150 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 50 | 100 | kHz | 1.8 | 6400 | 1.7354 | 6419.2
| ACA0.000000010.000150 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 50 | 100 | kHz | 1.8 | 6400 | 1.7363 | 6418.8
| ACA0.00000010.000150 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 50 | 100 | kHz | 1.8113 | 6388.8 | 1.8106 | 6365.7
| ACA0.0000010.000150 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 50 | 100 | kHz | 5.8316 | 4181 | 5.8282 | 4167.5
| ACA0.000010.000150 | ACA | 0.0001 | 0.001 | 1e-05 | A | µA | 50 | 100 | kHz | 57.727 | 595.61 | 57.726 | 584.99
| ACA0.00000010.0010.001 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 2.5 | 4600 | 2.3097 | 4618.8
| ACA0.0000010.0010.001 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 2.5002 | 4600.2 | 2.3362 | 4616.5
| ACA0.000010.0010.001 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 4.6001 | 4424.5 | 4.5951 | 4424.6
| ACA0.00010.0010.001 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.001 | 0.02 | kHz | 56.235 | 1918 | 56.233 | 1917.1
| ACA0.0010.0010.001 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 0.001 | 0.02 | kHz | 577.18 | 221.92 | 577.18 | 221.43
| ACA0.00000010.0010.02 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 2.7 | 1700 | 2.3102 | 1732
| ACA0.0000010.0010.02 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 2.7 | 1700 | 2.3551 | 1728.4
| ACA0.000010.0010.02 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.564 | 1491 | 5.5589 | 1490.5
| ACA0.00010.0010.02 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 57.534 | 346.39 | 57.532 | 345.24
| ACA0.0010.0010.02 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 0.02 | 0.045 | kHz | 577.33 | 36.111 | 577.33 | 35.619
| ACA0.00000010.0010.045 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.045 | 0.1 | kHz | 2.4 | 690 | 2.3102 | 692.83
| ACA0.0000010.0010.045 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.045 | 0.1 | kHz | 2.4 | 690 | 2.3694 | 688.8
| ACA0.000010.0010.045 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.045 | 0.1 | kHz | 6.0266 | 489.04 | 6.0214 | 487.7
| ACA0.00010.0010.045 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.045 | 0.1 | kHz | 57.743 | 74.446 | 57.741 | 73.235
| ACA0.0010.0010.045 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 0.045 | 0.1 | kHz | 577.36 | 7.959 | 577.36 | 7.467
| ACA0.00000010.0010.1 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.1 | 5 | kHz | 2.4 | 350 | 2.3064 | 351.39
| ACA0.0000010.0010.1 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.1 | 5 | kHz | 2.4 | 350 | 2.3722 | 347.81
| ACA0.000010.0010.1 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.1 | 5 | kHz | 6.1567 | 207.16 | 6.1521 | 205.17
| ACA0.00010.0010.1 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.1 | 5 | kHz | 57.773 | 27.244 | 57.771 | 26.022
| ACA0.0010.0010.1 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 0.1 | 5 | kHz | 577.36 | 3.1944 | 577.36 | 2.7024
| ACA0.00000010.0015 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 5 | 20 | kHz | 2.4 | 700 | 2.3049 | 700.68
| ACA0.0000010.0015 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 5 | 20 | kHz | 2.4 | 700 | 2.364 | 696.65
| ACA0.000010.0015 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 5 | 20 | kHz | 6.021 | 495.8 | 6.0158 | 494.47
| ACA0.00010.0015 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 5 | 20 | kHz | 57.742 | 75.721 | 57.74 | 74.51
| ACA0.0010.0015 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 5 | 20 | kHz | 577.36 | 8.0881 | 577.36 | 7.5962
| ACA0.00000010.00120 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 20 | 50 | kHz | 4.7 | 4700 | 4.6092 | 4653.9
| ACA0.0000010.00120 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 20 | 50 | kHz | 4.7 | 4700 | 4.6286 | 4652.3
| ACA0.000010.00120 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 20 | 50 | kHz | 6.3882 | 4531.2 | 6.409 | 4506.5
| ACA0.00010.00120 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 20 | 50 | kHz | 56.403 | 2074.3 | 56.401 | 2073.5
| ACA0.0010.00120 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 20 | 50 | kHz | 577.19 | 243.6 | 577.19 | 243.11
| ACA0.00000010.00150 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 50 | 100 | kHz | 18 | 6400 | 17.277 | 6419.6
| ACA0.0000010.00150 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 50 | 100 | kHz | 18 | 6400 | 17.286 | 6419.1
| ACA0.000010.00150 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 50 | 100 | kHz | 18.019 | 6398.1 | 18.03 | 6365.4
| ACA0.00010.00150 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 50 | 100 | kHz | 58.284 | 4168.3 | 58.255 | 4162.3
| ACA0.0010.00150 | ACA | 0.001 | 0.01 | 0.001 | A | µA | 50 | 100 | kHz | 577.27 | 587.92 | 577.26 | 582.84
| ACA0.0000010.010.001 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 25 | 4600 | 23.097 | 4618.8
| ACA0.000010.010.001 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 25 | 4600.2 | 23.361 | 4616.5
| ACA0.00010.010.001 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.001 | 0.02 | kHz | 45.971 | 4424.6 | 45.946 | 4424.6
| ACA0.0010.010.001 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.001 | 0.02 | kHz | 562.35 | 1917.4 | 562.33 | 1917
| ACA0.010.010.001 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 0.001 | 0.02 | kHz | 5771.8 | 221.62 | 5771.8 | 221.37
| ACA0.0000010.010.02 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 27 | 1700 | 23.101 | 1732
| ACA0.000010.010.02 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 27 | 1700 | 23.548 | 1728.4
| ACA0.00010.010.02 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 55.608 | 1490.7 | 55.578 | 1490.5
| ACA0.0010.010.02 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.02 | 0.045 | kHz | 575.33 | 345.69 | 575.32 | 345.1
| ACA0.010.010.02 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 0.02 | 0.045 | kHz | 5773.3 | 35.811 | 5773.3 | 35.56
| ACA0.0000010.010.045 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.045 | 0.1 | kHz | 24 | 690 | 23.101 | 692.83
| ACA0.000010.010.045 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.045 | 0.1 | kHz | 24 | 690 | 23.69 | 688.78
| ACA0.00010.010.045 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.045 | 0.1 | kHz | 60.235 | 488.22 | 60.208 | 487.54
| ACA0.0010.010.045 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.045 | 0.1 | kHz | 577.42 | 73.708 | 577.41 | 73.088
| ACA0.010.010.045 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 0.045 | 0.1 | kHz | 5773.6 | 7.6591 | 5773.6 | 7.4071
| ACA0.0000010.010.1 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.1 | 5 | kHz | 24 | 350 | 23.062 | 351.38
| ACA0.000010.010.1 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.1 | 5 | kHz | 24 | 350 | 23.716 | 347.76
| ACA0.00010.010.1 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.1 | 5 | kHz | 61.539 | 205.95 | 61.515 | 204.93
| ACA0.0010.010.1 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.1 | 5 | kHz | 577.72 | 26.5 | 577.71 | 25.873
| ACA0.010.010.1 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 0.1 | 5 | kHz | 5773.6 | 2.8946 | 5773.6 | 2.6425
| ACA0.0000010.015 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 5 | 20 | kHz | 24 | 700 | 23.045 | 700.7
| ACA0.000010.015 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 5 | 20 | kHz | 24 | 700 | 23.633 | 696.65
| ACA0.00010.015 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 5 | 20 | kHz | 60.178 | 495 | 60.151 | 494.32
| ACA0.0010.015 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 5 | 20 | kHz | 577.41 | 74.984 | 577.4 | 74.364
| ACA0.010.015 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 5 | 20 | kHz | 5773.6 | 7.7884 | 5773.6 | 7.5363
| ACA0.0000010.0120 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 20 | 50 | kHz | 47 | 4700 | 46.032 | 4653.9
| ACA0.000010.0120 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 20 | 50 | kHz | 47 | 4700 | 46.229 | 4652.3
| ACA0.00010.0120 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 20 | 50 | kHz | 63.904 | 4531 | 64.059 | 4506.3
| ACA0.0010.0120 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 20 | 50 | kHz | 564.07 | 2074.8 | 564.01 | 2073.3
| ACA0.010.0120 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 20 | 50 | kHz | 5771.9 | 244.08 | 5771.9 | 243.17
| ACA0.0000010.0150 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 50 | 100 | kHz | 180 | 6400 | 172.74 | 6420
| ACA0.000010.0150 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 50 | 100 | kHz | 180 | 6400 | 172.83 | 6419.6
| ACA0.00010.0150 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 50 | 100 | kHz | 180.17 | 6398.4 | 180.26 | 6366.5
| ACA0.0010.0150 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 50 | 100 | kHz | 582.7 | 4181.6 | 582.52 | 4165
| ACA0.010.0150 | ACA | 0.01 | 0.1 | 0.01 | A | µA | 50 | 100 | kHz | 5772.6 | 597.02 | 5772.6 | 584.66
| ACA0.000010.10.001 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.001 | 0.02 | kHz | 0.25 | 4.6 | 0.23097 | 4.6188
| ACA0.00010.10.001 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.001 | 0.02 | kHz | 0.25004 | 4.6001 | 0.2336 | 4.6165
| ACA0.0010.10.001 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.001 | 0.02 | kHz | 0.4598 | 4.4244 | 0.45947 | 4.4246
| ACA0.010.10.001 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.001 | 0.02 | kHz | 5.6235 | 1.9171 | 5.6234 | 1.9169
| ACA0.10.10.001 | ACA | 0.1 | 1 | 0.1 | A | mA | 0.001 | 0.02 | kHz | 57.718 | 0.22147 | 57.718 | 0.22134
| ACA0.000010.10.02 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.23099 | 1.8475
| ACA0.00010.10.02 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.23533 | 1.844
| ACA0.0010.10.02 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.02 | 0.045 | kHz | 0.55129 | 1.6063 | 0.55089 | 1.6064
| ACA0.010.10.02 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.02 | 0.045 | kHz | 5.7501 | 0.38681 | 5.7499 | 0.38651
| ACA0.10.10.02 | ACA | 0.1 | 1 | 0.1 | A | mA | 0.02 | 0.045 | kHz | 57.733 | 0.04006 | 57.733 | 0.039927
| ACA0.000010.10.045 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.045 | 0.1 | kHz | 0.34 | 0.82 | 0.231 | 0.92381
| ACA0.00010.10.045 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.045 | 0.1 | kHz | 0.34 | 0.82 | 0.23651 | 0.91976
| ACA0.0010.10.045 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 0.59229 | 0.69947 | 0.59189 | 0.69932
| ACA0.010.10.045 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 5.7711 | 0.11741 | 5.7709 | 0.11709
| ACA0.10.10.045 | ACA | 0.1 | 1 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 57.735 | 0.011991 | 57.735 | 0.011858
| ACA0.000010.10.1 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.1 | 5 | kHz | 0.24 | 1.2 | 0.23058 | 1.1597
| ACA0.00010.10.1 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.1 | 5 | kHz | 0.24 | 1.2 | 0.23578 | 1.1557
| ACA0.0010.10.1 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.1 | 5 | kHz | 0.5814 | 0.92455 | 0.581 | 0.92448
| ACA0.010.10.1 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.1 | 5 | kHz | 5.767 | 0.17222 | 5.7668 | 0.1719
| ACA0.10.10.1 | ACA | 0.1 | 1 | 0.1 | A | mA | 0.1 | 5 | kHz | 57.735 | 0.017607 | 57.735 | 0.017473
| ACA0.000010.15 | ACA | 0.1 | 1 | 1e-05 | A | mA | 5 | 20 | kHz | 0.23 | 3.5 | 0.22975 | 3.4776
| ACA0.00010.15 | ACA | 0.1 | 1 | 0.0001 | A | mA | 5 | 20 | kHz | 0.23056 | 3.4994 | 0.23289 | 3.4749
| ACA0.0010.15 | ACA | 0.1 | 1 | 0.001 | A | mA | 5 | 20 | kHz | 0.49087 | 3.2614 | 0.49055 | 3.2615
| ACA0.010.15 | ACA | 0.1 | 1 | 0.01 | A | mA | 5 | 20 | kHz | 5.6849 | 1.177 | 5.6848 | 1.1767
| ACA0.10.15 | ACA | 0.1 | 1 | 0.1 | A | mA | 5 | 20 | kHz | 57.726 | 0.12907 | 57.726 | 0.12894
| ACA0.000010.120 | ACA | 0.1 | 1 | 1e-05 | A | mA | 20 | 50 | kHz | 0.46 | 12 | 0.45898 | 11.587
| ACA0.00010.120 | ACA | 0.1 | 1 | 0.0001 | A | mA | 20 | 50 | kHz | 0.46 | 12 | 0.4602 | 11.586
| ACA0.0010.120 | ACA | 0.1 | 1 | 0.001 | A | mA | 20 | 50 | kHz | 0.52661 | 11.933 | 0.56908 | 11.491
| ACA0.010.120 | ACA | 0.1 | 1 | 0.01 | A | mA | 20 | 50 | kHz | 5.1799 | 8.187 | 5.1783 | 8.1815
| ACA0.10.120 | ACA | 0.1 | 1 | 0.1 | A | mA | 20 | 50 | kHz | 57.623 | 1.3639 | 57.623 | 1.3577
| ACA0.000110.001 | ACA | 1 | 10 | 0.0001 | A | mA | 0.001 | 0.05 | kHz | 0.5 | 0.19 | 0.49922 | 0.12453
| ACA0.00110.001 | ACA | 1 | 10 | 0.001 | A | mA | 0.001 | 0.05 | kHz | 0.67982 | 0.17202 | 0.73877 | 0.10982
| ACA0.0110.001 | ACA | 1 | 10 | 0.01 | A | mA | 0.001 | 0.05 | kHz | 5.7843 | 0.024904 | 5.7824 | 0.024912
| ACA0.110.001 | ACA | 1 | 10 | 0.1 | A | mA | 0.001 | 0.05 | kHz | 57.737 | 0.0025534 | 57.737 | 0.0025534
| ACA110.001 | ACA | 1 | 10 | 1 | A | mA | 0.001 | 0.05 | kHz | 577.36 | 0.00025541 | 577.36 | 0.00025541
| ACA0.000110.05 | ACA | 1 | 10 | 0.0001 | A | mA | 0.05 | 1 | kHz | 0.46 | 0.18 | 0.43867 | 0.1699
| ACA0.00110.05 | ACA | 1 | 10 | 0.001 | A | mA | 0.05 | 1 | kHz | 0.68308 | 0.15769 | 0.68455 | 0.15292
| ACA0.0110.05 | ACA | 1 | 10 | 0.01 | A | mA | 0.05 | 1 | kHz | 5.7686 | 0.038993 | 5.7667 | 0.039005
| ACA0.110.05 | ACA | 1 | 10 | 0.1 | A | mA | 0.05 | 1 | kHz | 57.736 | 0.0040392 | 57.735 | 0.0040393
| ACA110.05 | ACA | 1 | 10 | 1 | A | mA | 0.05 | 1 | kHz | 577.36 | 0.00040408 | 577.36 | 0.00040408
| ACA0.000111 | ACA | 1 | 10 | 0.0001 | A | mA | 1 | 5 | kHz | 0.71 | 0.17 | 0.68947 | 0.17149
| ACA0.00111 | ACA | 1 | 10 | 0.001 | A | mA | 1 | 5 | kHz | 0.87874 | 0.15946 | 0.87592 | 0.15963
| ACA0.0111 | ACA | 1 | 10 | 0.01 | A | mA | 1 | 5 | kHz | 5.7932 | 0.046295 | 5.7912 | 0.046309
| ACA0.111 | ACA | 1 | 10 | 0.1 | A | mA | 1 | 5 | kHz | 57.738 | 0.0048469 | 57.737 | 0.0048469
| ACA111 | ACA | 1 | 10 | 1 | A | mA | 1 | 5 | kHz | 577.36 | 0.00048493 | 577.36 | 0.00048493
| ACA0.0001100.001 | ACA | 10 | 20 | 0.0001 | A | mA | 0.001 | 0.05 | kHz | 0.92 | 0.49 | 0.84722 | 0.49356
| ACA0.001100.001 | ACA | 10 | 20 | 0.001 | A | mA | 0.001 | 0.05 | kHz | 0.92 | 0.49223 | 0.88886 | 0.49225
| ACA0.01100.001 | ACA | 10 | 20 | 0.01 | A | mA | 0.001 | 0.05 | kHz | 4.1706 | 0.40025 | 4.1688 | 0.40029
| ACA0.1100.001 | ACA | 10 | 20 | 0.1 | A | mA | 0.001 | 0.05 | kHz | 57.328 | 0.069761 | 57.327 | 0.069762
| ACA1100.001 | ACA | 10 | 20 | 1 | A | mA | 0.001 | 0.05 | kHz | 577.31 | 0.0070525 | 577.31 | 0.0070525
| ACA0.0001100.05 | ACA | 10 | 20 | 0.0001 | A | mA | 0.05 | 1 | kHz | 2 | 0.5 | -2.6161 | 0.62219
| ACA0.001100.05 | ACA | 10 | 20 | 0.001 | A | mA | 0.05 | 1 | kHz | 2 | 0.5 | -2.5419 | 0.61932
| ACA0.01100.05 | ACA | 10 | 20 | 0.01 | A | mA | 0.05 | 1 | kHz | 2 | 0.5 | 2.216 | 0.45911
| ACA0.1100.05 | ACA | 10 | 20 | 0.1 | A | mA | 0.05 | 1 | kHz | 57.131 | 0.071796 | 57.13 | 0.071797
| ACA1100.05 | ACA | 10 | 20 | 1 | A | mA | 0.05 | 1 | kHz | 577.29 | 0.0072378 | 577.29 | 0.0072378
| ACA0.0001101 | ACA | 10 | 20 | 0.0001 | A | mA | 1 | 5 | kHz | 2 | 0.85 | -10.611 | 1.4719
| ACA0.001101 | ACA | 10 | 20 | 0.001 | A | mA | 1 | 5 | kHz | 2 | 0.85 | -10.541 | 1.4688
| ACA0.01101 | ACA | 10 | 20 | 0.01 | A | mA | 1 | 5 | kHz | 2 | 0.91928 | -5.5198 | 1.2606
| ACA0.1101 | ACA | 10 | 20 | 0.1 | A | mA | 1 | 5 | kHz | 55.036 | 0.28461 | 55.036 | 0.28461
| ACA1101 | ACA | 10 | 20 | 1 | A | mA | 1 | 5 | kHz | 577.08 | 0.029227 | 577.08 | 0.029227
| Frequency0.001 | Frequency | 0.001 | 0.04 |  | kHz | Hz |  |  |  | 0 | 0.0006 | 0 | 0.00057779
| Frequency0.04 | Frequency | 0.04 | 10000 |  | kHz | Hz |  |  |  | 0 | 0.00012 | 0 | 0.00011563
| TC Ind, E 0.1-454 | TC Ind, E | -454 | 300 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.084427 | 0
| TC Ind, E 0.1300.001 | TC Ind, E | 300.001 | 1100 | 0.1 | °F | °F |  |  |  | 0.09 | 0 | 0.083674 | 0
| TC Ind, E 0.11100.001 | TC Ind, E | 1100.001 | 1832 | 0.1 | °F | °F |  |  |  | 0.091 | 0 | 0.086131 | 0
| TC Ind, E 1-454 | TC Ind, E | -454 | 300 | 1 | °F | °F |  |  |  | 0.60174 | 0 | 0.58063 | 0
| TC Ind, E 1300.001 | TC Ind, E | 300.001 | 1100 | 1 | °F | °F |  |  |  | 0.5829 | 0 | 0.58052 | 0
| TC Ind, E 11100.001 | TC Ind, E | 1100.001 | 1832 | 1 | °F | °F |  |  |  | 0.58326 | 0 | 0.58088 | 0
| TC Ind, E 0.1-270 | TC Ind, E | -270 | 149 | 0.1 | °C | °C |  |  |  | 0.080212 | 0 | 0.06714 | 0
| TC Ind, E 0.1149.001 | TC Ind, E | 149.001 | 593 | 0.1 | °C | °C |  |  |  | 0.068418 | 0 | 0.066835 | 0
| TC Ind, E 0.1593.001 | TC Ind, E | 593.001 | 1000 | 0.1 | °C | °C |  |  |  | 0.069362 | 0 | 0.067802 | 0
| TC Ind, E 1-270 | TC Ind, E | -270 | 149 | 1 | °C | °C |  |  |  | 0.58499 | 0 | 0.57837 | 0
| TC Ind, E 1149.001 | TC Ind, E | 149.001 | 593 | 1 | °C | °C |  |  |  | 0.57907 | 0 | 0.57833 | 0
| TC Ind, E 1593.001 | TC Ind, E | 593.001 | 1000 | 1 | °C | °C |  |  |  | 0.57918 | 0 | 0.57844 | 0
| TC Ind, J0.1-346 | TC Ind, J | -346 | 499.999 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.091294 | 0
| TC Ind, J0.1500 | TC Ind, J | 500 | 1350 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.084615 | 0
| TC Ind, J0.11350.001 | TC Ind, J | 1350.001 | 2192 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.088259 | 0
| TC Ind, J1-346 | TC Ind, J | -346 | 499.999 | 1 | °F | °F |  |  |  | 0.61864 | 0 | 0.58167 | 0
| TC Ind, J1500 | TC Ind, J | 500 | 1350 | 1 | °F | °F |  |  |  | 0.59111 | 0 | 0.58065 | 0
| TC Ind, J11350.001 | TC Ind, J | 1350.001 | 2192 | 1 | °F | °F |  |  |  | 0.5889 | 0 | 0.5812 | 0
| TC Ind, J0.1-210 | TC Ind, J | -210 | 259.999 | 0.1 | °C | °C |  |  |  | 0.09118 | 0 | 0.069924 | 0
| TC Ind, J0.1260 | TC Ind, J | 260 | 732 | 0.1 | °C | °C |  |  |  | 0.073531 | 0 | 0.067206 | 0
| TC Ind, J0.1732.001 | TC Ind, J | 732.001 | 1200 | 0.1 | °C | °C |  |  |  | 0.073545 | 0 | 0.068653 | 0
| TC Ind, J1-210 | TC Ind, J | -210 | 259.999 | 1 | °C | °C |  |  |  | 0.59041 | 0 | 0.5787 | 0
| TC Ind, J1260 | TC Ind, J | 260 | 732 | 1 | °C | °C |  |  |  | 0.58163 | 0 | 0.57837 | 0
| TC Ind, J1732.001 | TC Ind, J | 732.001 | 1200 | 1 | °C | °C |  |  |  | 0.58094 | 0 | 0.57854 | 0
| TC Ind, K0.1-454 | TC Ind, K | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.76 | 0 | 0.74713 | 0
| TC Ind, K0.1-400 | TC Ind, K | -400 | 499.999 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13406 | 0
| TC Ind, K0.1500 | TC Ind, K | 500 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.087399 | 0
| TC Ind, K0.11500.001 | TC Ind, K | 1500.001 | 2500 | 0.1 | °F | °F |  |  |  | 0.13 | 0 | 0.094912 | 0
| TC Ind, K1-454 | TC Ind, K | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.97819 | 0 | 0.94244 | 0
| TC Ind, K1-400 | TC Ind, K | -400 | 499.999 | 1 | °F | °F |  |  |  | 0.6379 | 0 | 0.58989 | 0
| TC Ind, K1500 | TC Ind, K | 500 | 1500 | 1 | °F | °F |  |  |  | 0.59928 | 0 | 0.58107 | 0
| TC Ind, K11500.001 | TC Ind, K | 1500.001 | 2500 | 1 | °F | °F |  |  |  | 0.60731 | 0 | 0.58224 | 0
| TC Ind, K0.1-270 | TC Ind, K | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.4291 | 0 | 0.42288 | 0
| TC Ind, K0.1-240 | TC Ind, K | -240 | 259.999 | 0.1 | °C | °C |  |  |  | 0.11175 | 0 | 0.089102 | 0
| TC Ind, K0.1260 | TC Ind, K | 260 | 816 | 0.1 | °C | °C |  |  |  | 0.079535 | 0 | 0.068316 | 0
| TC Ind, K0.1816.001 | TC Ind, K | 816.001 | 1371 | 0.1 | °C | °C |  |  |  | 0.085999 | 0 | 0.071379 | 0
| TC Ind, K1-270 | TC Ind, K | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.72802 | 0 | 0.71332 | 0
| TC Ind, K1-240 | TC Ind, K | -240 | 259.999 | 1 | °C | °C |  |  |  | 0.59677 | 0 | 0.58133 | 0
| TC Ind, K1260 | TC Ind, K | 260 | 816 | 1 | °C | °C |  |  |  | 0.58421 | 0 | 0.5785 | 0
| TC Ind, K1816.001 | TC Ind, K | 816.001 | 1371 | 1 | °C | °C |  |  |  | 0.58677 | 0 | 0.57887 | 0
| TC Ind, N0.1-454 | TC Ind, N | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 1.5 | 0 | 1.4092 | 0
| TC Ind, N0.1-400 | TC Ind, N | -400 | -150.001 | 0.1 | °F | °F |  |  |  | 0.24 | 0 | 0.1965 | 0
| TC Ind, N0.1-150 | TC Ind, N | -150 | 700 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.087995 | 0
| TC Ind, N0.1700.001 | TC Ind, N | 700.001 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.086893 | 0
| TC Ind, N0.11500.001 | TC Ind, N | 1500.001 | 2372 | 0.1 | °F | °F |  |  |  | 0.13 | 0 | 0.091887 | 0
| TC Ind, N1-454 | TC Ind, N | -454 | -400.001 | 1 | °F | °F |  |  |  | 1.5424 | 0 | 1.5218 | 0
| TC Ind, N1-400 | TC Ind, N | -400 | -150.001 | 1 | °F | °F |  |  |  | 0.65705 | 0 | 0.60713 | 0
| TC Ind, N1-150 | TC Ind, N | -150 | 700 | 1 | °F | °F |  |  |  | 0.61816 | 0 | 0.58116 | 0
| TC Ind, N1700.001 | TC Ind, N | 700.001 | 1500 | 1 | °F | °F |  |  |  | 0.60094 | 0 | 0.58099 | 0
| TC Ind, N11500.001 | TC Ind, N | 1500.001 | 2372 | 1 | °F | °F |  |  |  | 0.60341 | 0 | 0.58176 | 0
| TC Ind, N0.1-270 | TC Ind, N | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.79757 | 0 | 0.79451 | 0
| TC Ind, N0.1-240 | TC Ind, N | -240 | -101.001 | 0.1 | °C | °C |  |  |  | 0.1391 | 0 | 0.12034 | 0
| TC Ind, N0.1-101 | TC Ind, N | -101 | 371 | 0.1 | °C | °C |  |  |  | 0.090155 | 0 | 0.068582 | 0
| TC Ind, N0.1371.001 | TC Ind, N | 371.001 | 816 | 0.1 | °C | °C |  |  |  | 0.080367 | 0 | 0.068115 | 0
| TC Ind, N0.1816.001 | TC Ind, N | 816.001 | 1300 | 0.1 | °C | °C |  |  |  | 0.083056 | 0 | 0.070133 | 0
| TC Ind, N1-270 | TC Ind, N | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.99031 | 0 | 0.98043 | 0
| TC Ind, N1-240 | TC Ind, N | -240 | -101.001 | 1 | °C | °C |  |  |  | 0.60329 | 0 | 0.58692 | 0
| TC Ind, N1-101 | TC Ind, N | -101 | 371 | 1 | °C | °C |  |  |  | 0.59026 | 0 | 0.57854 | 0
| TC Ind, N1371.001 | TC Ind, N | 371.001 | 816 | 1 | °C | °C |  |  |  | 0.58474 | 0 | 0.57848 | 0
| TC Ind, N1816.001 | TC Ind, N | 816.001 | 1300 | 1 | °C | °C |  |  |  | 0.58552 | 0 | 0.57872 | 0
| TC Ind, R0.1-58 | TC Ind, R | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.26 | 0 | 0.22209 | 0
| TC Ind, R0.132 | TC Ind, R | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13901 | 0
| TC Ind, R0.11100.001 | TC Ind, R | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.16 | 0 | 0.10379 | 0
| TC Ind, R0.12200.001 | TC Ind, R | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.11772 | 0
| TC Ind, R1-58 | TC Ind, R | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.66463 | 0 | 0.61589 | 0
| TC Ind, R132 | TC Ind, R | 32 | 1100 | 1 | °F | °F |  |  |  | 0.64166 | 0 | 0.59104 | 0
| TC Ind, R11100.001 | TC Ind, R | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.62727 | 0 | 0.58376 | 0
| TC Ind, R12200.001 | TC Ind, R | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.62973 | 0 | 0.58639 | 0
| TC Ind, R0.1-50 | TC Ind, R | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.15076 | 0 | 0.13384 | 0
| TC Ind, R0.10 | TC Ind, R | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.11489 | 0 | 0.091563 | 0
| TC Ind, R0.1593.001 | TC Ind, R | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.098606 | 0 | 0.075211 | 0
| TC Ind, R0.11204.001 | TC Ind, R | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.1034 | 0 | 0.081389 | 0
| TC Ind, R1-50 | TC Ind, R | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60595 | 0 | 0.58984 | 0
| TC Ind, R10 | TC Ind, R | 0 | 593 | 1 | °C | °C |  |  |  | 0.59804 | 0 | 0.58171 | 0
| TC Ind, R1593.001 | TC Ind, R | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59323 | 0 | 0.57936 | 0
| TC Ind, R11204.001 | TC Ind, R | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.59405 | 0 | 0.58019 | 0
| TC Ind, S0.1-58 | TC Ind, S | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.22 | 0 | 0.17931 | 0
| TC Ind, S0.132 | TC Ind, S | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.1267 | 0
| TC Ind, S0.11100.001 | TC Ind, S | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.10379 | 0
| TC Ind, S0.12200.001 | TC Ind, S | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.11772 | 0
| TC Ind, S1-58 | TC Ind, S | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.65158 | 0 | 0.60179 | 0
| TC Ind, S132 | TC Ind, S | 32 | 1100 | 1 | °F | °F |  |  |  | 0.63802 | 0 | 0.58826 | 0
| TC Ind, S11100.001 | TC Ind, S | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.62798 | 0 | 0.58452 | 0
| TC Ind, S12200.001 | TC Ind, S | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.62973 | 0 | 0.58774 | 0
| TC Ind, S0.1-50 | TC Ind, S | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.13143 | 0 | 0.11162 | 0
| TC Ind, S0.10 | TC Ind, S | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.10979 | 0 | 0.085698 | 0
| TC Ind, S0.1593.001 | TC Ind, S | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.098606 | 0 | 0.07706 | 0
| TC Ind, S0.11204.001 | TC Ind, S | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.10578 | 0 | 0.084394 | 0
| TC Ind, S1-50 | TC Ind, S | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60143 | 0 | 0.5852 | 0
| TC Ind, S10 | TC Ind, S | 0 | 593 | 1 | °C | °C |  |  |  | 0.59681 | 0 | 0.58081 | 0
| TC Ind, S1593.001 | TC Ind, S | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59347 | 0 | 0.5796 | 0
| TC Ind, S11204.001 | TC Ind, S | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.59446 | 0 | 0.58062 | 0
| TC Ind, T0.1-454 | TC Ind, T | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.41 | 0 | 0.38495 | 0
| TC Ind, T0.1-400 | TC Ind, T | -400 | 199.999 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12271 | 0
| TC Ind, T0.1200 | TC Ind, T | 200 | 752 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.083059 | 0
| TC Ind, T1-454 | TC Ind, T | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.73948 | 0 | 0.69151 | 0
| TC Ind, T1-400 | TC Ind, T | -400 | 199.999 | 1 | °F | °F |  |  |  | 0.63397 | 0 | 0.58742 | 0
| TC Ind, T1200 | TC Ind, T | 200 | 752 | 1 | °F | °F |  |  |  | 0.59634 | 0 | 0.58043 | 0
| TC Ind, T0.1-270 | TC Ind, T | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.2333 | 0 | 0.22166 | 0
| TC Ind, T0.1-240 | TC Ind, T | -240 | 92.999 | 0.1 | °C | °C |  |  |  | 0.10679 | 0 | 0.083768 | 0
| TC Ind, T0.193 | TC Ind, T | 93 | 400 | 0.1 | °C | °C |  |  |  | 0.076683 | 0 | 0.066601 | 0
| TC Ind, T1-270 | TC Ind, T | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.63271 | 0 | 0.61574 | 0
| TC Ind, T1-240 | TC Ind, T | -240 | 92.999 | 1 | °C | °C |  |  |  | 0.59546 | 0 | 0.58053 | 0
| TC Ind, T193 | TC Ind, T | 93 | 400 | 1 | °C | °C |  |  |  | 0.58328 | 0 | 0.5783 | 0
| PT1000.01-320 | PT100 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.013 | 0 | 0.01065 | 0
| PT1000.0132.001 | PT100 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.01438 | 0
| PT1000.01400.001 | PT100 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.021 | 0 | 0.018636 | 0
| PT1000.01800.001 | PT100 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.027 | 0 | 0.025728 | 0
| PT1000.011200.001 | PT100 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.034 | 0 | 0.032361 | 0
| PT1000.1-320 | PT100 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.05993 | 0 | 0.058424 | 0
| PT1000.132.001 | PT100 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.060704 | 0 | 0.059218 | 0
| PT1000.1400.001 | PT100 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.062415 | 0 | 0.060393 | 0
| PT1000.1800.001 | PT100 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.064886 | 0 | 0.062944 | 0
| PT1000.11200.001 | PT100 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.069137 | 0 | 0.065934 | 0
| PT1001-320 | PT100 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57804 | 0 | 0.57742 | 0
| PT100132.001 | PT100 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57812 | 0 | 0.5775 | 0
| PT1001400.001 | PT100 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57848 | 0 | 0.57762 | 0
| PT1001800.001 | PT100 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57875 | 0 | 0.57789 | 0
| PT10011200.001 | PT100 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57972 | 0 | 0.57823 | 0
| PT1000.01-195.6 | PT100 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.0084736 | 0 | 0.007619 | 0
| PT1000.010.001 | PT100 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.010031 | 0 | 0.0093205 | 0
| PT1000.01204.445 | PT100 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.012223 | 0 | 0.011412 | 0
| PT1000.01426.668 | PT100 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015701 | 0 | 0.015078 | 0
| PT1000.01648.89 | PT100 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019485 | 0 | 0.018608 | 0
| PT1000.1-195.6 | PT100 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058421 | 0 | 0.057949 | 0
| PT1000.10.001 | PT100 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058668 | 0 | 0.058197 | 0
| PT1000.1204.445 | PT100 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.059219 | 0 | 0.058568 | 0
| PT1000.1426.668 | PT100 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060033 | 0 | 0.059392 | 0
| PT1000.1648.89 | PT100 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.06148 | 0 | 0.060384 | 0
| PT1001-195.6 | PT100 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57756 | 0 | 0.57737 | 0
| PT10010.001 | PT100 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57759 | 0 | 0.5774 | 0
| PT1001204.445 | PT100 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.5777 | 0 | 0.57743 | 0
| PT1001426.668 | PT100 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57778 | 0 | 0.57752 | 0
| PT1001648.89 | PT100 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57808 | 0 | 0.57762 | 0
| PT10000.01-320 | PT1000 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.009 | 0 | 0.007986 | 0
| PT10000.0132.001 | PT1000 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014357 | 0
| PT10000.01400.001 | PT1000 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.02 | 0 | 0.018611 | 0
| PT10000.01800.001 | PT1000 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.027 | 0 | 0.025699 | 0
| PT10000.011200.001 | PT1000 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.033 | 0 | 0.032328 | 0
| PT10000.1-320 | PT1000 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.058578 | 0 | 0.057998 | 0
| PT10000.132.001 | PT1000 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.059955 | 0 | 0.059212 | 0
| PT10000.1400.001 | PT1000 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.061114 | 0 | 0.060385 | 0
| PT10000.1800.001 | PT1000 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.064875 | 0 | 0.062932 | 0
| PT10000.11200.001 | PT1000 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.067054 | 0 | 0.065917 | 0
| PT10001-320 | PT1000 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57761 | 0 | 0.57738 | 0
| PT1000132.001 | PT1000 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57781 | 0 | 0.5775 | 0
| PT10001400.001 | PT1000 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57793 | 0 | 0.57762 | 0
| PT10001800.001 | PT1000 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57875 | 0 | 0.57789 | 0
| PT100011200.001 | PT1000 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57875 | 0 | 0.57823 | 0
| PT10000.01-195.6 | PT1000 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.0069243 | 0 | 0.0065367 | 0
| PT10000.010.001 | PT1000 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.009669 | 0 | 0.0093092 | 0
| PT10000.01204.445 | PT1000 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.011695 | 0 | 0.0114 | 0
| PT10000.01426.668 | PT1000 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015686 | 0 | 0.015063 | 0
| PT10000.01648.89 | PT1000 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.018902 | 0 | 0.018591 | 0
| PT10000.1-195.6 | PT1000 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.057997 | 0 | 0.057816 | 0
| PT10000.10.001 | PT1000 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058429 | 0 | 0.058195 | 0
| PT10000.1204.445 | PT1000 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.058799 | 0 | 0.058566 | 0
| PT10000.1426.668 | PT1000 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060029 | 0 | 0.059388 | 0
| PT10000.1648.89 | PT1000 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.060764 | 0 | 0.060379 | 0
| PT10001-195.6 | PT1000 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57743 | 0 | 0.57736 | 0
| PT100010.001 | PT1000 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57749 | 0 | 0.5774 | 0
| PT10001204.445 | PT1000 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.57753 | 0 | 0.57743 | 0
| PT10001426.668 | PT1000 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57778 | 0 | 0.57752 | 0
| PT10001648.89 | PT1000 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57778 | 0 | 0.57762 | 0

## Standard3
Updated: | 2019-08-23 | 34420A
| Main Name | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCVmeas0.000000010 | DCVmeas | 0 | 0.001 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.027845 | 71.95
| DCVmeas0.00000010 | DCVmeas | 0 | 0.001 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.062207 | 53.073
| DCVmeas0.0000010 | DCVmeas | 0 | 0.001 | 1e-06 | V | µV |  |  |  | 0.62 | 4.5 | 0.5802 | 8.2952
| DCVmeas0.000010 | DCVmeas | 0 | 0.001 | 1e-05 | V | µV |  |  |  | 5.8 | 0.89276 | 5.8 | 0.83626
| DCVmeas0.00010 | DCVmeas | 0 | 0.001 | 0.0001 | V | µV |  |  |  | 58 | 0.10624 | 58 | 0.083633
| DCVmeas0.0010 | DCVmeas | 0 | 0.001 | 0.001 | V | µV |  |  |  | 580 | 0.017406 | 580 | 0.0083633
| DCVmeas0.000000010.001 | DCVmeas | 0.001 | 0.01 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.060672 | 48.506
| DCVmeas0.00000010.001 | DCVmeas | 0.001 | 0.01 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.076236 | 47.254
| DCVmeas0.0000010.001 | DCVmeas | 0.001 | 0.01 | 1e-06 | V | µV |  |  |  | 0.62 | 19.155 | 0.56725 | 22.912
| DCVmeas0.000010.001 | DCVmeas | 0.001 | 0.01 | 1e-05 | V | µV |  |  |  | 5.7983 | 2.757 | 5.7983 | 2.7323
| DCVmeas0.00010.001 | DCVmeas | 0.001 | 0.01 | 0.0001 | V | µV |  |  |  | 58 | 0.2838 | 58 | 0.27385
| DCVmeas0.0010.001 | DCVmeas | 0.001 | 0.01 | 0.001 | V | µV |  |  |  | 580 | 0.031364 | 580 | 0.027386
| DCVmeas0.00000010.01 | DCVmeas | 0.01 | 0.1 | 1e-07 | V | µV |  |  |  | 0.48 | 37 | 0.47363 | 36.624
| DCVmeas0.0000010.01 | DCVmeas | 0.01 | 0.1 | 1e-06 | V | µV |  |  |  | 0.6696 | 35.075 | 0.66823 | 35.079
| DCVmeas0.000010.01 | DCVmeas | 0.01 | 0.1 | 1e-05 | V | µV |  |  |  | 5.7207 | 14.051 | 5.7199 | 14.036
| DCVmeas0.00010.01 | DCVmeas | 0.01 | 0.1 | 0.0001 | V | µV |  |  |  | 57.991 | 1.5774 | 57.99 | 1.5689
| DCVmeas0.0010.01 | DCVmeas | 0.01 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.16049 | 580 | 0.1571
| DCVmeas0.010.01 | DCVmeas | 0.01 | 0.1 | 0.01 | V | µV |  |  |  | 5800 | 0.017067 | 5800 | 0.01571
| DCVmeas0.0000010.1 | DCVmeas | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 4.7 | 41 | 4.6518 | 40.782
| DCVmeas0.000010.1 | DCVmeas | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 6.5523 | 39.255 | 6.5389 | 39.26
| DCVmeas0.00010.1 | DCVmeas | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 56.989 | 16.705 | 56.981 | 16.693
| DCVmeas0.0010.1 | DCVmeas | 0.1 | 1 | 0.001 | V | µV |  |  |  | 579.88 | 1.9084 | 579.88 | 1.9012
| DCVmeas0.010.1 | DCVmeas | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5800 | 0.19329 | 5800 | 0.19042
| DCVmeas0.10.1 | DCVmeas | 0.1 | 1 | 0.1 | V | µV |  |  |  | 58000 | 0.020188 | 58000 | 0.019042
| DCVmeas0.000011 | DCVmeas | 1 | 10 | 1e-05 | V | µV |  |  |  | 47 | 35 | 46.376 | 34.813
| DCVmeas0.00011 | DCVmeas | 1 | 10 | 0.0001 | V | µV |  |  |  | 66.458 | 33.232 | 66.382 | 33.232
| DCVmeas0.0011 | DCVmeas | 1 | 10 | 0.001 | V | µV |  |  |  | 572.81 | 12.879 | 572.76 | 12.867
| DCVmeas0.011 | DCVmeas | 1 | 10 | 0.01 | V | µV |  |  |  | 5799.2 | 1.4325 | 5799.1 | 1.4259
| DCVmeas0.11 | DCVmeas | 1 | 10 | 0.1 | V | µV |  |  |  | 58000 | 0.1454 | 58000 | 0.14276
| DCVmeas11 | DCVmeas | 1 | 10 | 1 | V | µV |  |  |  | 580000 | 0.015331 | 580000 | 0.014276
| DCVmeas0.000110 | DCVmeas | 10 | 100 | 0.0001 | V | mV |  |  |  | 0.58 | 0.041 | 0.57895 | 0.040741
| DCVmeas0.00110 | DCVmeas | 10 | 100 | 0.001 | V | mV |  |  |  | 0.7495 | 0.039399 | 0.74879 | 0.039399
| DCVmeas0.0110 | DCVmeas | 10 | 100 | 0.01 | V | mV |  |  |  | 5.7109 | 0.017264 | 5.7105 | 0.017251
| DCVmeas0.110 | DCVmeas | 10 | 100 | 0.1 | V | mV |  |  |  | 57.989 | 0.0019846 | 57.989 | 0.0019773
| DCVmeas110 | DCVmeas | 10 | 100 | 1 | V | mV |  |  |  | 580 | 0.000201 | 580 | 0.00019806
| Resistance0.0000000010 | Resistance | 0 | 1 | 1e-09 | kOhms | µΩ |  |  |  | 3 | 120 | 2.3273 | 80.8131
| Resistance0.000000010 | Resistance | 0 | 1 | 1e-08 | kOhms | µΩ |  |  |  | 3 | 120 | 3.9637 | 79.3768
| Resistance0.00000010 | Resistance | 0 | 1 | 1e-07 | kOhms | µΩ |  |  |  | 63.6928 | 62.6594 | 54.2074 | 47.1632
| Resistance0.0000010 | Resistance | 0 | 1 | 1e-06 | kOhms | µΩ |  |  |  | 600.322 | 96.4808 | 579.4447 | 6.4836
| Resistance0.000010 | Resistance | 0 | 1 | 1e-05 | kOhms | mΩ |  |  |  | 5.8513 | 0.24805 | 5.7999 | 0.00065169 | 0.0058513 | 0.00024805 | 0.0057999 | 6.5169e-07
| Resistance0.000000011 | Resistance | 1 | 10 | 1e-08 | kOhms | µΩ |  |  |  | 24.0351 | 69.2135 | 23.2955 | 69.2642
| Resistance0.00000011 | Resistance | 1 | 10 | 1e-07 | kOhms | µΩ |  |  |  | 66.5586 | 66.0557 | 41.3894 | 67.687
| Resistance0.0000011 | Resistance | 1 | 10 | 1e-06 | kOhms | µΩ |  |  |  | 643.2147 | 54.7509 | 550.1922 | 37.1183
| Resistance0.000011 | Resistance | 1 | 10 | 1e-05 | kOhms | mΩ |  |  |  | 6.0044 | 0.095076 | 5.7959 | 0.004809 | 0.0060044 | 9.5076e-05 | 0.0057959 | 4.809e-06
| Resistance0.00011 | Resistance | 1 | 10 | 0.0001 | kOhms | mΩ |  |  |  | 58.5135 | 0.24789 | 57.9996 | 0.00048274 | 0.058513 | 0.00024789 | 0.058 | 4.8274e-07
| Resistance0.000000110 | Resistance | 10 | 100 | 1e-07 | kOhms | mΩ |  |  |  | 0.24035 | 0.069214 | 0.23295 | 0.069264
| Resistance0.00000110 | Resistance | 10 | 100 | 1e-06 | kOhms | mΩ |  |  |  | 0.66559 | 0.066056 | 0.41389 | 0.067687
| Resistance0.0000110 | Resistance | 10 | 100 | 1e-05 | kOhms | mΩ |  |  |  | 6.4321 | 0.054751 | 5.5019 | 0.037118
| Resistance0.000110 | Resistance | 10 | 100 | 0.0001 | kOhms | mΩ |  |  |  | 60.0444 | 0.095076 | 57.9593 | 0.004809
| Resistance0.00110 | Resistance | 10 | 100 | 0.001 | kOhms | mΩ |  |  |  | 585.1347 | 0.24789 | 579.9959 | 0.000483
| Resistance0.001100 | Resistance | 100 | 1 | 0.001 | kOhm | mΩ |  |  |  | 2.4035 | 0.069214 | 2.3295 | 0.069264
| Resistance0.01100 | Resistance | 100 | 1 | 0.01 | kOhm | mΩ |  |  |  | 6.6559 | 0.066056 | 4.1389 | 0.067687
| Resistance0.1100 | Resistance | 100 | 1 | 0.1 | kOhm | mΩ |  |  |  | 64.3215 | 0.054751 | 55.0192 | 0.037118
| Resistance1100 | Resistance | 100 | 1 | 1 | kOhm | mΩ |  |  |  | 600.444 | 0.095076 | 579.5927 | 0.004809
| Resistance10100 | Resistance | 100 | 1 | 10 | kOhm | Ω |  |  |  | 5.8513 | 0.00024789 | 5.8 | 4.83e-07 | 0.0058513 | 2.4789e-07 | 0.0058 | 4.83e-10
| Resistance0.011 | Resistance | 1 | 10 | 0.01 | kOhm | mΩ |  |  |  | 24.0351 | 0.069214 | 23.2955 | 0.069264
| Resistance0.11 | Resistance | 1 | 10 | 0.1 | kOhm | mΩ |  |  |  | 66.5586 | 0.066056 | 41.3894 | 0.067688
| Resistance11 | Resistance | 1 | 10 | 1 | kOhm | mΩ |  |  |  | 643.2147 | 0.054751 | 550.192 | 0.037118
| Resistance101 | Resistance | 1 | 10 | 10 | kOhm | Ω |  |  |  | 6.0044 | 9.5076e-05 | 5.7959 | 4.809e-06 | 0.0060044 | 9.5076e-08 | 0.0057959 | 4.809e-09
| Resistance1001 | Resistance | 1 | 10 | 100 | kOhm | Ω |  |  |  | 58.5135 | 0.00024789 | 57.9996 | 4.83e-07 | 0.058513 | 2.4789e-07 | 0.058 | 4.83e-10
| Resistance0.110 | Resistance | 10 | 100 | 0.1 | kOhm | Ω |  |  |  | 0.46937 | 6.92e-05 | 0.46349 | 6.93e-05
| Resistance110 | Resistance | 10 | 100 | 1 | kOhm | Ω |  |  |  | 0.82928 | 6.67e-05 | 0.61213 | 6.81e-05
| Resistance1010 | Resistance | 10 | 100 | 10 | kOhm | Ω |  |  |  | 6.4547 | 5.59e-05 | 5.5271 | 3.87e-05
| Resistance10010 | Resistance | 10 | 100 | 100 | kOhm | Ω |  |  |  | 60.0461 | 9.53e-05 | 57.9607 | 5.09e-06
| Resistance100010 | Resistance | 10 | 100 | 1000 | kOhm | Ω |  |  |  | 585.1349 | 0.000248 | 579.996 | 5.1e-07
| Resistance1000100 | Resistance | 100 | 1 | 1000 | kOhm | Ω |  |  |  | 5 | 8.05e-05 | 4.6334 | 8.08e-05
| Resistance10000100 | Resistance | 100 | 1 | 10000 | kOhm | Ω |  |  |  | 8.0349 | 7.84e-05 | 5.9988 | 7.96e-05
| Resistance100000100 | Resistance | 100 | 1 | 100000 | kOhm | Ω |  |  |  | 63.945 | 6.39e-05 | 54.4969 | 4.88e-05
| Resistance1000000100 | Resistance | 100 | 1 | 1000000 | kOhm | Ω |  |  |  | 600.3396 | 9.67e-05 | 579.4589 | 6.8e-06
| Resistance10000000100 | Resistance | 100 | 1 | 10000000 | kOhm | kΩ |  |  |  | 5.8513 | 2.481e-07 | 5.7999 | 6.9e-10 | 0.0058513 | 2.481e-10 | 0.0057999 | 6.9e-13

## Standard4
Updated: | 2019-07-03 | Transmille
| Main Name | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.0000010 | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 12 | 92 | 11.629 | 92.013
| DCV0.000010 | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 12.975 | 86.384 | 12.968 | 86.412
| DCV0.00010 | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.892 | 24.898 | 58.886 | 24.901
| DCV0.0010 | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 577.47 | 2.5928 | 577.47 | 2.5928
| DCV0.010 | DCV | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5773.6 | 0.25939 | 5773.6 | 0.25939
| DCV0.000010.1 | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 36 | 92 | 35.476 | 91.91
| DCV0.00010.1 | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 65.438 | 74.399 | 65.327 | 74.414
| DCV0.0010.1 | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 577.75 | 13.553 | 577.7 | 13.513
| DCV0.010.1 | DCV | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5773.6 | 1.3863 | 5773.6 | 1.3693
| DCV0.10.1 | DCV | 0.1 | 1 | 0.1 | V | µV |  |  |  | 57736 | 0.14376 | 57736 | 0.13695
| DCV0.00011 | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 360 | 92 | 354.49 | 91.832
| DCV0.0011 | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 653.69 | 74.372 | 653.14 | 74.331
| DCV0.011 | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5777.2 | 13.559 | 5777 | 13.491
| DCV0.11 | DCV | 1 | 10 | 0.1 | V | µV |  |  |  | 57736 | 1.3954 | 57736 | 13.496
| DCV11 | DCV | 1 | 10 | 1 | V | µV |  |  |  | 577360 | 0.14808 | 577360 | 13.497
| DCV0.00110 | DCV | 10 | 100 | 0.001 | V | mV |  |  |  | 3.8 | 0.099 | 3.7872 | 0.098876
| DCV0.0110 | DCV | 10 | 100 | 0.01 | V | mV |  |  |  | 6.6632 | 0.081817 | 6.6527 | 0.081798
| DCV0.110 | DCV | 10 | 100 | 0.1 | V | mV |  |  |  | 57.779 | 0.015635 | 57.774 | 0.015558
| DCV110 | DCV | 10 | 100 | 1 | V | mV |  |  |  | 577.36 | 0.0016118 | 577.36 | 0.0015797
| DCV1010 | DCV | 10 | 100 | 10 | V | mV |  |  |  | 5773.6 | 0.00017086 | 5773.6 | 0.00015799
| DCV0.01100 | DCV | 100 | 1020 | 0.01 | V | mV |  |  |  | 36 | 0.092 | 35.438 | 0.091985
| DCV0.1100 | DCV | 100 | 1020 | 0.1 | V | mV |  |  |  | 65.389 | 0.0747 | 65.281 | 0.074678
| DCV1100 | DCV | 100 | 1020 | 1 | V | mV |  |  |  | 577.73 | 0.01374 | 577.68 | 0.013664
| DCV10100 | DCV | 100 | 1020 | 10 | V | mV |  |  |  | 5773.6 | 0.0014166 | 5773.6 | 0.0013851
| DCV100100 | DCV | 100 | 1020 | 100 | V | mV |  |  |  | 57736 | 0.00015112 | 57736 | 0.00013853
| DCA0.0000000010 | DCA | 0 | 0.0001 | 1e-09 | A | nA |  |  |  | 35 | 350000 | 34.652 | 346510
| DCA0.000000010 | DCA | 0 | 0.0001 | 1e-08 | A | nA |  |  |  | 35.128 | 344160 | 35.125 | 344160
| DCA0.00000010 | DCA | 0 | 0.0001 | 1e-07 | A | nA |  |  |  | 67.338 | 228700 | 67.334 | 228660
| DCA0.0000010 | DCA | 0 | 0.0001 | 1e-06 | A | nA |  |  |  | 578.4 | 31085 | 578.39 | 31056
| DCA0.000010 | DCA | 0 | 0.0001 | 1e-05 | A | nA |  |  |  | 5773.7 | 3131 | 5773.7 | 3119.4
| DCA0.000000010.0001 | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 120 | 350000 | 115.73 | 346500
| DCA0.00000010.0001 | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 127.17 | 338660 | 127.11 | 338680
| DCA0.0000010.0001 | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 580.75 | 158910 | 580.7 | 158870
| DCA0.000010.0001 | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.7 | 18379 | 5773.7 | 18351
| DCA0.00010.0001 | DCA | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 57736 | 1849.6 | 57736 | 1838.3
| DCA0.00000010.001 | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 1.2 | 350 | 1.1596 | 346.31
| DCA0.0000010.001 | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 1.2738 | 338.49 | 1.2733 | 338.5
| DCA0.000010.001 | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.8081 | 158.88 | 5.8076 | 158.83
| DCA0.00010.001 | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.737 | 18.38 | 57.737 | 18.348
| DCA0.0010.001 | DCA | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 1.8509 | 577.36 | 1.838
| DCA0.0000010.01 | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 12 | 350 | 11.575 | 346.48
| DCA0.000010.01 | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 12.719 | 338.64 | 12.714 | 338.65
| DCA0.00010.01 | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 58.076 | 158.91 | 58.071 | 158.86
| DCA0.0010.01 | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.37 | 18.382 | 577.37 | 18.35
| DCA0.010.01 | DCA | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 1.8511 | 5773.6 | 1.8382
| DCA0.000010.1 | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 180 | 350 | 173.39 | 346.71
| DCA0.00010.1 | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 181.73 | 341.58 | 181.69 | 341.58
| DCA0.0010.1 | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 595.57 | 181.58 | 595.52 | 181.54
| DCA0.010.1 | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5775.1 | 21.845 | 5775.1 | 21.813
| DCA0.10.1 | DCA | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 2.1992 | 57736 | 2.1864
| DCA0.00011 | DCA | 1 | 10.2 | 0.0001 | A | µA |  |  |  | 2.3 | 0.6 | 2.3096 | 0.58109
| DCA0.0011 | DCA | 1 | 10.2 | 0.001 | A | µA |  |  |  | 2.3155 | 0.63207 | 2.3702 | 0.57712
| DCA0.011 | DCA | 1 | 10.2 | 0.01 | A | µA |  |  |  | 5.8932 | 0.56391 | 6.065 | 0.39152
| DCA0.11 | DCA | 1 | 10.2 | 0.1 | A | µA |  |  |  | 57.625 | 0.18273 | 57.752 | 0.05568
| DCA11 | DCA | 1 | 10.2 | 1 | A | µA |  |  |  | 577.31 | 0.057423 | 577.36 | 0.0055994
| DCA 10 turn0.000110 | DCA 10 turn | 10 | 102 | 0.0001 | A | mA |  |  |  | 17 | 2 | 15.062 | 2.0092
| DCA 10 turn0.00110 | DCA 10 turn | 10 | 102 | 0.001 | A | mA |  |  |  | 17 | 2.0049 | 15.568 | 2.005
| DCA 10 turn0.0110 | DCA 10 turn | 10 | 102 | 0.01 | A | mA |  |  |  | 50.245 | 1.7376 | 50.216 | 1.7376
| DCA 10 turn0.110 | DCA 10 turn | 10 | 102 | 0.1 | A | mA |  |  |  | 574.15 | 0.42882 | 574.14 | 0.42855
| DCA 10 turn110 | DCA 10 turn | 10 | 102 | 1 | A | mA |  |  |  | 5773.2 | 0.044499 | 5773.2 | 0.044381
| DCA 10 turn1010 | DCA 10 turn | 10 | 102 | 10 | A | mA |  |  |  | 57735 | 0.0044871 | 57735 | 0.0044397
| DCA 20 turn0.00120 | DCA 20 turn | 20 | 204 | 0.001 | A | mA |  |  |  | 33 | 2 | 30.102 | 2.0093
| DCA 20 turn0.0120 | DCA 20 turn | 20 | 204 | 0.01 | A | mA |  |  |  | 33 | 2.0082 | 30.364 | 2.0082
| DCA 20 turn0.120 | DCA 20 turn | 20 | 204 | 0.1 | A | mA |  |  |  | 52.698 | 1.9173 | 52.609 | 1.9175
| DCA 20 turn120 | DCA 20 turn | 20 | 204 | 1 | A | mA |  |  |  | 565.98 | 0.78459 | 565.94 | 0.78417
| DCA 20 turn1020 | DCA 20 turn | 20 | 204 | 10 | A | mA |  |  |  | 5772.2 | 0.088899 | 5772.2 | 0.088663
| DCA 50 turn0.0015 | DCA 50 turn | 5 | 51 | 0.001 | A | mA |  |  |  | 4.7 | 2 | 4.5663 | 1.9762
| DCA 50 turn0.015 | DCA 50 turn | 5 | 51 | 0.01 | A | mA |  |  |  | 5.8994 | 1.9542 | 5.769 | 1.9557
| DCA 50 turn0.15 | DCA 50 turn | 5 | 51 | 0.1 | A | mA |  |  |  | 53.045 | 1.3193 | 52.925 | 1.3178
| DCA 50 turn15 | DCA 50 turn | 5 | 51 | 1 | A | mA |  |  |  | 576.56 | 0.20563 | 576.52 | 0.20331
| DCA 50 turn105 | DCA 50 turn | 5 | 51 | 10 | A | mA |  |  |  | 5773.5 | 0.021448 | 5773.5 | 0.020501
| DCA 50 turn0.00150 | DCA 50 turn | 50 | 510 | 0.001 | A | mA |  |  |  | 81 | 2 | 75.253 | 2.0093
| DCA 50 turn0.0150 | DCA 50 turn | 50 | 510 | 0.01 | A | mA |  |  |  | 81 | 2.0091 | 75.371 | 2.0091
| DCA 50 turn0.150 | DCA 50 turn | 50 | 510 | 0.1 | A | mA |  |  |  | 85.622 | 1.9922 | 85.336 | 1.9925
| DCA 50 turn150 | DCA 50 turn | 50 | 510 | 1 | A | mA |  |  |  | 534.37 | 1.3891 | 534.07 | 1.3887
| DCA 50 turn1050 | DCA 50 turn | 50 | 510 | 10 | A | mA |  |  |  | 5765.3 | 0.22054 | 5765.2 | 0.21996
| ACV0.00000100.01 | ACV | 0 | 0.1 | 1e-06 | V | µV | 0.01 | 2 | KHz | 37 | 1100 | 36.41 | 1079.2
| ACV0.0000100.01 | ACV | 0 | 0.1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 37 | 1075.9 | 36.783 | 1075.9
| ACV0.000100.01 | ACV | 0 | 0.1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 67.836 | 874.05 | 67.825 | 873.88
| ACV0.00100.01 | ACV | 0 | 0.1 | 0.001 | V | µV | 0.01 | 2 | KHz | 578.39 | 166.96 | 578.39 | 166.73
| ACV0.0100.01 | ACV | 0 | 0.1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5773.7 | 17.041 | 5773.7 | 16.943
| ACV0.00000102 | ACV | 0 | 0.1 | 1e-06 | V | µV | 2 | 20 | KHz | 48 | 910 | 47.858 | 908.34
| ACV0.0000102 | ACV | 0 | 0.1 | 1e-05 | V | µV | 2 | 20 | KHz | 48.224 | 905.23 | 48.22 | 905.2
| ACV0.000102 | ACV | 0 | 0.1 | 0.0001 | V | µV | 2 | 20 | KHz | 75.081 | 716.91 | 75.076 | 716.71
| ACV0.00102 | ACV | 0 | 0.1 | 0.001 | V | µV | 2 | 20 | KHz | 578.95 | 127.9 | 578.95 | 127.66
| ACV0.0102 | ACV | 0 | 0.1 | 0.01 | V | µV | 2 | 20 | KHz | 5773.7 | 13.028 | 5773.7 | 12.93
| ACV0.000010.10.01 | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 350 | 950 | 349.43 | 946.71
| ACV0.00010.10.01 | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 353.44 | 944.05 | 353.4 | 944.01
| ACV0.0010.10.01 | ACV | 0.1 | 1 | 0.001 | V | µV | 0.01 | 2 | KHz | 651.73 | 767.47 | 651.65 | 767.25
| ACV0.010.10.01 | ACV | 0.1 | 1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5776.6 | 140.98 | 5776.5 | 140.72
| ACV0.10.10.01 | ACV | 0.1 | 1 | 0.1 | V | µV | 0.01 | 2 | KHz | 57736 | 14.373 | 57736 | 14.266
| ACV0.000010.12 | ACV | 0.1 | 1 | 1e-05 | V | µV | 2 | 20 | KHz | 860 | 1700 | 809.52 | 1743.9
| ACV0.00010.12 | ACV | 0.1 | 1 | 0.0001 | V | µV | 2 | 20 | KHz | 860 | 1742.8 | 811.28 | 1742.8
| ACV0.0010.12 | ACV | 0.1 | 1 | 0.001 | V | µV | 2 | 20 | KHz | 977 | 1641.4 | 976.68 | 1641.2
| ACV0.010.12 | ACV | 0.1 | 1 | 0.01 | V | µV | 2 | 20 | KHz | 5806.3 | 507.46 | 5806.1 | 506.9
| ACV0.10.12 | ACV | 0.1 | 1 | 0.1 | V | µV | 2 | 20 | KHz | 57739 | 53.647 | 57739 | 53.394
| ACV0.000110.01 | ACV | 1 | 10 | 0.0001 | V | mV | 0.01 | 2 | KHz | 4 | 0.94 | 3.894 | 0.94247
| ACV0.00110.01 | ACV | 1 | 10 | 0.001 | V | mV | 0.01 | 2 | KHz | 4 | 0.94 | 3.928 | 0.93982
| ACV0.0110.01 | ACV | 1 | 10 | 0.01 | V | mV | 0.01 | 2 | KHz | 6.6043 | 0.76596 | 6.6016 | 0.76513
| ACV0.110.01 | ACV | 1 | 10 | 0.1 | V | mV | 0.01 | 2 | KHz | 57.77 | 0.14158 | 57.77 | 0.14061
| ACV110.01 | ACV | 1 | 10 | 1 | V | mV | 0.01 | 2 | KHz | 577.36 | 0.014653 | 577.36 | 0.014256
| ACV0.000112 | ACV | 1 | 10 | 0.0001 | V | mV | 2 | 20 | KHz | 7.6 | 2.5 | 7.5377 | 2.4794
| ACV0.00112 | ACV | 1 | 10 | 0.001 | V | mV | 2 | 20 | KHz | 7.6 | 2.478 | 7.5553 | 2.4782
| ACV0.0112 | ACV | 1 | 10 | 0.01 | V | mV | 2 | 20 | KHz | 9.2106 | 2.3643 | 9.1972 | 2.3646
| ACV0.112 | ACV | 1 | 10 | 0.1 | V | mV | 2 | 20 | KHz | 57.766 | 0.84266 | 57.757 | 0.84155
| ACV112 | ACV | 1 | 10 | 1 | V | mV | 2 | 20 | KHz | 577.35 | 0.091421 | 577.35 | 0.090855
| ACV0.001100.04 | ACV | 10 | 100 | 0.001 | V | mV | 0.04 | 1 | KHz | 37 | 1 | 34.646 | 1.0223
| ACV0.01100.04 | ACV | 10 | 100 | 0.01 | V | mV | 0.04 | 1 | KHz | 37 | 1.0199 | 35.038 | 1.0196
| ACV0.1100.04 | ACV | 10 | 100 | 0.1 | V | mV | 0.04 | 1 | KHz | 64.815 | 0.83966 | 64.736 | 0.83818
| ACV1100.04 | ACV | 10 | 100 | 1 | V | mV | 0.04 | 1 | KHz | 577.54 | 0.16047 | 577.51 | 0.15847
| ACV10100.04 | ACV | 10 | 100 | 10 | V | mV | 0.04 | 1 | KHz | 5773.6 | 0.016915 | 5773.6 | 0.016088
| ACV0.011000.04 | ACV | 100 | 1020 | 0.01 | V | mV | 0.04 | 1 | KHz | 340 | 1.3 | 333.77 | 1.26
| ACV0.11000.04 | ACV | 100 | 1020 | 0.1 | V | mV | 0.04 | 1 | KHz | 340 | 1.2574 | 337.65 | 1.2572
| ACV11000.04 | ACV | 100 | 1020 | 1 | V | mV | 0.04 | 1 | KHz | 632.25 | 1.0672 | 631.41 | 1.0661
| ACV101000.04 | ACV | 100 | 1020 | 10 | V | mV | 0.04 | 1 | KHz | 5769.9 | 0.22412 | 5769.6 | 0.2222
| ACV1001000.04 | ACV | 100 | 1020 | 100 | V | mV | 0.04 | 1 | KHz | 57735 | 0.023487 | 57735 | 0.022679
| ACA0.0000000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-09 | A | µA | 0.01 | 2 | KHz | 0.48 | 1100 | 0.46873 | 1141.5
| ACA0.000000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.48 | 1141.6 | 0.46876 | 1141.5
| ACA0.00000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.48 | 1135.6 | 0.47225 | 1135.1
| ACA0.0000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 0.74376 | 795.24 | 0.74373 | 793.72
| ACA0.000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8094 | 385.98 | 5.8094 | 385.1
| ACA0.000000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.94 | 1600 | 0.90594 | 1629.7
| ACA0.00000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.94 | 1628.9 | 0.9076 | 1628.7
| ACA0.0000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 1.0611 | 1543.3 | 1.0607 | 1542.4
| ACA0.000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8203 | 517.7 | 5.8201 | 515.72
| ACA0.00010.00010.01 | ACA | 0.0001 | 0.00104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 57.907 | 333.82 | 57.907 | 332.93
| ACA0.00000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 9.5 | 1300 | 9.452 | 1260.9
| ACA0.0000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 9.5 | 1260.3 | 9.4684 | 1260
| ACA0.000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 10.991 | 1185 | 10.986 | 1183.9
| ACA0.00010.0010.01 | ACA | 0.001 | 0.0104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 58.349 | 375.87 | 58.347 | 373.83
| ACA0.0010.0010.01 | ACA | 0.001 | 0.0104 | 0.001 | A | µA | 0.01 | 2 | KHz | 579.11 | 318.45 | 579.11 | 317.56
| ACA0.0000010.010.01 | ACA | 0.01 | 0.104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 110 | 1200 | 100.86 | 1199.9
| ACA0.000010.010.01 | ACA | 0.01 | 0.104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 110 | 1199.4 | 101.02 | 1199.2
| ACA0.00010.010.01 | ACA | 0.01 | 0.104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 115.51 | 1131.1 | 115.47 | 1130
| ACA0.0010.010.01 | ACA | 0.01 | 0.104 | 0.001 | A | µA | 0.01 | 2 | KHz | 584.68 | 364.46 | 584.66 | 362.42
| ACA0.010.010.01 | ACA | 0.01 | 0.104 | 0.01 | A | µA | 0.01 | 2 | KHz | 5791.2 | 317.28 | 5791.2 | 316.4
| ACA0.000010.10.01 | ACA | 0.1 | 1.04 | 1e-05 | A | mA | 0.01 | 2 | KHz | 1.1 | 1.2 | 1.0001 | 1.2081
| ACA0.00010.10.01 | ACA | 0.1 | 1.04 | 0.0001 | A | mA | 0.01 | 2 | KHz | 1.1 | 1.2076 | 1.0017 | 1.2073
| ACA0.0010.10.01 | ACA | 0.1 | 1.04 | 0.001 | A | mA | 0.01 | 2 | KHz | 1.2169 | 1.1384 | 1.1472 | 1.1372
| ACA0.010.10.01 | ACA | 0.1 | 1.04 | 0.01 | A | mA | 0.01 | 2 | KHz | 5.9335 | 0.36602 | 5.845 | 0.36398
| ACA0.10.10.01 | ACA | 0.1 | 1.04 | 0.1 | A | mA | 0.01 | 2 | KHz | 57.949 | 0.31744 | 57.912 | 0.31656
| ACA0.000110.01 | ACA | 1 | 10.4 | 0.0001 | A | mA | 0.01 | 2 | KHz | 20 | 1.4 | 19.737 | 1.4209
| ACA0.00110.01 | ACA | 1 | 10.4 | 0.001 | A | mA | 0.01 | 2 | KHz | 20 | 1.4208 | 19.745 | 1.4206
| ACA0.0110.01 | ACA | 1 | 10.4 | 0.01 | A | mA | 0.01 | 2 | KHz | 20.518 | 1.3907 | 20.514 | 1.3899
| ACA0.110.01 | ACA | 1 | 10.4 | 0.1 | A | mA | 0.01 | 2 | KHz | 60.821 | 0.63861 | 60.818 | 0.63675
| ACA110.01 | ACA | 1 | 10.4 | 1 | A | mA | 0.01 | 2 | KHz | 579.36 | 0.35001 | 579.36 | 0.34913
| Capacitance0.0000000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-09 | mF | nF |  |  |  | 0.0066 | 8700 | 0.0065002 | 8665.6
| Capacitance0.000000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-08 | mF | nF |  |  |  | 0.0076851 | 8569 | 0.0076487 | 8568.6
| Capacitance0.00000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-07 | mF | nF |  |  |  | 0.05416 | 5571.2 | 0.054131 | 5560.4
| Capacitance0.0000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-06 | mF | nF |  |  |  | 0.5773 | 1107.3 | 0.5773 | 1097.8
| Capacitance0.000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-05 | mF | nF |  |  |  | 5.7901 | 3024.4 | 5.7901 | 3020.5
| Capacitance0.000000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-08 | mF | nF |  |  |  | 0.05 | 8800 | 0.048294 | 8809
| Capacitance0.00000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-07 | mF | nF |  |  |  | 0.061389 | 8700 | 0.060987 | 8700
| Capacitance0.0000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-06 | mF | nF |  |  |  | 0.53772 | 5590 | 0.53742 | 5579.2
| Capacitance0.000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-05 | mF | nF |  |  |  | 5.7726 | 1107.4 | 5.7727 | 1097.8
| Capacitance0.00010.00001 | Capacitance | 1e-05 | 0.0001 | 0.0001 | mF | nF |  |  |  | 57.901 | 3024.4 | 57.901 | 3020.5
| Capacitance0.00000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-07 | mF | nF |  |  |  | 0.5 | 8800 | 0.48294 | 8809
| Capacitance0.0000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-06 | mF | nF |  |  |  | 0.61389 | 8700 | 0.60987 | 8700
| Capacitance0.000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-05 | mF | nF |  |  |  | 5.3772 | 5590 | 5.3742 | 5579.2
| Capacitance0.00010.0001 | Capacitance | 0.0001 | 0.001 | 0.0001 | mF | nF |  |  |  | 57.726 | 1107.4 | 57.727 | 1097.8
| Capacitance0.0010.0001 | Capacitance | 0.0001 | 0.001 | 0.001 | mF | nF |  |  |  | 579.01 | 3024.4 | 579.01 | 3020.5
| Frequency0.0010.001 | Frequency | 0.001 | 100 | 0.001 | kHz | Hz |  |  |  | 0.87 | 0.016 | 0.81638 | 0.016364
| Frequency0.010.001 | Frequency | 0.001 | 100 | 0.01 | kHz | Hz |  |  |  | 5.8379 | 0.0048237 | 5.8243 | 0.0046826
| Frequency0.10.001 | Frequency | 0.001 | 100 | 0.1 | kHz | Hz |  |  |  | 58.005 | 0.0031699 | 58 | 0.0031065
| Frequency10.001 | Frequency | 0.001 | 100 | 1 | kHz | Hz |  |  |  | 579.98 | 0.026531 | 579.97 | 0.026505
| Frequency100.001 | Frequency | 0.001 | 100 | 10 | kHz | Hz |  |  |  | 5799.7 | 0.26461 | 5799.7 | 0.2646
| Insulation DCV11 | Insulation DCV | 1 | 100 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.4054 | 0.010453
| Insulation DCV101 | Insulation DCV | 1 | 100 | 10 | V | V |  |  |  | 5.9149 | 0.0038256 | 5.9069 | 0.0036668
| Insulation DCV1001 | Insulation DCV | 1 | 100 | 100 | V | V |  |  |  | 57.918 | 0.0033656 | 57.915 | 0.0032952
| Insulation DCV1100 | Insulation DCV | 100 | 250 | 1 | V | V |  |  |  | 1.4 | 0.011 | 1.3333 | 0.011162
| Insulation DCV10100 | Insulation DCV | 100 | 250 | 10 | V | V |  |  |  | 5.6899 | 0.0057678 | 5.6798 | 0.0056527
| Insulation DCV100100 | Insulation DCV | 100 | 250 | 100 | V | V |  |  |  | 57.719 | 0.002458 | 57.716 | 0.0023936
| Insulation DCV1250 | Insulation DCV | 250 | 500 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.2696 | 0.011414
| Insulation DCV10250 | Insulation DCV | 250 | 500 | 10 | V | V |  |  |  | 5.1032 | 0.0080276 | 5.0881 | 0.007932
| Insulation DCV100250 | Insulation DCV | 250 | 500 | 100 | V | V |  |  |  | 57.497 | 0.0022225 | 57.502 | 0.0021425
| Insulation DCV 1500 | Insulation DCV | 500 | 1000 | 1 | V | V |  |  |  | 1.8 | 0.011 | 1.2242 | 0.011504
| Insulation DCV 10500 | Insulation DCV | 500 | 1000 | 10 | V | V |  |  |  | 4.1693 | 0.0098749 | 4.0987 | 0.0098767
| Insulation DCV 100500 | Insulation DCV | 500 | 1000 | 100 | V | V |  |  |  | 57.066 | 0.0024926 | 57.041 | 0.0024497
| Insulation Res10250 | Insulation Res | 250 | 100000 | 10 | kΩ | kΩ |  |  |  | 18 | 0.0092 | 13.773 | 0.0092419
| Insulation Res100250 | Insulation Res | 250 | 100000 | 100 | kΩ | kΩ |  |  |  | 57.523 | 0.0088223 | 57.453 | 0.0088228
| Insulation Res1000250 | Insulation Res | 250 | 100000 | 1000 | kΩ | kΩ |  |  |  | 576.26 | 0.0052659 | 576.24 | 0.0052655
| Insulation Res10000250 | Insulation Res | 250 | 100000 | 10000 | kΩ | kΩ |  |  |  | 5775 | 0.0010208 | 5775 | 0.0010204
| Insulation Res10100000 | Insulation Res | 100000 | 250000 | 10 | kΩ | kΩ |  |  |  | 14 | 0.011 | 13.511 | 0.010788
| Insulation Res100100000 | Insulation Res | 100000 | 250000 | 100 | kΩ | kΩ |  |  |  | 57.138 | 0.010616 | 57.125 | 0.010616
| Insulation Res1000100000 | Insulation Res | 100000 | 250000 | 1000 | kΩ | kΩ |  |  |  | 575.4 | 0.0087861 | 575.35 | 0.0087861
| Insulation Res10000100000 | Insulation Res | 100000 | 250000 | 10000 | kΩ | kΩ |  |  |  | 5773 | 0.002517 | 5772.9 | 0.0025167
| Insulation Res100000100000 | Insulation Res | 100000 | 250000 | 100000 | kΩ | kΩ |  |  |  | 57896 | 0.0013129 | 57896 | 0.0013128
| Insulation Res10250000 | Insulation Res | 250000 | 500000 | 10 | kΩ | kΩ |  |  |  | 18 | 0.011 | 12.764 | 0.011009
| Insulation Res100250000 | Insulation Res | 250000 | 500000 | 100 | kΩ | kΩ |  |  |  | 55.079 | 0.010925 | 54.77 | 0.010926
| Insulation Res1000250000 | Insulation Res | 250000 | 500000 | 1000 | kΩ | kΩ |  |  |  | 572.67 | 0.0099501 | 572.64 | 0.0099501
| Insulation Res10000250000 | Insulation Res | 250000 | 500000 | 10000 | kΩ | kΩ |  |  |  | 5771.4 | 0.0044676 | 5771.3 | 0.0044674
| Insulation Res10500000 | Insulation Res | 500000 | 1000000 | 10 | kΩ | kΩ |  |  |  | 25 | 0.011 | -76.51 | 0.011211
| Insulation Res100500000 | Insulation Res | 500000 | 1000000 | 100 | kΩ | kΩ |  |  |  | 25 | 0.011209 | -75.391 | 0.011209
| Insulation Res1000500000 | Insulation Res | 500000 | 1000000 | 1000 | kΩ | kΩ |  |  |  | 25 | 0.011045 | 24.396 | 0.011045
| Insulation Res10000500000 | Insulation Res | 500000 | 1000000 | 10000 | kΩ | kΩ |  |  |  | 5114.5 | 0.007343 | 5114.5 | 0.0073429
| Insulation Res100000500000 | Insulation Res | 500000 | 1000000 | 100000 | kΩ | kΩ |  |  |  | 57638 | 0.0014395 | 57638 | 0.0014395
| TcindicatorJ0.1-346 | TcindicatorJ | -346 | -148 | 0.1 | °F | °F |  |  |  | 2.4168 | 0 | 2.4145 | 0
| TcindicatorJ0.1-148 | TcindicatorJ | -148 | 302 | 0.1 | °F | °F |  |  |  | 1.9165 | 0 | 1.9145 | 0
| TcindicatorJ0.1302 | TcindicatorJ | 302 | 1400 | 0.1 | °F | °F |  |  |  | 2.0393 | 0 | 2.0387 | 0
| TcindicatorJ0.11400 | TcindicatorJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 2.2469 | 0 | 2.2464 | 0
| TcindicatorJ1-346 | TcindicatorJ | -346 | -148 | 1 | °F | °F |  |  |  | 2.4908 | 0 | 2.4819 | 0
| TcindicatorJ1-148 | TcindicatorJ | -148 | 302 | 1 | °F | °F |  |  |  | 2.0067 | 0 | 1.9988 | 0
| TcindicatorJ1302 | TcindicatorJ | 302 | 1400 | 1 | °F | °F |  |  |  | 2.1202 | 0 | 2.1181 | 0
| TcindicatorJ11400 | TcindicatorJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 2.3207 | 0 | 2.3187 | 0
| TcindicatorJ0.1-210 | TcindicatorJ | -210 | -100 | 0.1 | °C | °C |  |  |  | 1.3461 | 0 | 1.3448 | 0
| TcindicatorJ0.1-100 | TcindicatorJ | -100 | 150 | 0.1 | °C | °C |  |  |  | 1.0665 | 0 | 1.0653 | 0
| TcindicatorJ0.1150 | TcindicatorJ | 150 | 760 | 0.1 | °C | °C |  |  |  | 1.1342 | 0 | 1.1339 | 0
| TcindicatorJ0.1760 | TcindicatorJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 1.2495 | 0 | 1.2492 | 0
| TcindicatorJ1-210 | TcindicatorJ | -210 | -100 | 1 | °C | °C |  |  |  | 1.467 | 0 | 1.4624 | 0
| TcindicatorJ1-100 | TcindicatorJ | -100 | 150 | 1 | °C | °C |  |  |  | 1.2144 | 0 | 1.2103 | 0
| TcindicatorJ1150 | TcindicatorJ | 150 | 760 | 1 | °C | °C |  |  |  | 1.2722 | 0 | 1.2711 | 0
| TcindicatorJ1760 | TcindicatorJ | 760 | 1200 | 1 | °C | °C |  |  |  | 1.376 | 0 | 1.375 | 0
| TcindicatorK0.1-328 | TcindicatorK | -328 | -148 | 0.1 | °F | °F |  |  |  | 2.5841 | 0 | 2.5813 | 0
| TcindicatorK0.1-148 | TcindicatorK | -148 | 248 | 0.1 | °F | °F |  |  |  | 2.0844 | 0 | 2.0809 | 0
| TcindicatorK0.1248 | TcindicatorK | 248 | 2500 | 0.1 | °F | °F |  |  |  | 2.5391 | 0 | 2.5376 | 0
| TcindicatorK1-328 | TcindicatorK | -328 | -148 | 1 | °F | °F |  |  |  | 2.6555 | 0 | 2.6444 | 0
| TcindicatorK1-148 | TcindicatorK | -148 | 248 | 1 | °F | °F |  |  |  | 2.1723 | 0 | 2.1587 | 0
| TcindicatorK1248 | TcindicatorK | 248 | 2500 | 1 | °F | °F |  |  |  | 2.6075 | 0 | 2.6018 | 0
| TcindicatorK0.1-200 | TcindicatorK | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.4398 | 0 | 1.4382 | 0
| TcindicatorK0.1-100 | TcindicatorK | -100 | 120 | 0.1 | °C | °C |  |  |  | 1.16 | 0 | 1.1581 | 0
| TcindicatorK0.1120 | TcindicatorK | 120 | 1370 | 0.1 | °C | °C |  |  |  | 1.4121 | 0 | 1.4113 | 0
| TcindicatorK1-200 | TcindicatorK | -200 | -100 | 1 | °C | °C |  |  |  | 1.5545 | 0 | 1.5487 | 0
| TcindicatorK1-100 | TcindicatorK | -100 | 120 | 1 | °C | °C |  |  |  | 1.2998 | 0 | 1.2927 | 0
| TcindicatorK1120 | TcindicatorK | 120 | 1370 | 1 | °C | °C |  |  |  | 1.5267 | 0 | 1.5237 | 0
| TcindicatorN0.1-328 | TcindicatorN | -328 | -148 | 0.1 | °F | °F |  |  |  | 3.205 | 0 | 3.2025 | 0
| TcindicatorN0.1-148 | TcindicatorN | -148 | 770 | 0.1 | °F | °F |  |  |  | 2.2918 | 0 | 2.2894 | 0
| TcindicatorN0.1770 | TcindicatorN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 2.4558 | 0 | 2.4545 | 0
| TcindicatorN1-328 | TcindicatorN | -328 | -148 | 1 | °F | °F |  |  |  | 3.2633 | 0 | 3.2536 | 0
| TcindicatorN1-148 | TcindicatorN | -148 | 770 | 1 | °F | °F |  |  |  | 2.3698 | 0 | 2.3604 | 0
| TcindicatorN1770 | TcindicatorN | 770 | 2372 | 1 | °F | °F |  |  |  | 2.5259 | 0 | 2.5208 | 0
| TcindicatorN0.1-200 | TcindicatorN | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.782 | 0 | 1.7806 | 0
| TcindicatorN0.1-100 | TcindicatorN | -100 | 410 | 0.1 | °C | °C |  |  |  | 1.2762 | 0 | 1.2749 | 0
| TcindicatorN0.1410 | TcindicatorN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 1.3658 | 0 | 1.3651 | 0
| TcindicatorN1-200 | TcindicatorN | -200 | -100 | 1 | °C | °C |  |  |  | 1.8762 | 0 | 1.871 | 0
| TcindicatorN1-100 | TcindicatorN | -100 | 410 | 1 | °C | °C |  |  |  | 1.4032 | 0 | 1.3983 | 0
| TcindicatorN1410 | TcindicatorN | 410 | 1300 | 1 | °C | °C |  |  |  | 1.4837 | 0 | 1.481 | 0
| TcindicatorR0.132 | TcindicatorR | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
| TcindicatorR0.1482 | TcindicatorR | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5823 | 0 | 3.5805 | 0
| TcindicatorR132 | TcindicatorR | 32 | 482 | 1 | °F | °F |  |  |  | 4.8335 | 0 | 4.8271 | 0
| TcindicatorR1482 | TcindicatorR | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6336 | 0 | 3.6263 | 0
| TcindicatorR0.10 | TcindicatorR | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6784 | 0 | 2.6775 | 0
| TcindicatorR0.1250 | TcindicatorR | 250 | 1760 | 0.1 | °C | °C |  |  |  | 1.9965 | 0 | 1.9955 | 0
| TcindicatorR10 | TcindicatorR | 0 | 250 | 1 | °C | °C |  |  |  | 2.7419 | 0 | 2.7384 | 0
| TcindicatorR1250 | TcindicatorR | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0804 | 0 | 2.0765 | 0
| TcindicatorS0.132 | TcindicatorS | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
| TcindicatorS0.1482 | TcindicatorS | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5846 | 0 | 3.5828 | 0
| TcindicatorS132 | TcindicatorS | 32 | 482 | 1 | °F | °F |  |  |  | 4.8334 | 0 | 4.8271 | 0
| TcindicatorS1482 | TcindicatorS | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6358 | 0 | 3.6285 | 0
| TcindicatorS0.10 | TcindicatorS | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6783 | 0 | 2.6775 | 0
| TcindicatorS0.1250 | TcindicatorS | 250 | 1760 | 0.1 | °C | °C |  |  |  | 2.0006 | 0 | 1.9996 | 0
| TcindicatorS10 | TcindicatorS | 0 | 250 | 1 | °C | °C |  |  |  | 2.7418 | 0 | 2.7384 | 0
| TcindicatorS1250 | TcindicatorS | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0844 | 0 | 2.0805 | 0
| TcindicatorT0.1-418 | TcindicatorT | -418 | -238 | 0.1 | °F | °F |  |  |  | 3.9624 | 0 | 3.9602 | 0
| TcindicatorT0.1-238 | TcindicatorT | -238 | 752 | 0.1 | °F | °F |  |  |  | 1.9193 | 0 | 1.9156 | 0
| TcindicatorT1-418 | TcindicatorT | -418 | -238 | 1 | °F | °F |  |  |  | 4.0103 | 0 | 4.0017 | 0
| TcindicatorT1-238 | TcindicatorT | -238 | 752 | 1 | °F | °F |  |  |  | 2.0141 | 0 | 1.9999 | 0
| TcindicatorT0.1-250 | TcindicatorT | -250 | -150 | 0.1 | °C | °C |  |  |  | 2.2146 | 0 | 2.2134 | 0
| TcindicatorT0.1-150 | TcindicatorT | -150 | 400 | 0.1 | °C | °C |  |  |  | 1.0694 | 0 | 1.0674 | 0
| TcindicatorT1-250 | TcindicatorT | -250 | -150 | 1 | °C | °C |  |  |  | 2.2914 | 0 | 2.2868 | 0
| TcindicatorT1-150 | TcindicatorT | -150 | 400 | 1 | °C | °C |  |  |  | 1.2194 | 0 | 1.2121 | 0

## Software Validation
Date: | Action
2015-12-02 | Validated all functions and calculations, repeat readings, uncertainty derivations from table
2017-02-17 | Changed calculations so that any reading with no uncertainty is marked with an asterisk
2018-04-08 | Reordered standard 1 values to fix uncertainty calcualtor errors
2019-05-15 | Added Standard 2, HP 3458A uncertainties
2019-05-30 | Fixed lookup functions so that ACA works properly
2019-06-04 | Fixed error in lookup function
2019-06-04 | Changed 3458A uncertainties to Ω instead of Ohms
2019-07-03 | Added Standards 3 and 4, HP 34420A and Transmille uncertainties