## Unc. Lookup (2)
Standard | Item | Range1_Low | Range1_High | Res | Meas_Units | Uncert_Units | Range2_Low | Range2_High | Range2_Units | Base | Mult | Base2 | Mult2 | FDE | E | E | 0
3458A | FDE | -454 | 300 | 1 | °F | °F |  |  |  | 0.60177 | 0 | 0.58066 | 0 | FDK | K | K | E
3458A | FDE | -454 | 300 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.08462 | 0 | FDN | N | N | J
3458A | FDK | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.99057 | 0 | 0.95529 | 0 | FDT | T | T | K
3458A | FDK | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.77443 | 0 | 0.76327 | 0 | TcindicatorT | T | J | N
3458A | FDN | -454 | -400.001 | 1 | °F | °F |  |  |  | 1.5603 | 0 | 1.54 | 0 | FDJ | J | 0 | R
3458A | FDN | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 1.5 | 0 | 1.4288 | 0 | TcindicatorJ | J | R | S
3458A | FDT | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.74355 | 0 | 0.69586 | 0 | FDPT100 | 0 | S | T
3458A | FDT | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.41398 | 0 | 0.39271 | 0 | FDPT1000 | 0
5502A | FDE | -418 | -147.9 | 0.1 | °F | °F |  |  |  | 1.0462 | 0 | 1.0433 | 0 | TcindicatorK | K
5502A | FDE | -418 | -147.9 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0 | TcindicatorN | N
5502A | FDT | -418 | -237.9 | 0.1 | °F | °F |  |  |  | 0.55536 | 0 | 0.55058 | 0 | CDE | E
5502A | FDT | -418 | -237.9 | 1 | °F | °F |  |  |  | 0.8089 | 0 | 0.7957 | 0 | CDK | K
Transmille | TcindicatorT | -418 | -238 | 0.1 | °F | °F |  |  |  | 3.9624 | 0 | 3.9602 | 0 | CDN | N
Transmille | TcindicatorT | -418 | -238 | 1 | °F | °F |  |  |  | 4.0103 | 0 | 4.0017 | 0 | CDT | T
3458A | FDK | -400 | 499.999 | 1 | °F | °F |  |  |  | 0.63828 | 0 | 0.59031 | 0 | CDJ | J
3458A | FDK | -400 | 499.999 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13588 | 0 | CDPT100 | 0
3458A | FDN | -400 | -150.001 | 1 | °F | °F |  |  |  | 0.65772 | 0 | 0.60786 | 0 | CDPT1000 | 0
3458A | FDN | -400 | -150.001 | 0.1 | °F | °F |  |  |  | 0.24 | 0 | 0.19873 | 0 | FDR | R
3458A | FDT | -400 | 199.999 | 1 | °F | °F |  |  |  | 0.63425 | 0 | 0.58771 | 0 | FDS | S
3458A | FDT | -400 | 199.999 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12412 | 0 | CDR | R
5502A | FDJ | -346 | -147.9 | 0.1 | °F | °F |  |  |  | 0.57727 | 0 | 0.56758 | 0 | CDS | S
5502A | FDJ | -346 | -147.9 | 1 | °F | °F |  |  |  | 0.83458 | 0 | 0.80755 | 0
3458A | FDJ | -346 | 499.999 | 1 | °F | °F |  |  |  | 0.61872 | 0 | 0.58175 | 0
3458A | FDJ | -346 | 499.999 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.091814 | 0
Transmille | TcindicatorJ | -346 | -148 | 0.1 | °F | °F |  |  |  | 2.4168 | 0 | 2.4145 | 0
Transmille | TcindicatorJ | -346 | -148 | 1 | °F | °F |  |  |  | 2.4908 | 0 | 2.4819 | 0
5502A | FDK | -328 | -147.9 | 0.1 | °F | °F |  |  |  | 0.70186 | 0 | 0.69128 | 0
5502A | FDK | -328 | -147.9 | 1 | °F | °F |  |  |  | 0.93103 | 0 | 0.89881 | 0
5502A | FDN | -328 | -147.9 | 0.1 | °F | °F |  |  |  | 0.84492 | 0 | 0.83553 | 0
5502A | FDN | -328 | -147.9 | 1 | °F | °F |  |  |  | 1.0446 | 0 | 1.014 | 0
5502A | FDPT100 | -328 | 32.1 | 0.1 | °F | °F |  |  |  | 0.21928 | 0 | 0.19044 | 0
5502A | FDPT100 | -328 | 32.1 | 1 | °F | °F |  |  |  | 0.60123 | 0 | 0.58897 | 0
5502A | FDPT1000 | -328 | 32.1 | 0.1 | °F | °F |  |  |  | 0.16401 | 0 | 0.1228 | 0
5502A | FDPT1000 | -328 | 32.1 | 1 | °F | °F |  |  |  | 0.59551 | 0 | 0.58313 | 0
Transmille | TcindicatorK | -328 | -148 | 0.1 | °F | °F |  |  |  | 2.5841 | 0 | 2.5813 | 0
Transmille | TcindicatorK | -328 | -148 | 1 | °F | °F |  |  |  | 2.6555 | 0 | 2.6444 | 0
Transmille | TcindicatorN | -328 | -148 | 0.1 | °F | °F |  |  |  | 3.205 | 0 | 3.2025 | 0
Transmille | TcindicatorN | -328 | -148 | 1 | °F | °F |  |  |  | 3.2633 | 0 | 3.2536 | 0
3458A | FDPT100 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.013 | 0 | 0.010679 | 0
3458A | FDPT100 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57804 | 0 | 0.57742 | 0
3458A | FDPT100 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.059936 | 0 | 0.05843 | 0
3458A | FDPT1000 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.0090498 | 0 | 0.0080621 | 0
3458A | FDPT1000 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57761 | 0 | 0.57738 | 0
3458A | FDPT1000 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.058588 | 0 | 0.058009 | 0
3458A | CDE | -270 | 149 | 1 | °C | °C |  |  |  | 0.58502 | 0 | 0.57839 | 0
3458A | CDE | -270 | 149 | 0.1 | °C | °C |  |  |  | 0.080416 | 0 | 0.067383 | 0
3458A | CDK | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.74458 | 0 | 0.73021 | 0
3458A | CDK | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.45663 | 0 | 0.45079 | 0
3458A | CDN | -270 | -240.001 | 1 | °C | °C |  |  |  | 1.018 | 0 | 1.0084 | 0
3458A | CDN | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.8317 | 0 | 0.82877 | 0
3458A | CDT | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.63746 | 0 | 0.62062 | 0
3458A | CDT | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.2459 | 0 | 0.23488 | 0
5502A | CDE | -250 | -99.9 | 0.1 | °C | °C |  |  |  | 0.58423 | 0 | 0.58258 | 0
5502A | CDE | -250 | -99.9 | 1 | °C | °C |  |  |  | 0.82286 | 0 | 0.81817 | 0
5502A | CDT | -250 | -149.9 | 0.1 | °C | °C |  |  |  | 0.55536 | 0 | 0.55058 | 0
5502A | CDT | -250 | -149.9 | 1 | °C | °C |  |  |  | 0.8089 | 0 | 0.7957 | 0
Transmille | TcindicatorT | -250 | -150 | 0.1 | °C | °C |  |  |  | 2.2146 | 0 | 2.2134 | 0
Transmille | TcindicatorT | -250 | -150 | 1 | °C | °C |  |  |  | 2.2914 | 0 | 2.2868 | 0
3458A | CDK | -240 | 259.999 | 1 | °C | °C |  |  |  | 0.59718 | 0 | 0.58175 | 0
3458A | CDK | -240 | 259.999 | 0.1 | °C | °C |  |  |  | 0.11392 | 0 | 0.091809 | 0
3458A | CDN | -240 | -101.001 | 1 | °C | °C |  |  |  | 0.60402 | 0 | 0.58768 | 0
3458A | CDN | -240 | -101.001 | 0.1 | °C | °C |  |  |  | 0.14224 | 0 | 0.12394 | 0
3458A | CDT | -240 | 92.999 | 1 | °C | °C |  |  |  | 0.59575 | 0 | 0.58083 | 0
3458A | CDT | -240 | 92.999 | 0.1 | °C | °C |  |  |  | 0.10841 | 0 | 0.08582 | 0
5502A | FDT | -238 | 32.1 | 0.1 | °F | °F |  |  |  | 0.29308 | 0 | 0.28549 | 0
5502A | FDT | -238 | 32.1 | 1 | °F | °F |  |  |  | 0.65502 | 0 | 0.64149 | 0
Transmille | TcindicatorT | -238 | 752 | 0.1 | °F | °F |  |  |  | 1.9193 | 0 | 1.9156 | 0
Transmille | TcindicatorT | -238 | 752 | 1 | °F | °F |  |  |  | 2.0141 | 0 | 1.9999 | 0
5502A | CDJ | -210 | -99.9 | 0.1 | °C | °C |  |  |  | 0.32482 | 0 | 0.31951 | 0
5502A | CDJ | -210 | -99.9 | 1 | °C | °C |  |  |  | 0.66767 | 0 | 0.65733 | 0
3458A | CDJ | -210 | 259.999 | 1 | °C | °C |  |  |  | 0.59049 | 0 | 0.57878 | 0
3458A | CDJ | -210 | 259.999 | 0.1 | °C | °C |  |  |  | 0.0917 | 0 | 0.070601 | 0
Transmille | TcindicatorJ | -210 | -100 | 0.1 | °C | °C |  |  |  | 1.3461 | 0 | 1.3448 | 0
Transmille | TcindicatorJ | -210 | -100 | 1 | °C | °C |  |  |  | 1.467 | 0 | 1.4624 | 0
5502A | CDK | -200 | -99.9 | 0.1 | °C | °C |  |  |  | 0.39349 | 0 | 0.38767 | 0
5502A | CDK | -200 | -99.9 | 1 | °C | °C |  |  |  | 0.70603 | 0 | 0.69303 | 0
5502A | CDN | -200 | -99.9 | 0.1 | °C | °C |  |  |  | 0.47201 | 0 | 0.46682 | 0
5502A | CDN | -200 | -99.9 | 1 | °C | °C |  |  |  | 0.75326 | 0 | 0.74022 | 0
5502A | CDPT100 | -200 | 0.1 | 0.1 | °C | °C |  |  |  | 0.13102 | 0 | 0.11627 | 0
5502A | CDPT100 | -200 | 0.1 | 1 | °C | °C |  |  |  | 0.60095 | 0 | 0.58869 | 0
5502A | CDPT1000 | -200 | 0.1 | 0.1 | °C | °C |  |  |  | 0.10309 | 0 | 0.083548 | 0
5502A | CDPT1000 | -200 | 0.1 | 1 | °C | °C |  |  |  | 0.59549 | 0 | 0.58311 | 0
Transmille | TcindicatorK | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.4398 | 0 | 1.4382 | 0
Transmille | TcindicatorK | -200 | -100 | 1 | °C | °C |  |  |  | 1.5545 | 0 | 1.5487 | 0
Transmille | TcindicatorN | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.782 | 0 | 1.7806 | 0
Transmille | TcindicatorN | -200 | -100 | 1 | °C | °C |  |  |  | 1.8762 | 0 | 1.871 | 0
3458A | CDPT100 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.008485 | 0 | 0.0076317 | 0
3458A | CDPT100 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57756 | 0 | 0.57737 | 0
3458A | CDPT100 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058423 | 0 | 0.05795 | 0
3458A | CDPT1000 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.0069515 | 0 | 0.0065655 | 0
3458A | CDPT1000 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57743 | 0 | 0.57736 | 0
3458A | CDPT1000 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058 | 0 | 0.05782 | 0
5502A | CDT | -150 | 0.1 | 0.1 | °C | °C |  |  |  | 0.29308 | 0 | 0.28549 | 0
5502A | CDT | -150 | 0.1 | 1 | °C | °C |  |  |  | 0.65502 | 0 | 0.64149 | 0
3458A | FDN | -150 | 700 | 1 | °F | °F |  |  |  | 0.61817 | 0 | 0.58117 | 0
3458A | FDN | -150 | 700 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.088083 | 0
Transmille | TcindicatorT | -150 | 400 | 0.1 | °C | °C |  |  |  | 1.0694 | 0 | 1.0674 | 0
Transmille | TcindicatorT | -150 | 400 | 1 | °C | °C |  |  |  | 1.2194 | 0 | 1.2121 | 0
5502A | FDE | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.3466 | 0 | 0.34484 | 0
5502A | FDE | -148 | -12.9 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
5502A | FDJ | -148 | -21.9 | 0.1 | °F | °F |  |  |  | 0.35408 | 0 | 0.34263 | 0
5502A | FDJ | -148 | -21.9 | 1 | °F | °F |  |  |  | 0.69231 | 0 | 0.66888 | 0
5502A | FDK | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.40198 | 0 | 0.38321 | 0
5502A | FDK | -148 | -12.9 | 1 | °F | °F |  |  |  | 0.73198 | 0 | 0.69054 | 0
5502A | FDN | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.47674 | 0 | 0.46496 | 0
5502A | FDN | -148 | -12.9 | 1 | °F | °F |  |  |  | 0.76848 | 0 | 0.73904 | 0
Transmille | TcindicatorJ | -148 | 302 | 0.1 | °F | °F |  |  |  | 1.9165 | 0 | 1.9145 | 0
Transmille | TcindicatorJ | -148 | 302 | 1 | °F | °F |  |  |  | 2.0067 | 0 | 1.9988 | 0
Transmille | TcindicatorK | -148 | 248 | 0.1 | °F | °F |  |  |  | 2.0844 | 0 | 2.0809 | 0
Transmille | TcindicatorK | -148 | 248 | 1 | °F | °F |  |  |  | 2.1723 | 0 | 2.1587 | 0
Transmille | TcindicatorN | -148 | 770 | 0.1 | °F | °F |  |  |  | 2.2918 | 0 | 2.2894 | 0
Transmille | TcindicatorN | -148 | 770 | 1 | °F | °F |  |  |  | 2.3698 | 0 | 2.3604 | 0
3458A | CDN | -101 | 371 | 1 | °C | °C |  |  |  | 0.59027 | 0 | 0.57855 | 0
3458A | CDN | -101 | 371 | 0.1 | °C | °C |  |  |  | 0.090241 | 0 | 0.068695 | 0
5502A | CDE | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.19859 | 0 | 0.19765 | 0
5502A | CDE | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.60874 | 0 | 0.60751 | 0
5502A | CDJ | -100 | -29.9 | 0.1 | °C | °C |  |  |  | 0.20266 | 0 | 0.19649 | 0
5502A | CDJ | -100 | -29.9 | 1 | °C | °C |  |  |  | 0.61519 | 0 | 0.60713 | 0
5502A | CDK | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.22869 | 0 | 0.21853 | 0
5502A | CDK | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.62924 | 0 | 0.61462 | 0
5502A | CDN | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.26968 | 0 | 0.26325 | 0
5502A | CDN | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.64265 | 0 | 0.6319 | 0
Transmille | TcindicatorJ | -100 | 150 | 0.1 | °C | °C |  |  |  | 1.0665 | 0 | 1.0653 | 0
Transmille | TcindicatorJ | -100 | 150 | 1 | °C | °C |  |  |  | 1.2144 | 0 | 1.2103 | 0
Transmille | TcindicatorK | -100 | 120 | 0.1 | °C | °C |  |  |  | 1.16 | 0 | 1.1581 | 0
Transmille | TcindicatorK | -100 | 120 | 1 | °C | °C |  |  |  | 1.2998 | 0 | 1.2927 | 0
Transmille | TcindicatorN | -100 | 410 | 0.1 | °C | °C |  |  |  | 1.2762 | 0 | 1.2749 | 0
Transmille | TcindicatorN | -100 | 410 | 1 | °C | °C |  |  |  | 1.4032 | 0 | 1.3983 | 0
3458A | FDR | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.66465 | 0 | 0.61591 | 0
3458A | FDR | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.26 | 0 | 0.22215 | 0
3458A | FDS | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.6516 | 0 | 0.6018 | 0
3458A | FDS | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.22 | 0 | 0.17936 | 0
3458A | CDR | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60597 | 0 | 0.58986 | 0
3458A | CDR | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.15085 | 0 | 0.13394 | 0
3458A | CDS | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60145 | 0 | 0.58521 | 0
3458A | CDS | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.13149 | 0 | 0.11169 | 0
5502A | CDJ | -30 | 150.1 | 0.1 | °C | °C |  |  |  | 0.18177 | 0 | 0.17487 | 0
5502A | CDJ | -30 | 150.1 | 1 | °C | °C |  |  |  | 0.60863 | 0 | 0.60048 | 0
5502A | CDE | -25 | 350.1 | 0.1 | °C | °C |  |  |  | 0.17588 | 0 | 0.17482 | 0
5502A | CDE | -25 | 350.1 | 1 | °C | °C |  |  |  | 0.60171 | 0 | 0.60047 | 0
5502A | CDK | -25 | 120.1 | 0.1 | °C | °C |  |  |  | 0.20428 | 0 | 0.19653 | 0
5502A | CDK | -25 | 120.1 | 1 | °C | °C |  |  |  | 0.61728 | 0 | 0.60715 | 0
5502A | CDN | -25 | 120.1 | 0.1 | °C | °C |  |  |  | 0.23706 | 0 | 0.22973 | 0
5502A | CDN | -25 | 120.1 | 1 | °C | °C |  |  |  | 0.62966 | 0 | 0.61869 | 0
5502A | FDJ | -22 | 302.1 | 0.1 | °F | °F |  |  |  | 0.31531 | 0 | 0.3024 | 0
5502A | FDJ | -22 | 302.1 | 1 | °F | °F |  |  |  | 0.67331 | 0 | 0.64919 | 0
5502A | FDE | -13 | 662.1 | 0.1 | °F | °F |  |  |  | 0.30437 | 0 | 0.30237 | 0
5502A | FDE | -13 | 662.1 | 1 | °F | °F |  |  |  | 0.6529 | 0 | 0.64917 | 0
5502A | FDK | -13 | 248.1 | 0.1 | °F | °F |  |  |  | 0.35702 | 0 | 0.34265 | 0
5502A | FDK | -13 | 248.1 | 1 | °F | °F |  |  |  | 0.69831 | 0 | 0.66889 | 0
5502A | FDN | -13 | 248.1 | 0.1 | °F | °F |  |  |  | 0.41715 | 0 | 0.40364 | 0
5502A | FDN | -13 | 248.1 | 1 | °F | °F |  |  |  | 0.73301 | 0 | 0.70208 | 0
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 0.13 | 2300 | 0.11565 | 2314.2
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 0.13 | 2300 | 0.12522 | 2290.9
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 0.56558 | 1474.9 | 0.56533 | 1474.5
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 5.7994 | 124.28 | 5.7993 | 123.73
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 58.025 | -857.45 | 58.025 | -857.67
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.01 | 0.02001 | kHz | 580.26 | -8787.4 | 580.03 | -8787.5
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.01 | 0.02001 | kHz | 5802.6 | -87896 | 5800 | -87896
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.01 | 0.02001 | kHz | 58026 | -878960 | 58000 | -878960
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 0.01 | 0.02001 | kHz | 580260 | -8789600 | 580000 | -8789600
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 0.01 | 0.02001 | kHz | 5802600 | -87896000 | 5800000 | -87896000
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 0.13 | 1700 | 0.11582 | 1731.6
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 0.13 | 1700 | 0.1262 | 1707.4
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 0.57237 | 986.73 | 0.57213 | 986.13
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 5.7723 | 128.02 | 5.7722 | 127.48
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 57.735 | 13.015 | 57.735 | 12.795
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.02 | 0.04501 | kHz | 577.36 | 1.3674 | 577.35 | 1.2795
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.02 | 0.04501 | kHz | 5773.6 | 0.16308 | 5773.6 | 0.12796
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.02 | 0.04501 | kHz | 57736 | 0.026847 | 57736 | 0.012796
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 0.02 | 0.04501 | kHz | 577360 | 0.0068998 | 577360 | 0.0012794
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 0.02 | 0.04501 | kHz | 5773600 | 0.0023794 | 5773600 | 0.00012995
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.14 | 1400 | 0.11585 | 1442.9
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.14 | 1400 | 0.12669 | 1418.5
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.5767 | 759.64 | 0.57647 | 758.92
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.7731 | 93.964 | 5.773 | 93.42
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.736 | 9.5875 | 57.735 | 9.3679
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.045 | 1 | kHz | 577.36 | 1.0246 | 577.36 | 0.93682
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.045 | 1 | kHz | 5773.6 | 0.12881 | 5773.6 | 0.093682
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.045 | 1 | kHz | 57736 | 0.02342 | 57736 | 0.0093682
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 0.045 | 1 | kHz | 577360 | 0.0065572 | 577360 | 0.00093674
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 0.045 | 1 | kHz | 5773600 | 0.0023422 | 5773600 | 9.5917e-05
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 0.18 | 3500 | 0.17344 | 3463.7
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 0.18 | 3500 | 0.17984 | 3448.1
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 0.56231 | 2652.9 | 0.56207 | 2652.8
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 5.7665 | 471.27 | 5.7664 | 470.74
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 57.735 | 47.918 | 57.735 | 47.699
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 1 | 5.01 | kHz | 577.36 | 4.8583 | 577.35 | 4.7705
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 1 | 5.01 | kHz | 5773.6 | 0.51218 | 5773.6 | 0.47705
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 1 | 5.01 | kHz | 57736 | 0.061756 | 57736 | 0.047705
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 1 | 5.01 | kHz | 577360 | 0.010391 | 577360 | 0.0047703
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 1 | 5.01 | kHz | 5773600 | 0.0027259 | 5773600 | 0.00047649
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 0.25 | 9200 | 0.23106 | 9237.4
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 0.25 | 9200 | 0.23463 | 9228.1
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 0.5163 | 8526.7 | 0.51576 | 8527.5
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 5.7139 | 2807.7 | 5.7137 | 2806.8
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 57.729 | 302.48 | 57.729 | 302.01
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 5 | 10.01 | kHz | 577.35 | 30.413 | 577.35 | 30.226
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 5 | 10.01 | kHz | 5773.6 | 3.0976 | 5773.5 | 3.0226
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 5 | 10.01 | kHz | 57736 | 0.33224 | 57736 | 0.30226
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 5 | 10.01 | kHz | 577360 | 0.04222 | 577360 | 0.030226
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 5 | 10.01 | kHz | 5773600 | 0.0078188 | 5773600 | 0.0030229
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46198 | 18475
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46389 | 18470
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 10 | 30 | kHz | 0.6298 | 18044 | 0.62958 | 18044
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 10 | 30 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 10 | 30 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 10 | 30 | kHz | 577.35 | 121.09 | 577.32 | 120.9
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 10 | 30 | kHz | 5773.6 | 12.165 | 5773.5 | 12.09
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 10 | 30 | kHz | 57736 | 1.239 | 57735 | 1.209
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 10 | 30 | kHz | 577360 | 0.1329 | 577360 | 0.1209
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 10 | 30 | kHz | 5773600 | 0.016888 | 5773600 | 0.012089
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 30 | 30.1 | kHz | 0.62 | 18000 | 0.46198 | 18475
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 30 | 30.1 | kHz | 0.62 | 18000 | 0.46389 | 18470
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 30 | 30.1 | kHz | 0.6298 | 18044 | 0.62958 | 18044
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 30 | 30.1 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 30 | 30.1 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 30 | 30.1 | kHz | 577.35 | 121.09 | 577.32 | 120.9
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 30 | 30.1 | kHz | 5773.6 | 12.165 | 5773.5 | 12.09
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 30 | 30.1 | kHz | 57736 | 1.239 | 57735 | 1.209
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 30 | 30.1 | kHz | 577360 | 0.1329 | 577360 | 0.1209
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 30 | 30.1 | kHz | 5773600 | 0.016888 | 5773600 | 0.012089
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 0.01 | 0.04501 | Hz | 25 | 1700 | 23.181 | 1732
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 0.01 | 0.04501 | Hz | 25 | 1718.2 | 23.854 | 1717.9
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 0.01 | 0.04501 | Hz | 62.01 | 1125.5 | 61.999 | 1123.8
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 0.01 | 0.04501 | Hz | 580.38 | 157.54 | 580.38 | 156.33
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 0.01 | 0.04501 | Hz | 5800 | 16.202 | 5800 | 15.714
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 0.01 | 0.04501 | Hz | 58000 | 1.7665 | 58000 | 1.5715
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 0.01 | 0.04501 | Hz | 580000 | 0.23514 | 580000 | 0.15715
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 0.01 | 0.04501 | Hz | 5800000 | 0.046912 | 5800000 | 0.015715
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 1200 | 23.185 | 1155.4
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 1143.4 | 23.87 | 1142.8
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 0.045 | 10 | Hz | 62.237 | 673.84 | 62.225 | 671.86
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 0.045 | 10 | Hz | 580.43 | 86.247 | 580.43 | 85.036
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800.1 | 9.0181 | 5800.1 | 8.5307
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 0.045 | 10 | Hz | 58000 | 1.0481 | 58000 | 0.8531
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 0.045 | 10 | Hz | 580000 | 0.1633 | 580000 | 0.08531
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 0.045 | 10 | Hz | 5800000 | 0.039728 | 5800000 | 0.008531
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 10 | 20.01 | kHz | 25 | 1700 | 23.181 | 1732
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 10 | 20.01 | kHz | 25 | 1718.2 | 23.854 | 1717.9
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 10 | 20.01 | kHz | 62.01 | 1125.5 | 61.999 | 1123.8
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 10 | 20.01 | kHz | 580.38 | 157.54 | 580.38 | 156.33
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 10 | 20.01 | kHz | 5800 | 16.202 | 5800 | 15.714
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 10 | 20.01 | kHz | 58000 | 1.7665 | 58000 | 1.5715
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 10 | 20.01 | kHz | 580000 | 0.23514 | 580000 | 0.15715
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 10 | 20.01 | kHz | 5800000 | 0.046912 | 5800000 | 0.015715
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 20 | 50.01 | kHz | 24 | 2300 | 23.565 | 2300.9
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 20 | 50.01 | kHz | 24.224 | 2286.5 | 24.216 | 2286.3
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 20 | 50.01 | kHz | 61.904 | 1615.7 | 61.89 | 1614.4
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 20 | 50.01 | kHz | 580.34 | 247.93 | 580.33 | 246.73
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 20 | 50.01 | kHz | 5800 | 25.351 | 5800 | 24.864
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 20 | 50.01 | kHz | 58000 | 2.6816 | 58000 | 2.4866
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 20 | 50.01 | kHz | 580000 | 0.32665 | 580000 | 0.24866
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 20 | 50.01 | kHz | 5800000 | 0.056063 | 5800000 | 0.024866
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 50 | 100 | kHz | 40 | 4000 | 38.937 | 4027
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 50 | 100.01 | kHz | 40 | 4017.5 | 39.332 | 4018
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 50 | 100.01 | kHz | 68.858 | 3411.4 | 68.766 | 3411.7
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 50 | 100.01 | kHz | 580.9 | 730.69 | 580.86 | 729
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 50 | 100.01 | kHz | 5800.1 | 75.275 | 5800.1 | 74.549
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 50 | 100.01 | kHz | 58000 | 7.7474 | 58000 | 7.4566
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 50 | 100.01 | kHz | 580000 | 0.86197 | 580000 | 0.74566
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 50 | 100.01 | kHz | 5800000 | 0.12109 | 5800000 | 0.074566
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 100 | 500 | kHz | 89 | 11000 | 72.86 | 11473
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 100 | 500 | kHz | 89 | 11468 | 73.063 | 11468
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 100 | 500 | kHz | 91.464 | 11024 | 91.322 | 11026
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 100 | 500 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 100 | 500 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 100 | 500 | kHz | 58000 | 53.659 | 58000 | 52.994
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 100 | 500 | kHz | 580000 | 5.5657 | 580000 | 5.2995
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 500 | 501 | kHz | 89 | 11000 | 72.86 | 11473
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 500 | 501 | kHz | 89 | 11468 | 73.063 | 11468
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 500 | 501 | kHz | 91.464 | 11024 | 91.322 | 11026
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 500 | 501 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 500 | 501 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 500 | 501 | kHz | 58000 | 53.659 | 58000 | 52.994
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 500 | 501 | kHz | 580000 | 5.5657 | 580000 | 5.2995
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 100 | 500.01 | kHz | 5800000 | 0.63642 | 5800000 | 0.52995
5502A | ACW | 0 | 1.09e-05 | 1e-10 | WAC | WAC |  |  |  | 0 | 0.025 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-09 | WAC | WAC |  |  |  | 0 | 0.025178 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.062989 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.58053 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-06 | WAC | WAC |  |  |  | 0 | 5.8001 | 0 | 0
5502A | CDPT100 | 0 | 100.1 | 0.1 | °C | °C |  |  |  | 0.16571 | 0 | 0.15431 | 0
5502A | CDPT100 | 0 | 100.1 | 1 | °C | °C |  |  |  | 0.60946 | 0 | 0.59736 | 0
5502A | CDPT1000 | 0 | 100.1 | 0.1 | °C | °C |  |  |  | 0.11707 | 0 | 0.10029 | 0
5502A | CDPT1000 | 0 | 100.1 | 1 | °C | °C |  |  |  | 0.59806 | 0 | 0.58574 | 0
5502A | CDR | 0 | 250.1 | 0.1 | °C | °C |  |  |  | 0.66953 | 0 | 0.66592 | 0
5502A | CDR | 0 | 250.1 | 1 | °C | °C |  |  |  | 0.89034 | 0 | 0.87946 | 0
5502A | CDS | 0 | 250.1 | 0.1 | °C | °C |  |  |  | 0.55632 | 0 | 0.55207 | 0
5502A | CDS | 0 | 250.1 | 1 | °C | °C |  |  |  | 0.80847 | 0 | 0.79673 | 0
5502A | CDT | 0 | 120.1 | 0.1 | °C | °C |  |  |  | 0.20015 | 0 | 0.19651 | 0
5502A | CDT | 0 | 120.1 | 1 | °C | °C |  |  |  | 0.61188 | 0 | 0.60714 | 0
5502A | DCA | 0 | 0.00033001 | 1e-09 | A | µA |  |  |  | 0.024 | 180 | 0.023107 | 175.06
5502A | DCA | 0 | 0.00033001 | 1e-08 | A | µA |  |  |  | 0.024 | 173.53 | 0.023814 | 173.54
5502A | DCA | 0 | 0.00033001 | 1e-07 | A | µA |  |  |  | 0.062453 | 112.75 | 0.062384 | 112.54
5502A | DCA | 0 | 0.00033001 | 1e-06 | A | µA |  |  |  | 0.58048 | 15.83 | 0.58045 | 15.637
5502A | DCA | 0 | 0.00033001 | 1e-05 | A | µA |  |  |  | 5.8001 | 1.6501 | 5.8001 | 1.5718
5502A | DCA | 0 | 0.00033001 | 0.0001 | A | µA |  |  |  | 58 | 0.18852 | 58 | 0.15719
5502A | DCA | 0 | 0.00033001 | 0.001 | A | µA |  |  |  | 580 | 0.028252 | 580 | 0.015719
5502A | DCA | 0 | 0.00033001 | 0.01 | A | mA |  |  |  | 5.8 | 6.59e-06 | 5.8 | 1.57e-06
5502A | DCA | 0 | 0.00033001 | 0.1 | A | mA |  |  |  | 58 | 2.16e-06 | 58 | 1.57e-07
5502A | DCV | 0 | 0.33 | 1e-07 | V | µV |  |  |  | 3.6 | 69 | 3.4727 | 69.314
5502A | DCV | 0 | 0.33 | 1e-06 | V | µV |  |  |  | 3.6 | 69.151 | 3.5196 | 69.191
5502A | DCV | 0 | 0.33 | 1e-05 | V | µV |  |  |  | 6.7738 | 61.287 | 6.7344 | 61.342
5502A | DCV | 0 | 0.33 | 0.0001 | V | µV |  |  |  | 58.11 | 17.06 | 58.091 | 17.007
5502A | DCV | 0 | 0.33 | 0.001 | V | µV |  |  |  | 580.02 | 1.8105 | 580.01 | 1.785
5502A | DCV | 0 | 0.33 | 0.01 | V | µV |  |  |  | 5800 | 0.1888 | 5800 | 0.17859
5502A | DCV | 0 | 0.33 | 0.1 | V | mV |  |  |  | 58 | 2.1943e-05 | 58 | 1.7859e-05
5502A | DCV | 0 | 0.33 | 1 | V | mV |  |  |  | 580 | 3.42e-06 | 580 | 1.79e-06
5502A | DCV | 0 | 0.33 | 10 | V | mV |  |  |  | 5800 | 8.32e-07 | 5800 | 1.79e-07
5502A | DCW | 0 | 0.000109 | 1e-10 | WDC | WDC |  |  |  | 0 | 0.0036 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-09 | WDC | WDC |  |  |  | 0 | 0.0036 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-08 | WDC | WDC |  |  |  | 0 | 0.0068024 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.05811 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.58001 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-05 | WDC | WDC |  |  |  | 0 | 5.8 | 0 | 0
5502A | DCW | 0 | 0.000109 | 0.0001 | WDC | WDC |  |  |  | 0 | 58 | 0 | 0
5502A | DCW | 0 | 0.000109 | 0.001 | WDC | WDC |  |  |  | 0 | 580 | 0 | 0
5502A | DCW | 0 | 0.000109 | 0.01 | WDC | WDC |  |  |  | 0 | 5800 | 0 | 0
5502A | Frequency | 0 | 120.001 | 0.001 | Hz | Hz |  |  |  | 0.006 | 1.3e-05 | 0.0012889 | 2.8049e-05
5502A | Frequency | 0 | 120.001 | 0.01 | Hz | Hz |  |  |  | 0.0060576 | 1.5807e-05 | 0.005881 | 1.2606e-05
5502A | Frequency | 0 | 120.001 | 0.1 | Hz | Hz |  |  |  | 0.057819 | 3.3129e-06 | 0.057746 | 1.4481e-06
5502A | Frequency | 0 | 120.001 | 1 | Hz | Hz |  |  |  | 0.57739 | 8.9636e-07 | 0.57736 | 1.4506e-07
5502A | Frequency | 0 | 120.001 | 10 | Hz | Hz |  |  |  | 5.7736 | 3.1508e-07 | 5.7736 | 1.4506e-08
5502A | Frequency | 0 | 120.001 | 100 | Hz | Hz |  |  |  | 57.736 | 1.2168e-07 | 57.736 | 1.45e-09
5502A | Frequency | 0 | 120.001 | 1000 | Hz | Hz |  |  |  | 577.36 | 4.8236e-08 | 577.36 | 1.45e-10
5502A | Resistance | 0 | 11.001 | 10000000 | Ω | mΩ |  |  |  | 5800000000 | 0 | 5800000000 | 0
5502A | Resistance | 0 | 11.001 | 1000000 | Ω | mΩ |  |  |  | 580000000 | 9.5149e-07 | 580000000 | 0
5502A | Resistance | 0 | 11.001 | 0.001 | Ω | mΩ |  |  |  | 12 | 0.14 | 11.576 | 0.1383
5502A | Resistance | 0 | 11.001 | 0.01 | Ω | mΩ |  |  |  | 12.939 | 0.12581 | 12.935 | 0.12525
5502A | Resistance | 0 | 11.001 | 0.1 | Ω | mΩ |  |  |  | 59.144 | 0.029356 | 59.141 | 0.028788
5502A | Resistance | 0 | 11.001 | 1 | Ω | mΩ |  |  |  | 580.12 | 0.003176 | 580.12 | 0.0029427
5502A | Resistance | 0 | 11.001 | 10 | Ω | mΩ |  |  |  | 5800 | 0.00038767 | 5800 | 0.00029434
5502A | Resistance | 0 | 11.001 | 100 | Ω | mΩ |  |  |  | 58000 | 6.6765e-05 | 58000 | 2.9434e-05
5502A | Resistance | 0 | 11.001 | 1000 | Ω | mΩ |  |  |  | 580000 | 1.7876e-05 | 580000 | 2.9434e-06
5502A | Resistance | 0 | 11.001 | 10000 | Ω | mΩ |  |  |  | 5800000 | 6.2674e-06 | 5800000 | 2.9435e-07
5502A | Resistance | 0 | 11.001 | 100000 | Ω | mΩ |  |  |  | 58000000 | 2.4177e-06 | 58000000 | 2.9392e-08
3458A | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0 | 5 | kHz | 0.037 | 700 | 0.035478 | 705.46
3458A | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0 | 5 | kHz | 0.037 | 701.57 | 0.035484 | 705.45
3458A | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0 | 5 | kHz | 0.037 | 703.27 | 0.035899 | 702.95
3458A | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0 | 5 | kHz | 0.066315 | 546.37 | 0.066275 | 544.96
3458A | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0 | 5 | kHz | 0.57804 | 91.593 | 0.57802 | 90.228
3458A | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0 | 5 | kHz | 5.7736 | 9.7715 | 5.7728 | 9.2143
3458A | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0 | 5 | kHz | 57.736 | 1.1862 | 57.735 | 0.96332
3458A | ACA | 0 | 0.0001 | 0.001 | A | µA | 0 | 5 | kHz | 577.36 | 0.20221 | 577.35 | 0.11305
3458A | ACA | 0 | 0.0001 | 0.01 | A | µA | 0 | 5 | kHz | 5773.6 | 0.053657 | 5773.6 | 0.017992
3458A | ACA | 0 | 0.0001 | 0.1 | A | µA | 0 | 5 | kHz | 57736 | 0.01874 | 57736 | 0.0044741
3458A | ACA | 0 | 0.0001 | 1 | A | µA | 0 | 5 | kHz | 577360 | 0.0072242 | 577360 | 0.0015173
3458A | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.077 | 4200 | 0.076632 | 4198.9
3458A | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.077 | 4200 | 0.076633 | 4198.9
3458A | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.077 | 4200 | 0.076839 | 4197.2
3458A | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.095326 | 4045.9 | 0.095288 | 4045.8
3458A | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0 | 0 | kHz | 0.58121 | 1804 | 0.58119 | 1803.1
3458A | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0 | 0 | kHz | 5.7739 | 210.05 | 5.7723 | 209.55
3458A | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0 | 0 | kHz | 57.736 | 21.197 | 57.734 | 20.994
3458A | ACA | 0 | 0.0001 | 0.001 | A | µA | 0 | 0 | kHz | 577.36 | 2.1805 | 577.35 | 2.0995
3458A | ACA | 0 | 0.0001 | 0.01 | A | µA | 0 | 0 | kHz | 5773.6 | 0.24237 | 5773.6 | 0.20995
3458A | ACA | 0 | 0.0001 | 0.1 | A | µA | 0 | 0 | kHz | 57736 | 0.033964 | 57736 | 0.020995
3458A | ACA | 0 | 0.0001 | 1 | A | µA | 0 | 0 | kHz | 577360 | 0.0072871 | 577360 | 0.002099
3458A | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.038 | 1700 | 0.034644 | 1732.1
3458A | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.038 | 1700 | 0.034648 | 1732.1
3458A | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.038 | 1700 | 0.034998 | 1729.4
3458A | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.062399 | 1534.5 | 0.062349 | 1534
3458A | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0 | 0 | kHz | 0.57594 | 378.74 | 0.57592 | 377.44
3458A | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0 | 0 | kHz | 5.7734 | 39.656 | 5.77 | 39.099
3458A | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0 | 0 | kHz | 57.736 | 4.1759 | 57.732 | 3.953
3458A | ACA | 0 | 0.0001 | 0.001 | A | µA | 0 | 0 | kHz | 577.36 | 0.50118 | 577.35 | 0.41202
3458A | ACA | 0 | 0.0001 | 0.01 | A | µA | 0 | 0 | kHz | 5773.6 | 0.083554 | 5773.5 | 0.047889
3458A | ACA | 0 | 0.0001 | 0.1 | A | µA | 0 | 0 | kHz | 57736 | 0.02173 | 57736 | 0.0074638
3458A | ACA | 0 | 0.0001 | 1 | A | µA | 0 | 0 | kHz | 577360 | 0.007523 | 577360 | 0.0018174
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 1000 | 4000 | kHz | 23 | 80000 | 22.679 | 79411
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 1000 | 4000 | kHz | 23 | 79583 | 22.679 | 79411
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 1000 | 4000 | kHz | 23 | 79581 | 22.688 | 79410
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 1000 | 4000 | kHz | 23.293 | 79521 | 23.228 | 79358
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 1000 | 4000 | kHz | 57.909 | 76095 | 57.746 | 76108
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 1000 | 4000 | kHz | 573.98 | 42637 | 570.56 | 42634
3458A | ACV | 0 | 0.01 | 0.01 | V | µV | 1000 | 4000 | kHz | 5773.1 | 5803.4 | 5766 | 5798.9
3458A | ACV | 0 | 0.01 | 0.1 | V | µV | 1000 | 4000 | kHz | 57735 | 584.6 | 57728 | 582.75
3458A | ACV | 0 | 0.01 | 1 | V | µV | 1000 | 4000 | kHz | 577360 | 59.02 | 577350 | 58.278
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 4000 | 8000 | kHz | 31 | 230000 | 30.119 | 228890
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 4000 | 8000 | kHz | 31 | 228990 | 30.12 | 228890
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 4000 | 8000 | kHz | 31 | 228990 | 30.12 | 228890
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 4000 | 8000 | kHz | 31 | 228940 | 30.434 | 228860
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 4000 | 8000 | kHz | 56.084 | 226390 | 55.738 | 226400
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 4000 | 8000 | kHz | 561.69 | 182820 | 557.14 | 182830
3458A | ACV | 0 | 0.01 | 0.01 | V | µV | 4000 | 8000 | kHz | 5769.3 | 45269 | 5751.2 | 45261
3458A | ACV | 0 | 0.01 | 0.1 | V | µV | 4000 | 8000 | kHz | 57735 | 4703.5 | 57713 | 4700
3458A | ACV | 0 | 0.01 | 1 | V | µV | 4000 | 8000 | kHz | 577360 | 471.57 | 577330 | 470.19
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 0 | 0 | kHz | 4.5 | 310 | 4.3849 | 312.12
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 0 | 0 | kHz | 4.5 | 312.16 | 4.3853 | 312.11
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 0 | 0 | kHz | 4.5 | 310.77 | 4.4226 | 310.57
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 0 | 0 | kHz | 7.2506 | 223.08 | 7.2466 | 222.31
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 0 | 0 | kHz | 57.903 | 32.603 | 57.901 | 32.045
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 0 | 0 | kHz | 577.37 | 3.4481 | 577.37 | 3.2224
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 0 | 1 | kHz | 3 | 160 | 2.9725 | 162.69
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 0 | 1 | kHz | 3 | 162.76 | 2.9731 | 162.67
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 0 | 1 | kHz | 3.0304 | 161.11 | 3.028 | 160.75
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 0 | 1 | kHz | 6.4968 | 89.969 | 6.4924 | 88.922
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 0 | 1 | kHz | 57.814 | 11.229 | 57.812 | 10.667
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 0 | 1 | kHz | 577.36 | 1.2948 | 577.36 | 1.0691
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 1 | 20 | kHz | 3.1 | 250 | 3.0556 | 254.12
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 1 | 20 | kHz | 3.1 | 254.16 | 3.0563 | 254.09
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 1 | 20 | kHz | 3.1119 | 251.96 | 3.1096 | 251.7
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 1 | 20 | kHz | 6.5335 | 152.12 | 6.5291 | 151.19
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 1 | 20 | kHz | 57.818 | 19.6 | 57.816 | 19.039
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 1 | 20 | kHz | 577.36 | 2.1354 | 577.36 | 1.9097
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 20 | 50 | kHz | 5.6 | 840 | 5.5358 | 839.05
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 20 | 50 | kHz | 5.6 | 839.02 | 5.5362 | 839.03
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 20 | 50 | kHz | 5.6 | 837.24 | 5.5656 | 837.27
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 20 | 50 | kHz | 7.9956 | 709.13 | 7.9861 | 708.96
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 20 | 50 | kHz | 58 | 140.31 | 57.994 | 139.7
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 20 | 50 | kHz | 577.38 | 14.459 | 577.37 | 14.2
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 50 | 100 | kHz | 8.8 | 5100 | 8.3103 | 5147.5
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 50 | 100 | kHz | 8.8 | 5147.5 | 8.3103 | 5147.5
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 50 | 100 | kHz | 8.8 | 5145.7 | 8.3291 | 5145.9
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 50 | 100 | kHz | 10.062 | 5001.3 | 10.046 | 5001.8
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 50 | 100 | kHz | 58.168 | 2497.7 | 58.157 | 2495.6
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 50 | 100 | kHz | 577.4 | 306.39 | 577.17 | 305.02
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.8932 | 46131
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.8935 | 46131
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 100 | 300 | kHz | 4.2005 | 46000 | 2.919 | 46128
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 100 | 300 | kHz | 4.9503 | 45929 | 4.8763 | 45936
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 100 | 300 | kHz | 54.13 | 41367 | 54.086 | 41369
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 100 | 300 | kHz | 575.77 | 16511 | 573.27 | 16507
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 300 | 1000 | kHz | 18 | 45000 | 12.904 | 45496
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 300 | 1000 | kHz | 18 | 45496 | 12.904 | 45496
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 300 | 1000 | kHz | 18 | 45495 | 12.915 | 45495
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 300 | 1000 | kHz | 18 | 45403 | 13.84 | 45406
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 300 | 1000 | kHz | 56.207 | 41523 | 56.163 | 41525
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 300 | 1000 | kHz | 575.97 | 16722 | 573.47 | 16718
3458A | CDR | 0 | 593 | 1 | °C | °C |  |  |  | 0.59803 | 0 | 0.5817 | 0
3458A | CDR | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.11485 | 0 | 0.091518 | 0
3458A | CDS | 0 | 593 | 1 | °C | °C |  |  |  | 0.5968 | 0 | 0.58081 | 0
3458A | CDS | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.10976 | 0 | 0.085663 | 0
3458A | DCA | 0 | 1e-07 | 1e-12 | A | nA |  |  |  | 0.048 | 34000 | 0.046246 | 34649
3458A | DCA | 0 | 1e-07 | 1e-11 | A | nA |  |  |  | 0.048 | 34000 | 0.046601 | 34403
3458A | DCA | 0 | 1e-07 | 1e-10 | A | nA |  |  |  | 0.074257 | 24027 | 0.07397 | 22146
3458A | DCA | 0 | 1e-07 | 1e-09 | A | nA |  |  |  | 0.57935 | 3897.9 | 0.5792 | 2870.4
3458A | DCA | 0 | 1e-07 | 1e-08 | A | nA |  |  |  | 5.7738 | 700.9 | 5.7737 | 288.02
3458A | DCA | 0 | 1e-07 | 1e-07 | A | nA |  |  |  | 57.736 | 193.97 | 57.736 | 28.803
3458A | DCA | 0 | 1e-07 | 1e-06 | A | nA |  |  |  | 577.36 | 68.945 | 577.36 | 2.8803
3458A | DCA | 0 | 1e-07 | 1e-05 | A | nA |  |  |  | 5773.6 | 26.714 | 5773.6 | 0.28803
3458A | DCA | 0 | 1e-07 | 0.0001 | A | nA |  |  |  | 57736 | 10.599 | 57736 | 0.028736
3458A | DCA | 0 | 1e-07 | 0.001 | A | nA |  |  |  | 577360 | 4.2309 | 577360 | 0
3458A | DCA | 0 | 1e-07 | 0.01 | A | nA |  |  |  | 5773600 | 1.6933 | 5773600 | 0
3458A | DCA | 0 | 1e-07 | 0.1 | A | nA |  |  |  | 57736000 | 0.75258 | 57736000 | 0
3458A | DCV | 0 | 0.1 | 1e-08 | V | µV |  |  |  | 0.62 | 8 | 0.60546 | 8.1092
3458A | DCV | 0 | 0.1 | 1e-07 | V | µV |  |  |  | 0.62 | 8.0238 | 0.60819 | 8.0937
3458A | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 0.83938 | 6.9596 | 0.83737 | 6.9316
3458A | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 5.8321 | 1.4327 | 5.831 | 1.3947
3458A | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.004 | 0.15763 | 58.003 | 0.14188
3458A | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.020493 | 580 | 0.014191
3458A | DCV | 0 | 0.1 | 0.01 | V | mV |  |  |  | 5.8 | 3.94e-06 | 5.7999 | 1.42e-06
3458A | DCV | 0 | 0.1 | 0.1 | V | mV |  |  |  | 58 | 1.15e-06 | 58 | 1.42e-07
3458A | DCV | 0 | 0.1 | 1 | V | mV |  |  |  | 580 | 4.18e-07 | 580 | 1.42e-08
3458A | Frequency | 0 | 40 | 1e-06 | Hz | Hz |  |  |  | 0 | 6e-10 | 1.1e-11 | 5.77e-10
3458A | Frequency | 0 | 40 | 1e-05 | Hz | Hz |  |  |  | 0 | 6e-10 | 2.97e-11 | 5.77e-10
3458A | Frequency | 0 | 40 | 0.0001 | Hz | Hz |  |  |  | 0 | 6e-10 | 2.95e-09 | 5.77e-10
3458A | Frequency | 0 | 40 | 0.001 | Hz | Hz |  |  |  | 2.23e-07 | 5.94e-10 | 2.45e-07 | 5.71e-10
3458A | Frequency | 0 | 40 | 0.01 | Hz | Hz |  |  |  | 5.34e-06 | 4.67e-10 | 5.34e-06 | 4.62e-10
3458A | Frequency | 0 | 40 | 0.1 | Hz | Hz |  |  |  | 5.7625e-05 | 1.14e-10 | 5.7277e-05 | 1.14e-10
3458A | Frequency | 0 | 40 | 1 | Hz | Hz |  |  |  | 0.00057734 | 1.18e-11 | 0.00057689 | 1.18e-11
3458A | Frequency | 0 | 40 | 10 | Hz | Hz |  |  |  | 0.0057736 | 1.18e-12 | 0.0057731 | 1.18e-12
3458A | Frequency | 0 | 40 | 100 | Hz | Hz |  |  |  | 0.057736 | 1.18e-13 | 0.057735 | 1.18e-13
3458A | Frequency | 0 | 40 | 1000 | Hz | Hz |  |  |  | 0.57736 | 1.18e-14 | 0.57735 | 1.18e-14
3458A | Frequency | 0 | 40 | 10000 | Hz | Hz |  |  |  | 5.7736 | 1.18e-15 | 5.7736 | 1.18e-15
3458A | Resistance | 0 | 10 | 1e-05 | Ω | µΩ |  |  |  | 65 | 17 | 61.668 | 17.321
3458A | Resistance | 0 | 10 | 0.0001 | Ω | µΩ |  |  |  | 81.734 | 16.018 | 81.655 | 16.015
3458A | Resistance | 0 | 10 | 0.001 | Ω | µΩ |  |  |  | 578.24 | 4.5207 | 578.2 | 4.5079
3458A | Resistance | 0 | 10 | 0.01 | Ω | µΩ |  |  |  | 5773.6 | 0.47627 | 5773.6 | 0.47059
3458A | Resistance | 0 | 10 | 0.1 | Ω | µΩ |  |  |  | 57736 | 0.049352 | 57736 | 0.047081
3458A | Resistance | 0 | 10 | 1 | Ω | µΩ |  |  |  | 577360 | 0.0056165 | 577360 | 0.0047081
3458A | Resistance | 0 | 10 | 10 | Ω | µΩ |  |  |  | 5773600 | 0.00083417 | 5773600 | 0.00047081
3458A | Resistance | 0 | 10 | 100 | Ω | µΩ |  |  |  | 57736000 | 0.00019242 | 57736000 | 4.708e-05
3458A | Resistance | 0 | 10 | 1000 | Ω | µΩ |  |  |  | 577360000 | 6.285e-05 | 577360000 | 4.7021e-06
3458A | Resistance | 0 | 10 | 10000 | Ω | µΩ |  |  |  | 5773600000 | 2.3736e-05 | 5773600000 | 0
3458A | Resistance | 0 | 10 | 100000 | Ω | µΩ |  |  |  | 57736000000 | 8.4771e-06 | 57736000000 | 0
3458A | Resistance | 0 | 10 | 1000000 | Ω | µΩ |  |  |  | 577000000000 | 0 | 577000000000 | 0
34420A | DCVmeas | 0 | 0.001 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.027845 | 71.95
34420A | DCVmeas | 0 | 0.001 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.062207 | 53.073
34420A | DCVmeas | 0 | 0.001 | 1e-06 | V | µV |  |  |  | 0.62 | 4.5 | 0.5802 | 8.2952
34420A | DCVmeas | 0 | 0.001 | 1e-05 | V | µV |  |  |  | 5.8 | 0.89276 | 5.8 | 0.83626
34420A | DCVmeas | 0 | 0.001 | 0.0001 | V | µV |  |  |  | 58 | 0.10624 | 58 | 0.083633
34420A | DCVmeas | 0 | 0.001 | 0.001 | V | µV |  |  |  | 580 | 0.017406 | 580 | 0.0083633
34420A | Resistance | 0 | 1000 | 1e-06 | Ohms | µΩ |  |  |  | 3 | 0.12 | 2.3273 | 0.080813
34420A | Resistance | 0 | 1000 | 1e-05 | Ohms | µΩ |  |  |  | 3 | 0.12 | 3.9637 | 0.079377
34420A | Resistance | 0 | 1000 | 0.0001 | Ohms | µΩ |  |  |  | 63.6928 | 0.062659 | 54.2074 | 0.047163
34420A | Resistance | 0 | 1000 | 0.001 | Ohms | µΩ |  |  |  | 600.322 | 0.096481 | 579.4447 | 0.0064836
34420A | Resistance | 0 | 1000 | 0.01 | Ohms | mΩ |  |  |  | 5.8513 | 0.00024805 | 5.7999 | 6.5169e-07
Transmille | ACV | 0 | 0.1 | 1e-06 | V | µV | 0.01 | 2 | KHz | 35 | 940 | 34.74 | 937
Transmille | ACV | 0 | 0.1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 35.162 | 934 | 35.157 | 934
Transmille | ACV | 0 | 0.1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 67.116 | 737 | 67.105 | 737
Transmille | ACV | 0 | 0.1 | 0.001 | V | µV | 0.01 | 2 | KHz | 578.33 | 132 | 578.32 | 131
Transmille | ACV | 0 | 0.1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5773.7 | 13.4 | 5773.6 | 13.3
Transmille | ACV | 0 | 0.1 | 1e-06 | V | µV | 2 | 20 | KHz | 36 | 930 | 35.471 | 926
Transmille | ACV | 0 | 0.1 | 1e-05 | V | µV | 2 | 20 | KHz | 36 | 923 | 35.865 | 923
Transmille | ACV | 0 | 0.1 | 0.0001 | V | µV | 2 | 20 | KHz | 67.126 | 727 | 67.116 | 727
Transmille | ACV | 0 | 0.1 | 0.001 | V | µV | 2 | 20 | KHz | 578.33 | 129 | 578.32 | 129
Transmille | ACV | 0 | 0.1 | 0.01 | V | µV | 2 | 20 | KHz | 5773.7 | 13.2 | 5773.6 | 13.1
Transmille | DCA | 0 | 0.0001 | 1e-09 | A | nA |  |  |  | 35 | 350000 | 34.933 | 349000
Transmille | DCA | 0 | 0.0001 | 1e-08 | A | nA |  |  |  | 35.368 | 347000 | 35.366 | 347000
Transmille | DCA | 0 | 0.0001 | 1e-07 | A | nA |  |  |  | 67.34 | 231000 | 67.336 | 231000
Transmille | DCA | 0 | 0.0001 | 1e-06 | A | nA |  |  |  | 578.4 | 31400 | 578.39 | 31400
Transmille | DCA | 0 | 0.0001 | 1e-05 | A | nA |  |  |  | 5773.7 | 3160 | 5773.7 | 3150
Transmille | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 12 | 92 | 11.624 | 92.1
Transmille | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 12.963 | 86.4 | 12.956 | 86.5
Transmille | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.889 | 24.9 | 58.883 | 24.9
Transmille | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 577.47 | 2.59 | 577.47 | 2.59
Transmille | DCV | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5773.6 | 0.259 | 5773.6 | 0.259
Transmille | TcindicatorR | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6784 | 0 | 2.6775 | 0
Transmille | TcindicatorR | 0 | 250 | 1 | °C | °C |  |  |  | 2.7419 | 0 | 2.7384 | 0
Transmille | TcindicatorS | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6783 | 0 | 2.6775 | 0
Transmille | TcindicatorS | 0 | 250 | 1 | °C | °C |  |  |  | 2.7418 | 0 | 2.7384 | 0
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-13 | F | nF |  |  |  | 0.012 | 5800000 | 0.01155 | 5775200
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-12 | F | nF |  |  |  | 0.012 | 5800000 | 0.011563 | 5772800
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-11 | F | nF |  |  |  | 0.012845 | 5547800 | 0.012838 | 5547800
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-10 | F | nF |  |  |  | 0.058705 | 2016900 | 0.058698 | 2014500
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-09 | F | nF |  |  |  | 0.57745 | 218190 | 0.57745 | 217000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-12 | F | nF |  |  |  | 0.036 | 6600000 | 0.035476 | 6560000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-11 | F | nF |  |  |  | 0.036 | 6540000 | 0.035893 | 6530000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-10 | F | nF |  |  |  | 0.066425 | 5020000 | 0.066402 | 5010000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-09 | F | nF |  |  |  | 0.57863 | 1100000 | 0.57863 | 1100000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-08 | F | nF |  |  |  | 5.7903 | 3020000 | 5.7903 | 3020000
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-12 | F | nF |  |  |  | 0.012 | 2900000 | 0.011576 | 2885200
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-11 | F | nF |  |  |  | 0.012482 | 2856200 | 0.012511 | 2834700
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-10 | F | nF |  |  |  | 0.056911 | 1398700 | 0.056878 | 1390600
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-09 | F | nF |  |  |  | 0.57722 | 165280 | 0.57721 | 160660
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-11 | F | nF |  |  |  | 0.5 | 5800000 | 0.48989 | 5800000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-10 | F | nF |  |  |  | 0.5 | 5790000 | 0.49306 | 5790000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-09 | F | nF |  |  |  | 0.74933 | 4690000 | 0.74914 | 4680000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-08 | F | nF |  |  |  | 5.797 | 1100000 | 5.797 | 1090000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-07 | F | nF |  |  |  | 57.904 | 3020000 | 57.904 | 3020000
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-12 | F | nF |  |  |  | 0.12 | 2900000 | 0.11547 | 2886700
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-11 | F | nF |  |  |  | 0.12 | 2900000 | 0.11561 | 2885200
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-10 | F | nF |  |  |  | 0.12801 | 2745100 | 0.12797 | 2743500
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-09 | F | nF |  |  |  | 0.58647 | 856770 | 0.58644 | 853690
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-08 | F | nF |  |  |  | 5.7745 | 853270 | 5.7744 | 89445
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-11 | F | nF |  |  |  | 0.12 | 2900000 | 0.11565 | 2886200
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-10 | F | nF |  |  |  | 0.12459 | 2858300 | 0.125 | 2835700
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-09 | F | nF |  |  |  | 0.56898 | 1398000 | 0.56874 | 1390900
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-08 | F | nF |  |  |  | 5.7722 | 164700 | 5.7721 | 160690
3458A | DCA | 1e-07 | 1e-06 | 1e-12 | A | nA |  |  |  | 0.048 | 23000 | 0.046073 | 24826
3458A | DCA | 1e-07 | 1e-06 | 1e-11 | A | nA |  |  |  | 0.048 | 23271 | 0.046424 | 24708
3458A | DCA | 1e-07 | 1e-06 | 1e-10 | A | nA |  |  |  | 0.074091 | 17690 | 0.07366 | 17772
3458A | DCA | 1e-07 | 1e-06 | 1e-09 | A | nA |  |  |  | 0.57936 | 2553.2 | 0.57914 | 2554.1
3458A | DCA | 1e-07 | 1e-06 | 1e-08 | A | nA |  |  |  | 5.7738 | 256.81 | 5.7737 | 256.81
3458A | DCA | 1e-07 | 1e-06 | 1e-07 | A | nA |  |  |  | 57.736 | 25.682 | 57.736 | 25.682
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-10 | F | nF |  |  |  | 6.2 | 5200000 | 6.1439 | 5210000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-09 | F | nF |  |  |  | 6.2 | 5200000 | 6.1698 | 5200000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-08 | F | nF |  |  |  | 8.3825 | 4380000 | 8.3808 | 4370000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-07 | F | nF |  |  |  | 58.094 | 1100000 | 58.094 | 1090000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-06 | F | nF |  |  |  | 579.05 | 3020000 | 579.05 | 3020000
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-11 | F | nF |  |  |  | 0.35 | 2900000 | 0.34649 | 2886700
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-10 | F | nF |  |  |  | 0.35 | 2900000 | 0.34958 | 2881200
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-09 | F | nF |  |  |  | 0.60931 | 2463100 | 0.60906 | 2462300
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-08 | F | nF |  |  |  | 5.7585 | 484500 | 5.7585 | 483140
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-10 | F | nF |  |  |  | 1.2 | 2900000 | 1.1566 | 2886200
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-09 | F | nF |  |  |  | 1.2438 | 2860200 | 1.2501 | 2835600
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-08 | F | nF |  |  |  | 5.6891 | 1394000 | 5.6874 | 1390900
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-07 | F | nF |  |  |  | 57.722 | 162480 | 57.721 | 160690
3458A | DCA | 1e-06 | 1e-05 | 1e-12 | A | nA |  |  |  | 0.13 | 25000 | 0.11375 | 25625
3458A | DCA | 1e-06 | 1e-05 | 1e-11 | A | nA |  |  |  | 0.13 | 25119 | 0.11389 | 25617
3458A | DCA | 1e-06 | 1e-05 | 1e-10 | A | nA |  |  |  | 0.13 | 24832 | 0.12602 | 24846
3458A | DCA | 1e-06 | 1e-05 | 1e-09 | A | nA |  |  |  | 0.58396 | 10197 | 0.58374 | 10200
3458A | DCA | 1e-06 | 1e-05 | 1e-08 | A | nA |  |  |  | 5.7742 | 1129.1 | 5.7741 | 1129.1
3458A | DCA | 1e-06 | 1e-05 | 1e-07 | A | nA |  |  |  | 57.736 | 113.04 | 57.735 | 113.04
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-10 | F | nF |  |  |  | 3.5 | 2900000 | 3.4649 | 2886700
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-09 | F | nF |  |  |  | 3.5 | 2900000 | 3.4958 | 2881200
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-08 | F | nF |  |  |  | 6.0937 | 2463200 | 6.0906 | 2462300
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-07 | F | nF |  |  |  | 57.585 | 484680 | 57.585 | 483140
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-09 | F | nF |  |  |  | 12 | 2900000 | 11.566 | 2886200
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-08 | F | nF |  |  |  | 12.429 | 2861100 | 12.501 | 2835600
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-07 | F | nF |  |  |  | 56.88 | 1394300 | 56.874 | 1390900
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-06 | F | nF |  |  |  | 577.21 | 162540 | 577.21 | 160690
3458A | DCA | 1e-05 | 0.0001 | 1e-11 | A | nA |  |  |  | 0.16 | 23000 | 0.11411 | 23000
3458A | DCA | 1e-05 | 0.0001 | 1e-10 | A | nA |  |  |  | 0.16 | 23403 | 0.11928 | 23403
3458A | DCA | 1e-05 | 0.0001 | 1e-09 | A | nA |  |  |  | 0.47124 | 20574 | 0.46866 | 20574
3458A | DCA | 1e-05 | 0.0001 | 1e-08 | A | nA |  |  |  | 5.7304 | 5486.9 | 5.7294 | 5486.9
3458A | DCA | 1e-05 | 0.0001 | 1e-07 | A | nA |  |  |  | 57.731 | 581.51 | 57.731 | 581.51
3458A | DCA | 1e-05 | 0.0001 | 1e-06 | A | nA |  |  |  | 577.35 | 61.585 | 577.35 | 61.585
Transmille | ACA | 1e-05 | 0.000104 | 1e-09 | A | µA | 0.01 | 2 | KHz | 0.47 | 1200 | 0.46304 | 1230
Transmille | ACA | 1e-05 | 0.000104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.47 | 1230 | 0.46308 | 1230
Transmille | ACA | 1e-05 | 0.000104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.47 | 1230 | 0.4666 | 1230
Transmille | ACA | 1e-05 | 0.000104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 0.73981 | 857 | 0.73979 | 856
Transmille | ACA | 1e-05 | 0.000104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8088 | 394 | 5.8088 | 394
5502A | ACW | 1.09e-05 | 0.000109 | 1e-09 | WAC | WAC |  |  |  | 0 | 0.0026 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.0026 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0058884 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.05332 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.53261 | 0 | 0
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-09 | F | nF |  |  |  | 36 | 4600000 | 34.648 | 4618700
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-08 | F | nF |  |  |  | 36 | 4600000 | 34.894 | 4613900
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-07 | F | nF |  |  |  | 56.814 | 4211300 | 56.808 | 4211000
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-06 | F | nF |  |  |  | 572.02 | 1057800 | 572.02 | 1057200
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-08 | F | nF |  |  |  | 120 | 5200000 | 115.61 | 5195600
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-07 | F | nF |  |  |  | 121.85 | 5183200 | 122.72 | 5152800
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-06 | F | nF |  |  |  | 536.63 | 3284200 | 536.45 | 3282000
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-05 | F | nF |  |  |  | 5766.3 | 438500 | 5766.3 | 436530
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 5 | 20 | kHz | 0.25 | 710 | 0.23471 | 715.83
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 5 | 20 | kHz | 0.25 | 711.93 | 0.23479 | 715.8
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 5 | 20 | kHz | 0.25 | 712.22 | 0.24058 | 711.87
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 5 | 20 | kHz | 0.6031 | 511.67 | 0.60272 | 509.95
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 5 | 20 | kHz | 5.7741 | 79.21 | 5.774 | 77.713
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 20 | 50 | kHz | 0.47 | 4800 | 0.46627 | 4761.7
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 20 | 50 | kHz | 0.47 | 4761.6 | 0.46636 | 4761.6
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 20 | 50 | kHz | 0.47 | 4759.9 | 0.46837 | 4760
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 20 | 50 | kHz | 0.64627 | 4616 | 0.64414 | 4616.1
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 20 | 50 | kHz | 5.636 | 2159.7 | 5.6349 | 2155.3
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 50 | 100 | kHz | 1.9 | 6500 | 1.7592 | 6558
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 50 | 100 | kHz | 1.9 | 6531.9 | 1.7594 | 6558.1
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 50 | 100 | kHz | 1.9 | 6537.4 | 1.7607 | 6558
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 50 | 100 | kHz | 1.9 | 6516.3 | 1.8354 | 6506.9
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 50 | 100 | kHz | 5.8456 | 4362.3 | 5.8321 | 4309.3
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 5 | kHz | 0.25 | 350 | 0.23001 | 364.09
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 5 | kHz | 0.25 | 354.91 | 0.23009 | 364.09
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 5 | kHz | 0.25 | 359.01 | 0.23663 | 360.5
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 5 | kHz | 0.61486 | 217.08 | 0.61454 | 214.58
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 5 | kHz | 5.7771 | 28.936 | 5.777 | 27.424
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.25 | 4600 | 0.23095 | 4618.8
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.25 | 4600 | 0.23098 | 4618.8
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.25 | 4600.3 | 0.23363 | 4616.5
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.4599 | 4424.7 | 0.45952 | 4424.6
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 0 | kHz | 5.6235 | 1918.3 | 5.6233 | 1917.2
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.27 | 1700 | 0.23094 | 1732.1
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.27 | 1700 | 0.231 | 1732.1
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.27 | 1700 | 0.23546 | 1728.5
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.5563 | 1491.3 | 0.55586 | 1490.6
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 0 | kHz | 5.7533 | 346.75 | 5.7532 | 345.32
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.24 | 690 | 0.23094 | 693.02
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.24 | 690 | 0.23102 | 692.99
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.24 | 690 | 0.23691 | 689
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.60249 | 489.7 | 0.6021 | 487.94
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 0 | kHz | 5.7742 | 74.832 | 5.7741 | 73.333
3458A | DCA | 0.0001 | 0.001 | 1e-10 | A | nA |  |  |  | 6.3 | 23000 | 5.7456 | 23552
3458A | DCA | 0.0001 | 0.001 | 1e-09 | A | nA |  |  |  | 6.3 | 23340 | 5.7675 | 23536
3458A | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 7.7388 | 22133 | 7.7351 | 22126
3458A | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 57.586 | 7177.2 | 57.585 | 7158.5
3458A | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 577.34 | 770.78 | 577.34 | 762.27
3458A | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.6 | 79.686 | 5773.5 | 76.28
Transmille | ACA | 0.0001 | 0.00104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.97 | 1200 | 0.9653 | 1190
Transmille | ACA | 0.0001 | 0.00104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.97 | 1190 | 0.96691 | 1190
Transmille | ACA | 0.0001 | 0.00104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 1.1172 | 1120 | 1.1168 | 1120
Transmille | ACA | 0.0001 | 0.00104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8401 | 354 | 5.8398 | 352
Transmille | ACA | 0.0001 | 0.00104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 57.912 | 316 | 57.912 | 315
Transmille | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 120 | 350000 | 115.86 | 347000
Transmille | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 127.29 | 339000 | 127.24 | 339000
Transmille | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 580.76 | 159000 | 580.72 | 159000
Transmille | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.7 | 18400 | 5773.7 | 18400
Transmille | DCA | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 57736 | 1850 | 57736 | 1840
5502A | ACW | 0.000109 | 0.0010901 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.0017 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0017 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.0055645 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.053285 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-08 | WDC | WDC |  |  |  | 0 | 0.00035 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.00063401 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.0053337 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.053213 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.53211 | 0 | 0
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-08 | F | µF |  |  |  | 0.35 | 5200 | 0.34613 | 5199.8
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-07 | F | µF |  |  |  | 0.35 | 5200 | 0.34836 | 5195.3
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-06 | F | µF |  |  |  | 0.55623 | 4803.8 | 0.55608 | 4803.7
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-05 | F | µF |  |  |  | 5.7038 | 1294.4 | 5.7038 | 1293.8
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 0.23 | 2300 | 0.17326 | 2315.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 0.23 | 2300 | 0.17481 | 2314.9
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 0.35294 | 2267.3 | 0.35227 | 2267.5
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 5.4199 | 1301.8 | 5.4195 | 1301.8
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 57.686 | 174.69 | 57.685 | 174.66
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 0.32 | 1400 | 0.17327 | 1443.8
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 0.32 | 1400 | 0.17604 | 1443.1
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 0.41399 | 1381.1 | 0.41349 | 1381.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 5.6118 | 601.75 | 5.6114 | 601.72
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 57.716 | 69.766 | 57.716 | 69.735
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17375 | 1155.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17651 | 1154.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.43013 | 1124.2 | 0.44201 | 1086.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.6654 | 409.29 | 5.6651 | 409.25
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.723 | 45.397 | 57.723 | 45.366
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 0.27 | 2300 | 0.23112 | 2309.7
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 0.27 | 2300 | 0.2328 | 2309.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 0.40201 | 2264.3 | 0.40156 | 2264.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 5.4266 | 1309.3 | 5.4263 | 1309.3
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 57.686 | 176.15 | 57.686 | 176.12
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 0.35 | 5800 | 0.34654 | 5773.6
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 0.35 | 5800 | 0.34724 | 5773.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 0.41769 | 5779.5 | 0.4264 | 5752
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 4.6373 | 4728.3 | 4.6369 | 4728.3
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 57.432 | 1053.3 | 57.432 | 1053.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69332 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69339 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.7339 | 11536
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 10 | 30 | kHz | 3.6592 | 11103 | 3.7773 | 10742
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 10 | 30 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.69332 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.69339 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.7339 | 11536
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 30 | 30.1 | kHz | 3.6592 | 11103 | 3.7773 | 10742
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 30 | 30.1 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
5502A | Capacitance | 0.00033 | 0.0011001 | 1e-07 | F | µF |  |  |  | 1.2 | 5200 | 1.1558 | 5198.3
5502A | Capacitance | 0.00033 | 0.0011001 | 1e-06 | F | µF |  |  |  | 1.2193 | 5182.5 | 1.2268 | 5155.5
5502A | Capacitance | 0.00033 | 0.0011001 | 1e-05 | F | µF |  |  |  | 5.3658 | 3286.5 | 5.3641 | 3284.2
5502A | Capacitance | 0.00033 | 0.0011001 | 0.0001 | F | µF |  |  |  | 57.663 | 438.86 | 57.663 | 436.89
5502A | DCA | 0.00033 | 0.0033001 | 1e-08 | A | µA |  |  |  | 0.068 | 150 | 0.057598 | 152.89
5502A | DCA | 0.00033 | 0.0033001 | 1e-07 | A | µA |  |  |  | 0.076187 | 148.47 | 0.07332 | 149.02
5502A | DCA | 0.00033 | 0.0033001 | 1e-06 | A | µA |  |  |  | 0.56792 | 73.551 | 0.56576 | 73.312
5502A | DCA | 0.00033 | 0.0033001 | 1e-05 | A | µA |  |  |  | 5.7989 | 9.0557 | 5.7981 | 8.8116
5502A | DCA | 0.00033 | 0.0033001 | 0.0001 | A | µA |  |  |  | 58 | 0.98192 | 58 | 0.88328
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 5 | 20 | kHz | 2.5 | 710 | 2.2891 | 721.65
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 5 | 20 | kHz | 2.5 | 712.46 | 2.2904 | 721.65
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 5 | 20 | kHz | 2.5 | 716.59 | 2.3512 | 717.7
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 5 | 20 | kHz | 6.0251 | 518.68 | 6.0044 | 513.52
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 5 | 20 | kHz | 57.745 | 83.689 | 57.738 | 78.853
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 20 | 50 | kHz | 5.1 | 4700 | 4.5719 | 4747.9
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 20 | 50 | kHz | 5.1 | 4739.1 | 4.5722 | 4747.9
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 20 | 50 | kHz | 5.1 | 4739.9 | 4.5921 | 4746.3
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 20 | 50 | kHz | 6.3814 | 4600.8 | 6.366 | 4600.8
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 20 | 50 | kHz | 56.354 | 2143.3 | 56.346 | 2139.9
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 50 | 100 | kHz | 19 | 6500 | 17.151 | 6602.1
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 50 | 100 | kHz | 19 | 6531.1 | 17.153 | 6602.1
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 50 | 100 | kHz | 19 | 6534.1 | 17.164 | 6601.7
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 50 | 100 | kHz | 19 | 6550 | 17.919 | 6548
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 50 | 100 | kHz | 58.261 | 4342.2 | 58.146 | 4318.8
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 5 | kHz | 2.5 | 350 | 2.2952 | 364.62
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 5 | kHz | 2.5 | 355.75 | 2.2965 | 364.68
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 5 | kHz | 2.5 | 362.33 | 2.3641 | 361.33
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 5 | kHz | 6.1659 | 223.94 | 6.1472 | 216.14
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 5 | kHz | 57.778 | 33.176 | 57.771 | 28.291
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 0 | kHz | 2.5 | 4600 | 2.3096 | 4618.8
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 0 | kHz | 2.5 | 4600 | 2.3101 | 4618.8
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 0 | kHz | 2.5 | 4600.7 | 2.3373 | 4616.5
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 0 | kHz | 4.6177 | 4424.3 | 4.5991 | 4424.5
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 0 | kHz | 56.243 | 1921.2 | 56.235 | 1917.8
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 0 | kHz | 2.7 | 1700 | 2.3095 | 1732.1
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 0 | kHz | 2.7 | 1700 | 2.311 | 1732
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 0 | kHz | 2.7 | 1700 | 2.3566 | 1728.5
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 0 | kHz | 5.585 | 1492.6 | 5.5626 | 1490.9
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 0 | kHz | 57.541 | 350.7 | 57.534 | 346.11
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 0 | kHz | 2.4 | 690 | 2.3095 | 693.04
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 0 | kHz | 2.4 | 690 | 2.3107 | 693.04
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 0 | kHz | 2.4 | 690 | 2.3717 | 689.14
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 0 | kHz | 6.0458 | 494.17 | 6.0252 | 488.84
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 0 | kHz | 57.749 | 79.009 | 57.742 | 74.17
3458A | CDPT100 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.010087 | 0 | 0.0093806 | 0
3458A | CDPT100 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57759 | 0 | 0.5774 | 0
3458A | CDPT100 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058677 | 0 | 0.058206 | 0
3458A | CDPT1000 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.0097193 | 0 | 0.0093614 | 0
3458A | CDPT1000 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57749 | 0 | 0.5774 | 0
3458A | CDPT1000 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058438 | 0 | 0.058203 | 0
3458A | DCA | 0.001 | 0.01 | 1e-09 | A | µA |  |  |  | 0.064 | 23 | 0.057395 | 23.593
3458A | DCA | 0.001 | 0.01 | 1e-08 | A | µA |  |  |  | 0.064 | 23.49 | 0.057621 | 23.576
3458A | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 0.07733 | 22.173 | 0.077294 | 22.166
3458A | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 0.57584 | 7.1949 | 0.57583 | 7.1763
3458A | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.7734 | 0.77277 | 5.7734 | 0.76426
3458A | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.736 | 0.079885 | 57.735 | 0.076479
34420A | DCVmeas | 0.001 | 0.01 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.060672 | 48.506
34420A | DCVmeas | 0.001 | 0.01 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.076236 | 47.254
34420A | DCVmeas | 0.001 | 0.01 | 1e-06 | V | µV |  |  |  | 0.62 | 19.155 | 0.56725 | 22.912
34420A | DCVmeas | 0.001 | 0.01 | 1e-05 | V | µV |  |  |  | 5.7983 | 2.757 | 5.7983 | 2.7323
34420A | DCVmeas | 0.001 | 0.01 | 0.0001 | V | µV |  |  |  | 58 | 0.2838 | 58 | 0.27385
34420A | DCVmeas | 0.001 | 0.01 | 0.001 | V | µV |  |  |  | 580 | 0.031364 | 580 | 0.027386
Transmille | ACA | 0.001 | 0.0104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 9.7 | 1200 | 9.653 | 1190
Transmille | ACA | 0.001 | 0.0104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 9.7 | 1190 | 9.6691 | 1190
Transmille | ACA | 0.001 | 0.0104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 11.172 | 1120 | 11.168 | 1120
Transmille | ACA | 0.001 | 0.0104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 58.401 | 354 | 58.398 | 352
Transmille | ACA | 0.001 | 0.0104 | 0.001 | A | µA | 0.01 | 2 | KHz | 579.12 | 316 | 579.12 | 315
Transmille | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 1.2 | 350 | 1.1597 | 347
Transmille | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 1.2739 | 339 | 1.2734 | 339
Transmille | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.8079 | 159 | 5.8074 | 159
Transmille | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.737 | 18.4 | 57.737 | 18.4
Transmille | DCA | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 1.86 | 577.36 | 1.84
5502A | ACW | 0.00109 | 0.0109 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0013 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.0013483 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.0054702 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.053275 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 0.001 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.00024 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.00058093 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.0053277 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.053212 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 0.001 | WDC | WDC |  |  |  | 0 | 0.53211 | 0 | 0
5502A | Capacitance | 0.0011 | 0.0033001 | 1e-07 | F | µF |  |  |  | 3.5 | 5200 | 3.4636 | 5198.9
5502A | Capacitance | 0.0011 | 0.0033001 | 1e-06 | F | µF |  |  |  | 3.5 | 5200 | 3.4866 | 5194.4
5502A | Capacitance | 0.0011 | 0.0033001 | 1e-05 | F | µF |  |  |  | 5.5761 | 4803.9 | 5.5634 | 4802.9
5502A | Capacitance | 0.0011 | 0.0033001 | 0.0001 | F | µF |  |  |  | 57.043 | 1299.1 | 57.038 | 1293.7
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 2.4 | 2100 | 2.3005 | 2087.6
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 2.4 | 2100 | 2.3157 | 2087.1
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 4.0897 | 2048.8 | 4.1185 | 2039.5
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 54.774 | 1117.8 | 54.773 | 1117.7
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 576.95 | 144.81 | 576.95 | 144.75
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 3.8 | 1000 | 2.2968 | 1043.6
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 3.8 | 1000 | 2.3282 | 1042.8
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 4.917 | 977.92 | 4.9145 | 977.93
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 56.865 | 350.66 | 56.863 | 350.55
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 577.26 | 38.407 | 577.26 | 38.349
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 2.4 | 470 | 2.2998 | 465.62
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 2.4 | 470 | 2.3464 | 464.49
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 5.6396 | 392.41 | 5.6369 | 392.36
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 57.584 | 84.81 | 57.583 | 84.672
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 577.34 | 8.7266 | 577.34 | 8.6681
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 2.4 | 930 | 2.3048 | 925.75
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 2.4 | 930 | 2.3386 | 924.88
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 5.0491 | 857.92 | 5.0466 | 857.92
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 57.047 | 284.47 | 57.045 | 284.35
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 577.28 | 30.67 | 577.28 | 30.612
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 3.8 | 2300 | 3.4644 | 2309.4
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 3.8 | 2300 | 3.4807 | 2309
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 5.0186 | 2268.7 | 5.0129 | 2268.8
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 54.397 | 1333.5 | 54.39 | 1333.5
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 576.87 | 180.75 | 576.87 | 180.65
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 10 | 30 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 10 | 30 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 10 | 30 | kHz | 575.41 | 695.3 | 575.4 | 695.06
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 575.41 | 695.3 | 575.4 | 695.06
5502A | Capacitance | 0.0033 | 0.011001 | 1e-06 | F | µF |  |  |  | 13 | 5200 | 12.076 | 5198.3
5502A | Capacitance | 0.0033 | 0.011001 | 1e-05 | F | µF |  |  |  | 13 | 5200 | 12.713 | 5155.6
5502A | Capacitance | 0.0033 | 0.011001 | 0.0001 | F | µF |  |  |  | 53.666 | 3287.9 | 53.64 | 3284.3
5502A | Capacitance | 0.0033 | 0.011001 | 0.001 | F | µF |  |  |  | 576.63 | 439.99 | 576.63 | 436.9
5502A | DCA | 0.0033 | 0.033001 | 1e-07 | A | µA |  |  |  | 0.21 | 150 | 0.201 | 145.79
5502A | DCA | 0.0033 | 0.033001 | 1e-06 | A | µA |  |  |  | 0.44201 | 139.58 | 0.43216 | 139.79
5502A | DCA | 0.0033 | 0.033001 | 1e-05 | A | µA |  |  |  | 5.6422 | 61.544 | 5.6369 | 61.47
5502A | DCA | 0.0033 | 0.033001 | 0.0001 | A | µA |  |  |  | 57.983 | 7.2053 | 57.981 | 7.143
5502A | DCA | 0.0033 | 0.033001 | 0.001 | A | µA |  |  |  | 580 | 0.74076 | 580 | 0.71564
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 5 | 20 | kHz | 25 | 710 | 22.891 | 721.64
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 5 | 20 | kHz | 25 | 712.1 | 22.9 | 721.62
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 5 | 20 | kHz | 25 | 715.16 | 23.497 | 717.59
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 5 | 20 | kHz | 60.128 | 515.53 | 60.018 | 512.89
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 5 | 20 | kHz | 577.41 | 80.741 | 577.37 | 78.263
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 20 | 50 | kHz | 51 | 4700 | 45.602 | 4746.9
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 20 | 50 | kHz | 51 | 4736.8 | 45.606 | 4746.9
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 20 | 50 | kHz | 51 | 4738 | 45.823 | 4745.2
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 20 | 50 | kHz | 64.008 | 4598.6 | 63.621 | 4599.4
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 20 | 50 | kHz | 563.71 | 2145.2 | 563.49 | 2139.2
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 50 | 100 | kHz | 190 | 6500 | 171.51 | 6602.2
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 50 | 100 | kHz | 190 | 6538.1 | 171.63 | 6602.5
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 50 | 100 | kHz | 190 | 6565.8 | 179.15 | 6551.2
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 50 | 100 | kHz | 582.09 | 4394.2 | 581.36 | 4329.2
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 50 | 100 | kHz | 5772.5 | 669.65 | 5772.4 | 620.3
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 5 | kHz | 25 | 350 | 22.95 | 364.61
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 5 | kHz | 25 | 355.17 | 22.961 | 364.62
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 5 | kHz | 25 | 360.06 | 23.625 | 361.08
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 5 | kHz | 61.548 | 219.19 | 61.449 | 215.18
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 5 | kHz | 577.74 | 30.199 | 577.71 | 27.695
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 0 | kHz | 25 | 4600 | 23.095 | 4618.8
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 0 | kHz | 25 | 4600 | 23.099 | 4618.8
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 0 | kHz | 25 | 4600.4 | 23.366 | 4616.5
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 0 | kHz | 46.067 | 4424.4 | 45.964 | 4424.6
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 0 | kHz | 562.38 | 1919.1 | 562.34 | 1917.3
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 0 | kHz | 27 | 1700 | 23.095 | 1732.1
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 0 | kHz | 27 | 1700 | 23.102 | 1732.1
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 0 | kHz | 27 | 1700 | 23.554 | 1728.5
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 0 | kHz | 55.716 | 1491.6 | 55.6 | 1490.7
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 0 | kHz | 575.37 | 347.9 | 575.33 | 345.55
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 0 | kHz | 24 | 690 | 23.094 | 693.05
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 0 | kHz | 24 | 690 | 23.103 | 693.03
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 0 | kHz | 24 | 690 | 23.701 | 689.05
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 0 | kHz | 60.336 | 490.93 | 60.227 | 488.21
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 0 | kHz | 577.45 | 76.063 | 577.41 | 73.582
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 1000 | 2000 | kHz | 53 | 17000 | 11.488 | 17414
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.488 | 17414
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.489 | 17414
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.588 | 17413
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 53 | 17065 | 21.129 | 17327
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 1000 | 2000 | kHz | 469.05 | 13766 | 433.19 | 13767
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.778 | 46223
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.778 | 46223
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.779 | 46223
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.809 | 46223
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 84.139 | 46193
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 2000 | 4000 | kHz | 354.54 | 43842 | 330.66 | 43842
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.256 | 46261
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.256 | 46261
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.257 | 46261
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.287 | 46261
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 95.542 | 46232
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 4000 | 8000 | kHz | 363.53 | 43909 | 338.44 | 43920
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 115.4 | 173280
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 8000 | 10000 | kHz | 450 | 170030 | 115.4 | 173280
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 115.4 | 173280
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 115.41 | 173280
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 116.4 | 173270
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 8000 | 10000 | kHz | 450 | 170150 | 203.58 | 172400
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 0 | 0 | kHz | 5.4 | 83 | 5.316 | 83.165
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 0 | 0 | kHz | 5.4 | 83.166 | 5.3163 | 83.163
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 0 | 0 | kHz | 5.4 | 83.01 | 5.3447 | 83
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 0 | 0 | kHz | 7.7286 | 70.835 | 7.7258 | 70.789
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 0 | 0 | kHz | 57.923 | 14.075 | 57.922 | 14.014
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 0 | 0 | kHz | 577.37 | 1.4498 | 577.25 | 1.4244
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 0 | 1 | kHz | 3.5 | 81 | 3.4429 | 80.684
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 0 | 1 | kHz | 3.5 | 80.684 | 3.4434 | 80.682
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 0 | 1 | kHz | 3.5 | 80.42 | 3.4848 | 80.411
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 0 | 1 | kHz | 6.538 | 63.487 | 6.5346 | 63.435
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 0 | 1 | kHz | 57.784 | 10.953 | 57.783 | 10.891
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 0 | 1 | kHz | 577.36 | 1.1266 | 577.26 | 1.1012
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 1 | 20 | kHz | 3.4 | 160 | 3.3551 | 157.72
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 1 | 20 | kHz | 3.4 | 157.72 | 3.3552 | 157.72
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 1 | 20 | kHz | 3.4 | 157.45 | 3.3911 | 157.45
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 1 | 20 | kHz | 6.2204 | 137.65 | 6.2175 | 137.62
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 1 | 20 | kHz | 57.628 | 32.01 | 57.626 | 31.951
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 1 | 20 | kHz | 577.34 | 3.3107 | 577.06 | 3.2853
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 20 | 50 | kHz | 4.7 | 330 | 4.6688 | 326.81
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 20 | 50 | kHz | 4.7 | 327.36 | 4.6691 | 326.81
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 20 | 50 | kHz | 4.7 | 327.1 | 4.6916 | 326.63
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 20 | 50 | kHz | 6.7172 | 310.87 | 6.7063 | 310.87
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 20 | 50 | kHz | 57.119 | 116.68 | 57.114 | 116.5
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 20 | 50 | kHz | 577.28 | 12.896 | 576.24 | 12.803
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 50 | 100 | kHz | 6 | 910 | 5.9729 | 906.03
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 50 | 100 | kHz | 6 | 906.02 | 5.973 | 906.03
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 50 | 100 | kHz | 6 | 905.89 | 5.9854 | 905.92
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 50 | 100 | kHz | 7.1698 | 895.95 | 7.1433 | 896.05
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 50 | 100 | kHz | 53.812 | 587.62 | 53.788 | 587.3
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 50 | 100 | kHz | 576.69 | 87.306 | 571.68 | 86.954
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.521 | 3476.1
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.521 | 3476.1
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.522 | 3476.1
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.911 | 3472.7
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 100 | 300 | kHz | 41.904 | 3219 | 41.802 | 3219.4
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 100 | 300 | kHz | 568.06 | 1120 | 547.01 | 1119.2
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.635 | 11630
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 300 | 1000 | kHz | 21.308 | 11907 | 25.139 | 11509
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 300 | 1000 | kHz | 511.75 | 7972.6 | 476.24 | 7972.4
3458A | DCA | 0.01 | 0.1 | 1e-08 | A | µA |  |  |  | 0.67 | 40 | 0.57451 | 40.858
3458A | DCA | 0.01 | 0.1 | 1e-07 | A | µA |  |  |  | 0.67 | 40.354 | 0.57629 | 40.844
3458A | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 0.75326 | 39.474 | 0.74494 | 39.51
3458A | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 5.6886 | 17.431 | 5.683 | 17.368
3458A | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 57.726 | 2.0334 | 57.724 | 1.9935
3458A | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.35 | 0.21573 | 577.34 | 0.19968
34420A | DCVmeas | 0.01 | 0.1 | 1e-07 | V | µV |  |  |  | 0.48 | 37 | 0.47363 | 36.624
34420A | DCVmeas | 0.01 | 0.1 | 1e-06 | V | µV |  |  |  | 0.6696 | 35.075 | 0.66823 | 35.079
34420A | DCVmeas | 0.01 | 0.1 | 1e-05 | V | µV |  |  |  | 5.7207 | 14.051 | 5.7199 | 14.036
34420A | DCVmeas | 0.01 | 0.1 | 0.0001 | V | µV |  |  |  | 57.991 | 1.5774 | 57.99 | 1.5689
34420A | DCVmeas | 0.01 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.16049 | 580 | 0.1571
34420A | DCVmeas | 0.01 | 0.1 | 0.01 | V | µV |  |  |  | 5800 | 0.017067 | 5800 | 0.01571
Transmille | ACA | 0.01 | 0.104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 97 | 1200 | 96.53 | 1190
Transmille | ACA | 0.01 | 0.104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 97 | 1190 | 96.691 | 1190
Transmille | ACA | 0.01 | 0.104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 111.72 | 1120 | 111.68 | 1120
Transmille | ACA | 0.01 | 0.104 | 0.001 | A | µA | 0.01 | 2 | KHz | 583.96 | 354 | 583.98 | 352
Transmille | ACA | 0.01 | 0.104 | 0.01 | A | µA | 0.01 | 2 | KHz | 5791.2 | 316 | 5791.2 | 315
Transmille | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 12 | 350 | 11.589 | 348
Transmille | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 12.73 | 341 | 12.725 | 341
Transmille | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 58.071 | 160 | 58.066 | 160
Transmille | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.37 | 18.6 | 577.37 | 18.5
Transmille | DCA | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 1.87 | 5773.6 | 1.86
5502A | ACW | 0.0109 | 0.109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.00078 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.00094316 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.0053847 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 0.001 | WAC | WAC |  |  |  | 0 | 0.053267 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 0.01 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.00019 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.00056395 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.0053308 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 0.001 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 0.01 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | Capacitance | 0.011 | 0.033001 | 1e-06 | F | µF |  |  |  | 35 | 8700 | 34.642 | 8661.9
5502A | Capacitance | 0.011 | 0.033001 | 1e-05 | F | µF |  |  |  | 35 | 8700 | 34.806 | 8658.5
5502A | Capacitance | 0.011 | 0.033001 | 0.0001 | F | µF |  |  |  | 50.508 | 8339.7 | 50.437 | 8339.6
5502A | Capacitance | 0.011 | 0.033001 | 0.001 | F | µF |  |  |  | 557.55 | 3119.3 | 557.52 | 3115.8
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.01 | 0.02001 | kHz | 0.023 | 2.1 | 0.022917 | 2.0907
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.01 | 0.02001 | kHz | 0.023 | 2.1 | 0.023129 | 2.0901
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.01 | 0.02001 | kHz | 0.041052 | 2.0453 | 0.041148 | 2.0425
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.01 | 0.02001 | kHz | 0.54767 | 1.1202 | 0.54766 | 1.1202
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.01 | 0.02001 | kHz | 5.7695 | 0.14517 | 5.7695 | 0.14515
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.02 | 0.04501 | kHz | 0.038 | 1 | 0.023094 | 1.0436
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.02 | 0.04501 | kHz | 0.038 | 1 | 0.023358 | 1.0428
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.02 | 0.04501 | kHz | 0.049203 | 0.97804 | 0.049192 | 0.97804
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.02 | 0.04501 | kHz | 0.56864 | 0.35073 | 0.56864 | 0.35068
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.02 | 0.04501 | kHz | 5.7726 | 0.03839 | 5.7726 | 0.038365
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.02282 | 0.47204
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.023285 | 0.47091
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.056197 | 0.39853 | 0.056185 | 0.39851
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.57578 | 0.086645 | 0.57577 | 0.086584
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7734 | 0.0088941 | 5.7734 | 0.0088682
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 1 | 5.01 | kHz | 0.058 | 1.2 | 0.057624 | 1.1588
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 1 | 5.01 | kHz | 0.058 | 1.2 | 0.057825 | 1.1583
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 1 | 5.01 | kHz | 0.073934 | 1.1517 | 0.075029 | 1.1175
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 1 | 5.01 | kHz | 0.56965 | 0.47368 | 0.56963 | 0.47361
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 1 | 5.01 | kHz | 5.7726 | 0.053742 | 5.7726 | 0.053697
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 5 | 10.01 | kHz | 0.12 | 2.3 | 0.11543 | 2.3114
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 5 | 10.01 | kHz | 0.12 | 2.3 | 0.11552 | 2.3112
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 5 | 10.01 | kHz | 0.12471 | 2.2891 | 0.12466 | 2.2892
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 5 | 10.01 | kHz | 0.55922 | 1.4904 | 0.55916 | 1.4904
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 5 | 10.01 | kHz | 5.7697 | 0.21298 | 5.7697 | 0.21288
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 10 | 30 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 10 | 30 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 30 | 30.01 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 30 | 30.01 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 0.01 | 0.04501 | Hz | 33 | 550 | 23.236 | 577.95
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 0.01 | 0.04501 | Hz | 33 | 550.18 | 23.663 | 576.89
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 0.01 | 0.04501 | Hz | 55.234 | 504.47 | 55.135 | 504.68
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 0.01 | 0.04501 | Hz | 577.51 | 123.44 | 577.47 | 123.45
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 0.01 | 0.04501 | Hz | 5799.8 | 12.763 | 577.47 | 12.763
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 350 | 23.254 | 347.52
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 346.2 | 23.771 | 346.32
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 0.045 | 10 | Hz | 58.586 | 276.03 | 58.483 | 276.21
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 0.045 | 10 | Hz | 579.4 | 50.965 | 579.36 | 50.969
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800 | 5.1718 | 5799.9 | 5.1718
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 10 | 20.01 | kHz | 24 | 810 | 23.218 | 808.67
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 10 | 20.01 | kHz | 24 | 807.65 | 23.583 | 807.74
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 10 | 20.01 | kHz | 52.238 | 738.2 | 52.145 | 738.41
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 10 | 20.01 | kHz | 574.82 | 223.39 | 574.77 | 223.41
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 10 | 20.01 | kHz | 5799.5 | 23.686 | 5799.4 | 23.686
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 20 | 50.01 | kHz | 47 | 1200 | 46.389 | 1155.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 20 | 50.01 | kHz | 47 | 1154.8 | 46.618 | 1154.8
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 20 | 50.01 | kHz | 66.007 | 1107.9 | 65.939 | 1108
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 20 | 50.01 | kHz | 571.22 | 452.98 | 571.18 | 452.85
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 20 | 50.01 | kHz | 5799 | 51.02 | 5798.9 | 50.942
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 50 | 100.01 | kHz | 200 | 2700 | 196.5 | 2656.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 50 | 100.01 | kHz | 200 | 2656.2 | 196.56 | 2656.3
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 50 | 100.01 | kHz | 202.89 | 2641.9 | 202.83 | 2642
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 50 | 100.01 | kHz | 582.19 | 1932.6 | 582.1 | 1932.5
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 50 | 100.01 | kHz | 5796.8 | 308.19 | 5796.8 | 308.02
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 100 | 500 | kHz | 390 | 5800 | 381.75 | 5779.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 100 | 500 | kHz | 390 | 5779.3 | 381.8 | 5779.3
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 100 | 500 | kHz | 390 | 5771.9 | 384.93 | 5772
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 100 | 500 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 100 | 500 | kHz | 5783 | 1371.2 | 5783 | 1370.9
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 500 | 501 | kHz | 390 | 5800 | 381.75 | 5779.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 500 | 501 | kHz | 390 | 5779.3 | 381.8 | 5779.3
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 500 | 501 | kHz | 390 | 5771.9 | 384.93 | 5772
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 500 | 501 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 500 | 501 | kHz | 5783 | 1371.2 | 5783 | 1370.9
5502A | Capacitance | 0.033 | 0.11001 | 1e-05 | F | µF* |  |  |  | 120 | 13000 | 115.63 | 12701
5502A | Capacitance | 0.033 | 0.11001 | 0.0001 | F | µF* |  |  |  | 120 | 13000 | 119.5 | 12676
5502A | Capacitance | 0.033 | 0.11001 | 0.001 | F | µF* |  |  |  | 430.99 | 10809 | 430.28 | 10808
5502A | Capacitance | 0.033 | 0.11001 | 0.01 | F | µF* |  |  |  | 5725.6 | 2215.6 | 5725.3 | 2210.1
5502A | DCA | 0.033 | 0.33001 | 1e-06 | A | µA |  |  |  | 3 | 120 | 2.9414 | 115.33
5502A | DCA | 0.033 | 0.33001 | 1e-05 | A | µA |  |  |  | 5.2853 | 109.48 | 5.266 | 109.51
5502A | DCA | 0.033 | 0.33001 | 0.0001 | A | µA |  |  |  | 56.995 | 42.585 | 56.984 | 42.551
5502A | DCA | 0.033 | 0.33001 | 0.001 | A | µA |  |  |  | 579.89 | 4.761 | 579.88 | 4.7411
5502A | DCA | 0.033 | 0.33001 | 0.01 | A | µA |  |  |  | 5800 | 0.4827 | 5800 | 0.47471
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 0 | kHz | 0.25 | 4.6 | 0.23096 | 4.6188
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 0 | kHz | 0.25 | 4.6 | 0.231 | 4.6188
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 0 | kHz | 0.25 | 4.6003 | 0.23368 | 4.6165
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 0 | kHz | 0.46103 | 4.4237 | 0.45975 | 4.4244
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 0 | kHz | 5.6242 | 1.9179 | 5.6235 | 1.9171
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 0 | kHz | 0.28 | 1.8 | 0.23096 | 1.8476
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 0 | kHz | 0.28 | 1.8 | 0.23102 | 1.8476
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 0 | kHz | 0.28 | 1.8 | 0.23545 | 1.844
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 0 | kHz | 0.55279 | 1.6061 | 0.55118 | 1.6064
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 0 | kHz | 5.7507 | 0.38797 | 5.75 | 0.38676
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 0 | kHz | 0.24 | 0.92 | 0.23091 | 0.92432
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 0 | kHz | 0.24 | 0.92 | 0.23101 | 0.92427
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 0 | kHz | 0.24 | 0.92 | 0.23664 | 0.92018
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 0 | kHz | 0.59376 | 0.70044 | 0.59215 | 0.69986
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 0 | kHz | 5.7717 | 0.11872 | 5.771 | 0.11743
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 5 | kHz | 0.24 | 1.2 | 0.22816 | 1.1855
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 5 | kHz | 0.24 | 1.2 | 0.2283 | 1.1854
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 5 | kHz | 0.24 | 1.2 | 0.23356 | 1.1814
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 5 | kHz | 0.58062 | 0.9487 | 0.57901 | 0.94846
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 5 | kHz | 5.767 | 0.17953 | 5.7663 | 0.17826
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 5 | 20 | kHz | 0.27 | 3.5 | 0.22313 | 3.5466
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 5 | 20 | kHz | 0.27 | 3.5331 | 0.22319 | 3.5466
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 5 | 20 | kHz | 0.27 | 3.5338 | 0.22642 | 3.5438
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 5 | 20 | kHz | 0.48562 | 3.329 | 0.48418 | 3.3297
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 5 | 20 | kHz | 5.6818 | 1.2155 | 5.6811 | 1.2146
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44152 | 11.792
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44162 | 11.792
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44312 | 11.791
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 20 | 50 | kHz | 0.53612 | 11.934 | 0.5524 | 11.697
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 20 | 50 | kHz | 5.1678 | 8.3949 | 5.1615 | 8.3731
3458A | ACV | 0.1 | 1 | 1e-07 | V | mV | 2000 | 4000 | kHz | 1.2 | 46 | 0.80072 | 46.326
3458A | ACV | 0.1 | 1 | 1e-06 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.029 | 0.80072 | 46.326
3458A | ACV | 0.1 | 1 | 1e-05 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.029 | 0.80073 | 46.326
3458A | ACV | 0.1 | 1 | 0.0001 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.03 | 0.8015 | 46.325
3458A | ACV | 0.1 | 1 | 0.001 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.034 | 0.83431 | 46.296
3458A | ACV | 0.1 | 1 | 0.01 | V | mV | 2000 | 4000 | kHz | 3.5375 | 43.944 | 3.2985 | 43.945
3458A | ACV | 0.1 | 1 | 1e-07 | V | mV | 4000 | 8000 | kHz | 1.3 | 46 | 0.91606 | 46.326
3458A | ACV | 0.1 | 1 | 1e-06 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91606 | 46.326
3458A | ACV | 0.1 | 1 | 1e-05 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91607 | 46.326
3458A | ACV | 0.1 | 1 | 0.0001 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91665 | 46.325
3458A | ACV | 0.1 | 1 | 0.001 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.051 | 0.94915 | 46.296
3458A | ACV | 0.1 | 1 | 0.01 | V | mV | 4000 | 8000 | kHz | 3.6131 | 43.982 | 3.378 | 43.984
3458A | ACV | 0.1 | 1 | 1e-07 | V | mV | 8000 | 10000 | kHz | 4.6 | 170 | 1.146 | 173.36
3458A | ACV | 0.1 | 1 | 1e-06 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.146 | 173.36
3458A | ACV | 0.1 | 1 | 1e-05 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.146 | 173.36
3458A | ACV | 0.1 | 1 | 0.0001 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.1461 | 173.36
3458A | ACV | 0.1 | 1 | 0.001 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.156 | 173.35
3458A | ACV | 0.1 | 1 | 0.01 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.23 | 2.0278 | 172.49
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 590 | 17000 | 108.37 | 17479
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 108.37 | 17479
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 108.38 | 17479
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 109.37 | 17478
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 1000 | 2000 | kHz | 590 | 17163 | 204.79 | 17392
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 1000 | 2000 | kHz | 4683.4 | 13828 | 4325.4 | 13829
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 0 | 0 | kHz | 47 | 82 | 80.62 | 80.62
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 0 | 0 | kHz | 47 | 82 | 80.618 | 80.618
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 0 | 0 | kHz | 47 | 82 | 80.426 | 80.426
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 0 | 0 | kHz | 72.743 | 66.872 | 66.849 | 66.849
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 0 | 0 | kHz | 578.69 | 12.536 | 12.498 | 12.498
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 0 | 0 | kHz | 5773.7 | 1.2831 | 1.2674 | 1.2674
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 0 | 1 | kHz | 24 | 82 | 23.524 | 80.445
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 0 | 1 | kHz | 24 | 82 | 23.53 | 80.441
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 0 | 1 | kHz | 24 | 82 | 24.088 | 80.041
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 0 | 1 | kHz | 59.943 | 59.049 | 59.9 | 59.024
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 0 | 1 | kHz | 577.3 | 9.3986 | 577.28 | 9.3602
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 0 | 1 | kHz | 5773.6 | 0.9599 | 5772.7 | 0.94416
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.438 | 161.34
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.442 | 161.34
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.893 | 160.98
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 1 | 20 | kHz | 56.377 | 137.35 | 56.262 | 137.33
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 1 | 20 | kHz | 575.7 | 30.683 | 575.66 | 30.548
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.4 | 3.1917 | 5770.6 | 0.31603
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.379 | 346.14
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.383 | 346.14
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.693 | 345.87
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 20 | 50 | kHz | 49.713 | 324.48 | 49.371 | 324.63
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 20 | 50 | kHz | 568.73 | 117.17 | 568.57 | 116.91
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 20 | 50 | kHz | 5772.6 | 12.94 | 5762.2 | 12.802
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 50 | 100 | kHz | 27 | 940 | 23.141 | 938.86
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 50 | 100 | kHz | 27 | 935.56 | 23.148 | 938.85
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 50 | 100 | kHz | 27 | 935.62 | 23.298 | 938.72
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 50 | 100 | kHz | 38.506 | 925.39 | 37.913 | 925.82
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 50 | 100 | kHz | 530.37 | 592.16 | 529.89 | 592.07
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 50 | 100 | kHz | 5766.2 | 87.392 | 5715.5 | 87.123
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.46 | 3463.3
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.47 | 3463.3
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.55 | 3463.2
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 100 | 300 | kHz | 120 | 3500 | 120.41 | 3459.8
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 100 | 300 | kHz | 420.55 | 3205.9 | 419.31 | 3206.7
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 100 | 300 | kHz | 5681.4 | 1112.8 | 5471.4 | 1112.4
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 111.87 | 11674
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 300 | 1000 | kHz | 212.67 | 11907 | 246.94 | 11553
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 300 | 1000 | kHz | 5113.2 | 8011.6 | 4758 | 8011.9
3458A | DCA | 0.1 | 1 | 1e-07 | A | µA |  |  |  | 12 | 130 | 11.536 | 127.22
3458A | DCA | 0.1 | 1 | 1e-06 | A | µA |  |  |  | 12 | 130 | 11.548 | 127.21
3458A | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 12 | 130 | 12.277 | 126.6
3458A | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 52.949 | 97.342 | 52.884 | 97.405
3458A | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 576.12 | 17.674 | 576.1 | 17.7
3458A | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5773.4 | 1.7853 | 5771.8 | 1.7958
3458A | DCV | 0.1 | 1 | 1e-08 | V | µV |  |  |  | 0.44 | 6 | 0.43733 | 5.9926
3458A | DCV | 0.1 | 1 | 1e-07 | V | µV |  |  |  | 0.44094 | 5.99 | 0.43909 | 5.9911
3458A | DCV | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 0.60901 | 5.85 | 0.60244 | 5.8536
3458A | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 5.589 | 3.0792 | 5.5845 | 3.0749
3458A | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 57.973 | 0.3881 | 57.971 | 0.38452
3458A | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 580 | 0.040018 | 579.96 | 0.038571
34420A | DCVmeas | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 4.7 | 41 | 4.6518 | 40.782
34420A | DCVmeas | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 6.5523 | 39.255 | 6.5389 | 39.26
34420A | DCVmeas | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 56.989 | 16.705 | 56.981 | 16.693
34420A | DCVmeas | 0.1 | 1 | 0.001 | V | µV |  |  |  | 579.88 | 1.9084 | 579.88 | 1.9012
34420A | DCVmeas | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5800 | 0.19329 | 5800 | 0.19042
34420A | DCVmeas | 0.1 | 1 | 0.1 | V | µV |  |  |  | 58000 | 0.020188 | 58000 | 0.019042
Transmille | ACA | 0.1 | 1.04 | 1e-05 | A | mA | 0.01 | 2 | KHz | 1 | 1.5 | 0.99589 | 1.49
Transmille | ACA | 0.1 | 1.04 | 0.0001 | A | mA | 0.01 | 2 | KHz | 1 | 1.49 | 0.99738 | 1.49
Transmille | ACA | 0.1 | 1.04 | 0.001 | A | mA | 0.01 | 2 | KHz | 1.1829 | 1.41 | 1.1409 | 1.41
Transmille | ACA | 0.1 | 1.04 | 0.01 | A | mA | 0.01 | 2 | KHz | 5.9116 | 0.478 | 5.8383 | 0.476
Transmille | ACA | 0.1 | 1.04 | 0.1 | A | mA | 0.01 | 2 | KHz | 57.947 | 0.33 | 57.91 | 0.329
Transmille | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 350 | 940 | 347.15 | 937
Transmille | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 351.19 | 935 | 351.16 | 934
Transmille | ACV | 0.1 | 1 | 0.001 | V | µV | 0.01 | 2 | KHz | 650.74 | 758 | 650.66 | 757
Transmille | ACV | 0.1 | 1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5776.6 | 138 | 5776.5 | 138
Transmille | ACV | 0.1 | 1 | 0.1 | V | µV | 0.01 | 2 | KHz | 57736 | 14.1 | 57736 | 14
Transmille | ACV | 0.1 | 1 | 1e-05 | V | µV | 2 | 20 | KHz | 850 | 1700 | 808.55 | 1740
Transmille | ACV | 0.1 | 1 | 0.0001 | V | µV | 2 | 20 | KHz | 850 | 1740 | 810.33 | 1740
Transmille | ACV | 0.1 | 1 | 0.001 | V | µV | 2 | 20 | KHz | 976.25 | 1640 | 975.89 | 1640
Transmille | ACV | 0.1 | 1 | 0.01 | V | µV | 2 | 20 | KHz | 5806.3 | 505 | 5806.1 | 504
Transmille | ACV | 0.1 | 1 | 0.1 | V | µV | 2 | 20 | KHz | 57739 | 53.4 | 57739 | 53.1
Transmille | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 180 | 370 | 172.98 | 365
Transmille | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 181.28 | 360 | 181.24 | 360
Transmille | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 594.75 | 195 | 594.71 | 195
Transmille | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5775 | 23.6 | 5775 | 23.6
Transmille | DCA | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 2.38 | 57736 | 2.36
Transmille | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 36 | 92 | 35.885 | 91.6
Transmille | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 65.685 | 74.3 | 65.574 | 74.3
Transmille | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 577.78 | 13.6 | 577.73 | 13.5
Transmille | DCV | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5773.6 | 1.39 | 5773.6 | 1.37
Transmille | DCV | 0.1 | 1 | 0.1 | V | µV |  |  |  | 57736 | 0.144 | 57736 | 0.137
5502A | ACW | 0.109 | 1.0901 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.00078 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.00094283 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.001 | WAC | WAC |  |  |  | 0 | 0.0053846 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.01 | WAC | WAC |  |  |  | 0 | 0.053267 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.1 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.00057856 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.001 | WDC | WDC |  |  |  | 0 | 0.0053324 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.01 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.1 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | Capacitance | 0.11 | 0.1101 | 1e-05 | F | µF |  |  |  | 120 | 13000 | 115.63 | 12701
5502A | Capacitance | 0.11 | 0.1101 | 0.0001 | F | µF |  |  |  | 120 | 13000 | 119.5 | 12676
5502A | Capacitance | 0.11 | 0.1101 | 0.001 | F | µF |  |  |  | 430.99 | 10809 | 430.28 | 10808
5502A | Capacitance | 0.11 | 0.1101 | 0.01 | F | µF |  |  |  | 5725.6 | 2215.6 | 5725.3 | 2210.1
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11617 | 2.0784
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11882 | 2.0766
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.35291 | 1.9255 | 0.35284 | 1.9254
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 5.6471 | 0.55145 | 5.6471 | 0.55121
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 57.722 | 0.057755 | 57.722 | 0.057649
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.12 | 0.58 | 0.11734 | 0.57739
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.12267 | 0.57758 | 0.124 | 0.57333
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.52825 | 0.38235 | 0.52823 | 0.382
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7643 | 0.053023 | 5.7643 | 0.05276
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.735 | 0.005408 | 57.735 | 0.0053018
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 1 | 5.01 | kHz | 1.2 | 6.9 | 1.1549 | 6.9282
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 1 | 5.01 | kHz | 1.2 | 6.9 | 1.1555 | 6.9278
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 1 | 5.01 | kHz | 1.2164 | 6.8903 | 1.2155 | 6.8904
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 1 | 5.01 | kHz | 5.1011 | 4.9149 | 5.0998 | 4.9136
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 1 | 5.01 | kHz | 57.598 | 0.72965 | 57.598 | 0.72819
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7734 | 28.868
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7742 | 28.867
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7874 | 28.859
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 5 | 10 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 5 | 10 | kHz | 55.816 | 11.862 | 55.815 | 11.86
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7734 | 28.868
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7742 | 28.867
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7874 | 28.859
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 10 | 10.01 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 10 | 10.01 | kHz | 55.816 | 11.862 | 55.815 | 11.86
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 0.01 | 0.04501 | Hz | 0.11 | 0.56 | 0.10391 | 0.56221
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 0.01 | 0.04501 | Hz | 0.11 | 0.56033 | 0.10921 | 0.56037
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 0.48274 | 0.46448 | 0.4824 | 0.4645
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 5.7719 | 0.10338 | 5.7717 | 0.10332
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 57.997 | 0.010643 | 57.997 | 0.010612
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 0.045 | 10 | Hz | 0.07 | 0.35 | 0.069477 | 0.34756
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.079443 | 0.34499 | 0.079136 | 0.34505
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.52675 | 0.24887 | 0.5264 | 0.24885
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.7894 | 0.041576 | 5.7892 | 0.041503
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.999 | 0.0042263 | 57.999 | 0.004196
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 10 | 20.01 | kHz | 0.07 | 0.81 | 0.069435 | 0.80878
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 10 | 20.01 | kHz | 0.075008 | 0.80729 | 0.074829 | 0.80733
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 10 | 20.01 | kHz | 0.43427 | 0.71669 | 0.43391 | 0.71674
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 10 | 20.01 | kHz | 5.7428 | 0.2035 | 5.7426 | 0.20344
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 10 | 20.01 | kHz | 57.994 | 0.021456 | 57.994 | 0.021426
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 20 | 50.01 | kHz | 0.07 | 1.2 | 0.069506 | 1.1559
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 20 | 50.01 | kHz | 0.073711 | 1.1548 | 0.073371 | 1.1549
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 20 | 50.01 | kHz | 0.38065 | 1.0748 | 0.37987 | 1.0749
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 20 | 50.01 | kHz | 5.6886 | 0.39166 | 5.6883 | 0.39154
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 20 | 50.01 | kHz | 57.988 | 0.043219 | 57.988 | 0.043148
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 50 | 100.01 | kHz | 0.24 | 2.7 | 0.23108 | 2.6571
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 50 | 100.01 | kHz | 0.24 | 2.6566 | 0.23256 | 2.6567
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 50 | 100.01 | kHz | 0.38846 | 2.6152 | 0.38733 | 2.6154
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 50 | 100.01 | kHz | 5.3723 | 1.6168 | 5.3715 | 1.6167
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 50 | 100.01 | kHz | 57.935 | 0.23029 | 57.935 | 0.23012
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 100 | 500 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 100 | 500 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 100 | 500 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 100 | 500 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 100 | 500 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 500 | 501 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 500 | 501 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 500 | 501 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 500 | 501 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 500 | 501 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
5502A | DCA | 0.33 | 3.0001 | 1e-05 | A | mA |  |  |  | 0.052 | 0.44 | 0.051 | 0.43873
5502A | DCA | 0.33 | 3.0001 | 0.0001 | A | mA |  |  |  | 0.060249 | 0.43614 | 0.060204 | 0.43607
5502A | DCA | 0.33 | 3.0001 | 0.001 | A | mA |  |  |  | 0.50421 | 0.32729 | 0.50424 | 0.32696
5502A | DCA | 0.33 | 3.0001 | 0.01 | A | mA |  |  |  | 5.784 | 0.058643 | 5.7841 | 0.058299
5502A | DCA | 0.33 | 3.0001 | 0.1 | A | mA |  |  |  | 57.998 | 0.0060523 | 57.998 | 0.0059106
5502A | DCV | 0.33 | 3.3 | 1e-06 | V | µV |  |  |  | 5.9 | 58 | 5.7708 | 57.787
5502A | DCV | 0.33 | 3.3 | 1e-05 | V | µV |  |  |  | 6.8066 | 57.529 | 6.4957 | 57.593
5502A | DCV | 0.33 | 3.3 | 0.0001 | V | µV |  |  |  | 47.82 | 47.701 | 47.342 | 47.73
5502A | DCV | 0.33 | 3.3 | 0.001 | V | µV |  |  |  | 577.17 | 10.824 | 576.99 | 10.721
5502A | DCV | 0.33 | 3.3 | 0.01 | V | µV |  |  |  | 5799.8 | 1.1473 | 5799.7 | 1.1022
5502A | DCV | 0.33 | 3.3 | 0.1 | V | µV |  |  |  | 58000 | 0.12833 | 5799.8 | 0.11025
5502A | ACA 10 turn | 1 | 11.001 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 3.3 | 2.7 | 0.97318 | 2.873
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 4.3265 | 2.6149 | 4.3204 | 2.6148
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 57.287 | 0.84059 | 57.285 | 0.83952
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 579.92 | 0.090643 | 579.92 | 0.090136
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 5800 | 0.0092241 | 5800 | 0.0090207
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 58000 | 0.00098346 | 58000 | 0.00090208
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.66 | 2.1 | 0.65508 | 2.0543
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 4.6183 | 1.7587 | 4.6169 | 1.7583
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.618 | 0.44323 | 57.618 | 0.44244
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 0.045 | 1 | kHz | 579.96 | 0.046302 | 579.96 | 0.045959
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 0.045 | 1 | kHz | 5800 | 0.004735 | 5800 | 0.0045977
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.405 | 7.1882
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 12.361 | 7.1187 | 12.349 | 7.1191
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 56.262 | 4.6574 | 56.248 | 4.6564
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 1 | 5.01 | kHz | 579.63 | 0.67288 | 579.63 | 0.67163
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 1 | 5.01 | kHz | 5800 | 0.068091 | 5800 | 0.067582
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.703 | 28.929
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 28.914 | 57.905 | 28.915
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 5 | 10 | kHz | 76.668 | 27.611 | 76.638 | 27.612
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 5 | 10 | kHz | 575.98 | 10.476 | 575.96 | 10.474
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 5 | 10 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.703 | 28.929
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 58 | 28.914 | 57.905 | 28.915
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 76.668 | 27.611 | 76.638 | 27.612
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 10 | 10.01 | kHz | 575.98 | 10.476 | 575.96 | 10.474
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 10 | 10.01 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
5502A | DCA 10 turn | 1 | 11.001 | 0.0001 | A | mA |  |  |  | 0.082 | 2 | 0.080875 | 1.9947
5502A | DCA 10 turn | 1 | 11.001 | 0.001 | A | mA |  |  |  | 0.16836 | 1.9875 | 0.16627 | 1.9876
5502A | DCA 10 turn | 1 | 11.001 | 0.01 | A | mA |  |  |  | 4.4751 | 1.6631 | 4.4718 | 1.6632
5502A | DCA 10 turn | 1 | 11.001 | 0.1 | A | mA |  |  |  | 57.372 | 0.40227 | 57.371 | 0.40202
5502A | DCA 10 turn | 1 | 11.001 | 1 | A | mA |  |  |  | 577.32 | 0.041721 | 577.32 | 0.041613
3458A | ACA | 1 | 10 | 1e-05 | A | mA | 0 | 0 | kHz | 0.55 | 0.12 | 0.51999 | 0.093087
3458A | ACA | 1 | 10 | 0.0001 | A | mA | 0 | 0 | kHz | 0.55 | 0.12 | 0.52375 | 0.092861
3458A | ACA | 1 | 10 | 0.001 | A | mA | 0 | 0 | kHz | 0.75909 | 0.099091 | 0.76542 | 0.079775
3458A | ACA | 1 | 10 | 0.01 | A | mA | 0 | 0 | kHz | 5.7992 | 0.016313 | 5.7916 | 0.016333
3458A | ACA | 1 | 10 | 0.1 | A | mA | 0 | 0 | kHz | 57.741 | 0.0016634 | 57.738 | 0.0016635
3458A | ACA | 1 | 10 | 1e-05 | A | mA | 0 | 1 | kHz | 0.46 | 0.17 | 0.49897 | 0.12063
3458A | ACA | 1 | 10 | 0.0001 | A | mA | 0 | 1 | kHz | 0.46 | 0.16806 | 0.50269 | 0.12039
3458A | ACA | 1 | 10 | 0.001 | A | mA | 0 | 1 | kHz | 0.72094 | 0.14191 | 0.74419 | 0.10576
3458A | ACA | 1 | 10 | 0.01 | A | mA | 0 | 1 | kHz | 5.7925 | 0.023676 | 5.7849 | 0.023706
3458A | ACA | 1 | 10 | 0.1 | A | mA | 0 | 1 | kHz | 57.74 | 0.002428 | 57.737 | 0.0024282
3458A | ACA | 1 | 10 | 1e-05 | A | mA | 1 | 5 | kHz | 1.4 | 0.66 | 1.326 | 0.66218
3458A | ACA | 1 | 10 | 0.0001 | A | mA | 1 | 5 | kHz | 1.4 | 0.66168 | 1.3279 | 0.66196
3458A | ACA | 1 | 10 | 0.001 | A | mA | 1 | 5 | kHz | 1.4676 | 0.64572 | 1.4594 | 0.64665
3458A | ACA | 1 | 10 | 0.01 | A | mA | 1 | 5 | kHz | 5.82 | 0.34986 | 5.8127 | 0.35015
3458A | ACA | 1 | 10 | 0.1 | A | mA | 1 | 5 | kHz | 57.725 | 0.044886 | 57.722 | 0.044889
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 1000 | 2000 | kHz | 6.1 | 17 | 1.1338 | 17.487
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1338 | 17.487
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1338 | 17.487
3458A | ACV | 1 | 10 | 0.001 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1448 | 17.486
3458A | ACV | 1 | 10 | 0.01 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.167 | 2.0915 | 17.401
3458A | ACV | 1 | 10 | 0.1 | V | mV | 1000 | 2000 | kHz | 46.847 | 13.84 | 43.262 | 13.841
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 2000 | 4000 | kHz | 12 | 46 | 8.0599 | 46.328
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.06 | 46.328
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.06 | 46.328
3458A | ACV | 1 | 10 | 0.001 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.0603 | 46.328
3458A | ACV | 1 | 10 | 0.01 | V | mV | 2000 | 4000 | kHz | 12 | 46.042 | 8.3953 | 46.298
3458A | ACV | 1 | 10 | 0.1 | V | mV | 2000 | 4000 | kHz | 35.391 | 43.95 | 33.018 | 43.95
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 4000 | 8000 | kHz | 15 | 46 | 9.1632 | 46.576
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1632 | 46.576
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1632 | 46.576
3458A | ACV | 1 | 10 | 0.001 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1658 | 46.576
3458A | ACV | 1 | 10 | 0.01 | V | mV | 4000 | 8000 | kHz | 15 | 46.329 | 9.49 | 46.547
3458A | ACV | 1 | 10 | 0.1 | V | mV | 4000 | 8000 | kHz | 36.034 | 44.241 | 33.704 | 44.241
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 8000 | 10000 | kHz | 46 | 170 | 11.528 | 173.36
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.529 | 173.36
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.529 | 173.36
3458A | ACV | 1 | 10 | 0.001 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.584 | 173.35
3458A | ACV | 1 | 10 | 0.01 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.629 | 173.35
3458A | ACV | 1 | 10 | 0.1 | V | mV | 8000 | 10000 | kHz | 46 | 170.24 | 20.344 | 172.48
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 0 | 0 | kHz | 0.46 | 0.092 | 0.45053 | 0.092199
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 0 | 0 | kHz | 0.46 | 0.091968 | 0.45057 | 0.092197
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 0 | 0 | kHz | 0.46 | 0.092019 | 0.4538 | 0.091994
3458A | ACV | 1 | 10 | 0.001 | V | mV | 0 | 0 | kHz | 0.71619 | 0.077546 | 0.71499 | 0.077402
3458A | ACV | 1 | 10 | 0.01 | V | mV | 0 | 0 | kHz | 5.7845 | 0.015263 | 5.784 | 0.01505
3458A | ACV | 1 | 10 | 0.1 | V | mV | 0 | 0 | kHz | 57.737 | 0.0016174 | 57.723 | 0.001529
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 0 | 1 | kHz | 0.25 | 0.084 | 0.22625 | 0.085545
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 0 | 1 | kHz | 0.25 | 0.084438 | 0.22632 | 0.08554
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 0 | 1 | kHz | 0.25 | 0.084683 | 0.23198 | 0.085127
3458A | ACV | 1 | 10 | 0.001 | V | mV | 0 | 1 | kHz | 0.59428 | 0.063506 | 0.59284 | 0.06333
3458A | ACV | 1 | 10 | 0.01 | V | mV | 0 | 1 | kHz | 5.7723 | 0.010443 | 5.7717 | 0.010227
3458A | ACV | 1 | 10 | 0.1 | V | mV | 0 | 1 | kHz | 57.736 | 0.0011207 | 57.726 | 0.0010323
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 1 | 20 | kHz | 0.28 | 0.16 | 0.22835 | 0.16434
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 1 | 20 | kHz | 0.28 | 0.16134 | 0.22837 | 0.16434
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 1 | 20 | kHz | 0.28 | 0.16156 | 0.23295 | 0.16397
3458A | ACV | 1 | 10 | 0.001 | V | mV | 1 | 20 | kHz | 0.56386 | 0.14001 | 0.55818 | 0.14006
3458A | ACV | 1 | 10 | 0.01 | V | mV | 1 | 20 | kHz | 5.7578 | 0.031829 | 5.7555 | 0.031388
3458A | ACV | 1 | 10 | 0.1 | V | mV | 1 | 20 | kHz | 57.735 | 0.0034135 | 57.705 | 0.003222
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.22967 | 0.34776
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.22973 | 0.34775
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.23282 | 0.34749
3458A | ACV | 1 | 10 | 0.001 | V | mV | 20 | 50 | kHz | 0.50358 | 0.32533 | 0.49042 | 0.32615
3458A | ACV | 1 | 10 | 0.01 | V | mV | 20 | 50 | kHz | 5.6913 | 0.11808 | 5.6847 | 0.11766
3458A | ACV | 1 | 10 | 0.1 | V | mV | 20 | 50 | kHz | 57.728 | 0.013146 | 57.621 | 0.01289
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 50 | 100 | kHz | 0.27 | 0.94 | 0.22603 | 0.94121
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 50 | 100 | kHz | 0.27 | 0.93759 | 0.22604 | 0.94121
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 50 | 100 | kHz | 0.27 | 0.93768 | 0.22761 | 0.94107
3458A | ACV | 1 | 10 | 0.001 | V | mV | 50 | 100 | kHz | 0.38942 | 0.92692 | 0.37412 | 0.92813
3458A | ACV | 1 | 10 | 0.01 | V | mV | 50 | 100 | kHz | 5.3092 | 0.59368 | 5.2965 | 0.59386
3458A | ACV | 1 | 10 | 0.1 | V | mV | 50 | 100 | kHz | 57.665 | 0.087823 | 57.153 | 0.087462
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1359 | 3.4831
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1359 | 3.4831
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1365 | 3.483
3458A | ACV | 1 | 10 | 0.001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1754 | 3.4796
3458A | ACV | 1 | 10 | 0.01 | V | mV | 100 | 300 | kHz | 4.1895 | 3.2244 | 4.1679 | 3.2259
3458A | ACV | 1 | 10 | 0.1 | V | mV | 100 | 300 | kHz | 56.807 | 1.1229 | 54.694 | 1.1224
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 300 | 1000 | kHz | 0.83 | 12 | 1.127 | 11.688
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 300 | 1000 | kHz | 0.83 | 11.988 | 1.127 | 11.688
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 300 | 1000 | kHz | 0.83 | 11.988 | 1.1272 | 11.688
3458A | ACV | 1 | 10 | 0.001 | V | mV | 300 | 1000 | kHz | 0.84663 | 11.986 | 1.141 | 11.687
3458A | ACV | 1 | 10 | 0.01 | V | mV | 300 | 1000 | kHz | 2.2239 | 11.849 | 2.4896 | 11.566
3458A | ACV | 1 | 10 | 0.1 | V | mV | 300 | 1000 | kHz | 51.132 | 8.0252 | 47.575 | 8.0258
3458A | DCA | 1 | 20 | 1e-06 | A | mA |  |  |  | 0.068 | 0.0036 | 0.043 | 0.0047621
3458A | DCA | 1 | 20 | 1e-05 | A | mA |  |  |  | 0.068 | 0.0036 | 0.043318 | 0.0047475
3458A | DCA | 1 | 20 | 0.0001 | A | mA |  |  |  | 0.070893 | 0.0038965 | 0.068928 | 0.0038258
3458A | DCA | 1 | 20 | 0.001 | A | mA |  |  |  | 0.57834 | 0.00081073 | 0.57809 | 0.0006992
3458A | DCA | 1 | 20 | 0.01 | A | mA |  |  |  | 5.7739 | 0.00011677 | 5.7736 | 7.0888e-05
3458A | DCA | 1 | 20 | 0.1 | A | mA |  |  |  | 57.736 | 2.5449e-05 | 57.735 | 7.09e-06
3458A | DCV | 1 | 10 | 1e-07 | V | µV |  |  |  | 1.2 | 5.3 | 0.75905 | 5.3437
3458A | DCV | 1 | 10 | 1e-06 | V | µV |  |  |  | 1.2 | 5.3409 | 0.78898 | 5.341
3458A | DCV | 1 | 10 | 1e-05 | V | µV |  |  |  | 3.301 | 5.1205 | 3.2984 | 5.1207
3458A | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 55.981 | 2.3401 | 55.98 | 2.34
3458A | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 579.76 | 0.27724 | 579.76 | 0.27716
3458A | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5800 | 0.027808 | 5799.7 | 0.027777
34420A | DCVmeas | 1 | 10 | 1e-05 | V | µV |  |  |  | 47 | 35 | 46.376 | 34.813
34420A | DCVmeas | 1 | 10 | 0.0001 | V | µV |  |  |  | 66.458 | 33.232 | 66.382 | 33.232
34420A | DCVmeas | 1 | 10 | 0.001 | V | µV |  |  |  | 572.81 | 12.879 | 572.76 | 12.867
34420A | DCVmeas | 1 | 10 | 0.01 | V | µV |  |  |  | 5799.2 | 1.4325 | 5799.1 | 1.4259
34420A | DCVmeas | 1 | 10 | 0.1 | V | µV |  |  |  | 58000 | 0.1454 | 58000 | 0.14276
34420A | DCVmeas | 1 | 10 | 1 | V | µV |  |  |  | 580000 | 0.015331 | 580000 | 0.014276
Transmille | ACA | 1 | 10.4 | 0.0001 | A | mA | 0.01 | 2 | KHz | 18 | 1.4 | 17.128 | 1.4
Transmille | ACA | 1 | 10.4 | 0.001 | A | mA | 0.01 | 2 | KHz | 18 | 1.4 | 17.138 | 1.4
Transmille | ACA | 1 | 10.4 | 0.01 | A | mA | 0.01 | 2 | KHz | 18.047 | 1.36 | 18.045 | 1.36
Transmille | ACA | 1 | 10.4 | 0.1 | A | mA | 0.01 | 2 | KHz | 60.057 | 0.582 | 60.055 | 0.581
Transmille | ACA | 1 | 10.4 | 1 | A | mA | 0.01 | 2 | KHz | 579.27 | 0.343 | 579.27 | 0.342
Transmille | ACV | 1 | 10 | 0.0001 | V | mV | 0.01 | 2 | KHz | 3.5 | 0.94 | 3.4755 | 0.937
Transmille | ACV | 1 | 10 | 0.001 | V | mV | 0.01 | 2 | KHz | 3.5127 | 0.935 | 3.5116 | 0.934
Transmille | ACV | 1 | 10 | 0.01 | V | mV | 0.01 | 2 | KHz | 6.5091 | 0.758 | 6.5066 | 0.757
Transmille | ACV | 1 | 10 | 0.1 | V | mV | 0.01 | 2 | KHz | 57.766 | 0.139 | 57.765 | 0.138
Transmille | ACV | 1 | 10 | 1 | V | mV | 0.01 | 2 | KHz | 577.36 | 0.0144 | 577.36 | 0.014
Transmille | ACV | 1 | 10 | 0.0001 | V | mV | 2 | 20 | KHz | 8.2 | 1.8 | 8.1347 | 1.75
Transmille | ACV | 1 | 10 | 0.001 | V | mV | 2 | 20 | KHz | 8.2 | 1.75 | 8.1525 | 1.75
Transmille | ACV | 1 | 10 | 0.01 | V | mV | 2 | 20 | KHz | 9.8126 | 1.65 | 9.799 | 1.65
Transmille | ACV | 1 | 10 | 0.1 | V | mV | 2 | 20 | KHz | 58.073 | 0.513 | 58.064 | 0.512
Transmille | ACV | 1 | 10 | 1 | V | mV | 2 | 20 | KHz | 577.39 | 0.0545 | 577.39 | 0.054
Transmille | DCA | 1 | 10.2 | 0.0001 | A | µA |  |  |  | 2.4 | 0.59 | 2.3095 | 0.58
Transmille | DCA | 1 | 10.2 | 0.001 | A | µA |  |  |  | 2.4 | 0.631 | 2.3701 | 0.576
Transmille | DCA | 1 | 10.2 | 0.01 | A | µA |  |  |  | 5.8935 | 0.563 | 6.0655 | 0.39
Transmille | DCA | 1 | 10.2 | 0.1 | A | µA |  |  |  | 57.625 | 0.183 | 57.752 | 0.0555
Transmille | DCA | 1 | 10.2 | 1 | A | µA |  |  |  | 577.31 | 0.0574 | 577.36 | 0.00558
Transmille | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 360 | 92 | 354.54 | 91.8
Transmille | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 653.72 | 74.4 | 653.17 | 74.3
Transmille | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5777.2 | 13.6 | 5777 | 13.5
Transmille | DCV | 1 | 10 | 0.1 | V | µV |  |  |  | 57736 | 1.4 | 57736 | 13.5
Transmille | DCV | 1 | 10 | 1 | V | µV |  |  |  | 577360 | 0.148 | 577360 | 13.5
Transmille | Frequency | 1 | 100000 | 1 | Hz | Hz |  |  |  | 0.87 | 1.6e-05 | 0.81638 | 1.64e-05
Transmille | Frequency | 1 | 100000 | 10 | Hz | Hz |  |  |  | 5.8379 | 4.82e-06 | 5.8243 | 4.68e-06
Transmille | Frequency | 1 | 100000 | 100 | Hz | Hz |  |  |  | 58.005 | 3.17e-06 | 58 | 3.11e-06
Transmille | Frequency | 1 | 100000 | 1000 | Hz | Hz |  |  |  | 579.98 | 2.65e-05 | 579.97 | 2.65e-05
Transmille | Frequency | 1 | 100000 | 10000 | Hz | Hz |  |  |  | 5799.7 | 0.000265 | 5799.7 | 0.000265
Transmille | Insulation DCV | 1 | 100 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.4054 | 0.0105
Transmille | Insulation DCV | 1 | 100 | 10 | V | V |  |  |  | 5.9149 | 0.00383 | 5.9069 | 0.00367
Transmille | Insulation DCV | 1 | 100 | 100 | V | V |  |  |  | 57.918 | 0.00337 | 57.915 | 0.0033
5502A | ACW | 1.09 | 10.9001 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.0013 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.001 | WAC | WAC |  |  |  | 0 | 0.0013347 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0054669 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.053275 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 1 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00057825 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.0053322 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 1 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11589 | 2.0787
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11676 | 2.0785
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.18892 | 2.077 | 0.20868 | 2.0565
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 4.9071 | 1.2281 | 4.9043 | 1.2265
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 57.613 | 0.15838 | 57.613 | 0.15706
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11619 | 0.69369
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11868 | 0.6931
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.35192 | 0.64234 | 0.34551 | 0.6421
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.6463 | 0.18057 | 5.6448 | 0.17758
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.722 | 0.019812 | 57.722 | 0.018474
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1548 | 6.9283
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1552 | 6.9282
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1804 | 6.9223
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 3.4547 | 6.4116 | 3.4521 | 6.4115
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 56.451 | 1.7722 | 56.451 | 1.771
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.773 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7731 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.78 | 28.866
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 5 | 10 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 5 | 10 | kHz | 45.662 | 21.093 | 45.659 | 21.092
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.773 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7731 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.78 | 28.866
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 45.662 | 21.093 | 45.659 | 21.092
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 0.045 | 0.10001 | kHz | 2.4 | 0.69 | 2.3298 | 0.69373
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 0.045 | 0.10001 | kHz | 2.4 | 0.69 | 2.3709 | 0.69112
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 5.6644 | 0.53041 | 5.6622 | 0.53026
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 57.647 | 0.085551 | 57.646 | 0.085343
5502A | ACA | 3 | 11.001 | 1 | A | mA | 0.045 | 0.10001 | kHz | 577.35 | 0.0086938 | 577.35 | 0.0086087
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.324 | 1.1553
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.3551 | 1.1531
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 0.1 | 1 | kHz | 5.1972 | 0.99041 | 5.1951 | 0.99036
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 0.1 | 1 | kHz | 57.412 | 0.20433 | 57.411 | 0.20413
5502A | ACA | 3 | 11.001 | 1 | A | mA | 0.1 | 1 | kHz | 577.32 | 0.020884 | 577.32 | 0.020799
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3103 | 34.641
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3112 | 34.641
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.5084 | 34.627
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 1 | 5 | kHz | 20.888 | 33.346 | 20.867 | 33.347
5502A | ACA | 3 | 11.001 | 1 | A | mA | 1 | 5 | kHz | 547.31 | 13.25 | 547.3 | 13.249
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.3103 | 34.641
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.3112 | 34.641
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.5084 | 34.627
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 5 | 5.01 | kHz | 20.888 | 33.346 | 20.867 | 33.347
5502A | ACA | 3 | 11.001 | 1 | A | mA | 5 | 5.01 | kHz | 547.31 | 13.25 | 547.3 | 13.249
5502A | DCA | 3 | 11.001 | 0.0001 | A | mA |  |  |  | -0.024 | 0.92 | -0.024619 | 0.91523
5502A | DCA | 3 | 11.001 | 0.001 | A | mA |  |  |  | 0.055271 | 0.90973 | 0.052344 | 0.90974
5502A | DCA | 3 | 11.001 | 0.01 | A | mA |  |  |  | 4.4629 | 0.64947 | 4.4597 | 0.64887
5502A | DCA | 3 | 11.001 | 0.1 | A | mA |  |  |  | 57.764 | 0.1006 | 57.764 | 0.099909
5502A | DCV | 3 | 33.0001 | 1e-05 | V | mV |  |  |  | 0.059 | 0.058 | 0.058348 | 0.057919
5502A | DCV | 3 | 33.0001 | 0.0001 | V | mV |  |  |  | 0.067758 | 0.057661 | 0.064652 | 0.057725
5502A | DCV | 3 | 33.0001 | 0.001 | V | mV |  |  |  | 0.47784 | 0.047829 | 0.47304 | 0.047859
5502A | DCV | 3 | 33.0001 | 0.01 | V | mV |  |  |  | 5.7715 | 0.010868 | 5.7698 | 0.010764
5502A | DCV | 3 | 33.0001 | 0.1 | V | mV |  |  |  | 57.998 | 0.0011519 | 57.997 | 0.0011068
5502A | DCV | 3 | 33.0001 | 1 | V | mV |  |  |  | 580 | 0.00012879 | 580 | 0.00011071
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 0.01 | 0.04501 | Hz | 0.93 | 0.58 | 0.92292 | 0.5783
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 0.98726 | 0.57661 | 0.98668 | 0.57662
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 4.8574 | 0.48407 | 4.8565 | 0.48406
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 57.704 | 0.11064 | 57.704 | 0.11059
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 0.01 | 0.04501 | Hz | 579.97 | 0.011401 | 579.97 | 0.011382
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.7 | 0.35 | 0.69218 | 0.34795
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.78957 | 0.34543 | 0.78891 | 0.34543
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.2633 | 0.24921 | 5.2624 | 0.24918
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.892 | 0.04162 | 57.892 | 0.041575
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 0.045 | 10 | Hz | 579.99 | 0.0042219 | 579.99 | 0.0042034
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 10 | 20.01 | kHz | 0.7 | 0.81 | 0.69304 | 0.80896
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 10 | 20.01 | kHz | 0.74744 | 0.8075 | 0.74699 | 0.80751
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 10 | 20.01 | kHz | 4.3392 | 0.7169 | 4.3383 | 0.7169
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 10 | 20.01 | kHz | 57.426 | 0.20355 | 57.426 | 0.20351
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 10 | 20.01 | kHz | 579.94 | 0.021452 | 579.94 | 0.021433
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 20 | 50.01 | kHz | 0.7 | 1.2 | 0.69122 | 1.1565
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 20 | 50.01 | kHz | 0.73466 | 1.1553 | 0.73165 | 1.1554
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 20 | 50.01 | kHz | 3.8002 | 1.0754 | 3.7971 | 1.0754
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 20 | 50.01 | kHz | 56.883 | 0.3919 | 56.882 | 0.39182
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 20 | 50.01 | kHz | 579.88 | 0.043223 | 579.88 | 0.043182
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 50 | 100.01 | kHz | 2.3 | 2.7 | 2.3079 | 2.6577
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 50 | 100.01 | kHz | 2.3265 | 2.6572 | 2.3234 | 2.6573
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 50 | 100.01 | kHz | 3.8797 | 2.6158 | 3.8711 | 2.616
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 50 | 100.01 | kHz | 53.719 | 1.6172 | 53.713 | 1.6172
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 50 | 100.01 | kHz | 579.35 | 0.23028 | 579.35 | 0.23021
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3364 | 2.6995
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3501 | 2.699
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 3.831 | 2.6581
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 100 | 500 | kHz | 47.329 | 3.5597 | 53.608 | 1.6549
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 100 | 500 | kHz | 579.33 | 0.23723 | 579.33 | 0.23717
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 2.3364 | 2.6995
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 2.3501 | 2.699
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 3.831 | 2.6581
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 500 | 501 | kHz | 47.329 | 3.5597 | 53.608 | 1.6549
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 500 | 501 | kHz | 579.33 | 0.23723 | 579.33 | 0.23717
Transmille | DCA 50 turn | 5 | 51 | 0.001 | A | mA |  |  |  | 4.7 | 2 | 4.5993 | 1.98
Transmille | DCA 50 turn | 5 | 51 | 0.01 | A | mA |  |  |  | 5.9279 | 1.96 | 5.8001 | 1.96
Transmille | DCA 50 turn | 5 | 51 | 0.1 | A | mA |  |  |  | 53.039 | 1.32 | 52.916 | 1.32
Transmille | DCA 50 turn | 5 | 51 | 1 | A | mA |  |  |  | 576.56 | 0.206 | 576.52 | 0.204
Transmille | DCA 50 turn | 5 | 51 | 10 | A | mA |  |  |  | 5773.5 | 0.0215 | 5773.5 | 0.0206
3458A | ACA | 10 | 20 | 1e-05 | A | mA | 0 | 0 | kHz | 0.92 | 0.49 | 4.6374 | 0.081764
3458A | ACA | 10 | 20 | 0.0001 | A | mA | 0 | 0 | kHz | 0.92 | 0.49 | 4.6379 | 0.081759
3458A | ACA | 10 | 20 | 0.001 | A | mA | 0 | 0 | kHz | 0.92 | 0.49 | 4.6724 | 0.081363
3458A | ACA | 10 | 20 | 0.01 | A | mA | 0 | 0 | kHz | 5.1799 | 0.27701 | 7.3622 | 0.058218
3458A | ACA | 10 | 20 | 0.1 | A | mA | 0 | 0 | kHz | 57.914 | 0.0082611 | 57.911 | 0.0082616
3458A | ACA | 10 | 20 | 1e-05 | A | mA | 0 | 1 | kHz | 0 | 0.5 | 2.1832 | 0.11356
3458A | ACA | 10 | 20 | 0.0001 | A | mA | 0 | 1 | kHz | 0 | 0.5 | 2.1841 | 0.11354
3458A | ACA | 10 | 20 | 0.001 | A | mA | 0 | 1 | kHz | 0 | 0.5 | 2.2467 | 0.11228
3458A | ACA | 10 | 20 | 0.01 | A | mA | 0 | 1 | kHz | 3.3354 | 0.33323 | 6.0285 | 0.063256
3458A | ACA | 10 | 20 | 0.1 | A | mA | 0 | 1 | kHz | 57.758 | 0.0076264 | 57.755 | 0.0076268
3458A | ACA | 10 | 20 | 1e-05 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.0702 | 0.65987
3458A | ACA | 10 | 20 | 0.0001 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.0706 | 0.65986
3458A | ACA | 10 | 20 | 0.001 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.1025 | 0.65885
3458A | ACA | 10 | 20 | 0.01 | A | mA | 1 | 5 | kHz | 0.20999 | 0.9395 | 3.8087 | 0.57918
3458A | ACA | 10 | 20 | 0.1 | A | mA | 1 | 5 | kHz | 57.017 | 0.12296 | 57.014 | 0.12297
3458A | ACV | 10 | 100 | 1e-05 | V | mV | 100 | 300 | kHz | 15 | 4.6 | 11.575 | 4.6268
3458A | ACV | 10 | 100 | 0.0001 | V | mV | 100 | 300 | kHz | 15 | 4.6028 | 11.575 | 4.6268
3458A | ACV | 10 | 100 | 0.001 | V | mV | 100 | 300 | kHz | 15 | 4.6028 | 11.576 | 4.6268
3458A | ACV | 10 | 100 | 0.01 | V | mV | 100 | 300 | kHz | 15 | 4.6033 | 11.89 | 4.624
3458A | ACV | 10 | 100 | 0.1 | V | mV | 100 | 300 | kHz | 37.775 | 4.4 | 37.726 | 4.4003
3458A | ACV | 10 | 100 | 1 | V | mV | 100 | 300 | kHz | 561.72 | 1.8549 | 536.24 | 1.8547
3458A | ACV | 10 | 100 | 1e-05 | V | mV | 300 | 1000 | kHz | 46 | 17 | 12.317 | 17.334
3458A | ACV | 10 | 100 | 0.0001 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.318 | 17.334
3458A | ACV | 10 | 100 | 0.001 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.318 | 17.334
3458A | ACV | 10 | 100 | 0.01 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.427 | 17.333
3458A | ACV | 10 | 100 | 0.1 | V | mV | 300 | 1000 | kHz | 46 | 17.03 | 21.972 | 17.247
3458A | ACV | 10 | 100 | 1 | V | mV | 300 | 1000 | kHz | 469.61 | 13.691 | 433.99 | 13.692
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 0 | 0 | kHz | 4.7 | 0.23 | 4.6186 | 0.23291
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 0 | 0 | kHz | 4.7 | 0.23 | 4.619 | 0.2329
3458A | ACV | 10 | 20 | 0.001 | V | mV | 0 | 0 | kHz | 4.7 | 0.23 | 4.6486 | 0.23231
3458A | ACV | 10 | 20 | 0.01 | V | mV | 0 | 0 | kHz | 7.1399 | 0.18969 | 7.1403 | 0.18932
3458A | ACV | 10 | 20 | 0.1 | V | mV | 0 | 0 | kHz | 57.827 | 0.032722 | 57.828 | 0.0324
3458A | ACV | 10 | 20 | 1 | V | mV | 0 | 0 | kHz | 577.36 | 0.003403 | 577.07 | 0.0032722
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 0 | 1 | kHz | 2.4 | 0.23 | 2.298 | 0.23306
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 0 | 1 | kHz | 2.4 | 0.23 | 2.2985 | 0.23305
3458A | ACV | 10 | 20 | 0.001 | V | mV | 0 | 1 | kHz | 2.4 | 0.23 | 2.3458 | 0.23187
3458A | ACV | 10 | 20 | 0.01 | V | mV | 0 | 1 | kHz | 5.7569 | 0.1647 | 5.7574 | 0.16425
3458A | ACV | 10 | 20 | 0.1 | V | mV | 0 | 1 | kHz | 57.687 | 0.023591 | 57.688 | 0.023267
3458A | ACV | 10 | 20 | 1 | V | mV | 0 | 1 | kHz | 577.35 | 0.0024695 | 577.14 | 0.0023387
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.298 | 0.23306
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.2985 | 0.23305
3458A | ACV | 10 | 20 | 0.001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.3458 | 0.23187
3458A | ACV | 10 | 20 | 0.01 | V | mV | 1 | 20 | kHz | 5.7569 | 0.1647 | 5.7574 | 0.16425
3458A | ACV | 10 | 20 | 0.1 | V | mV | 1 | 20 | kHz | 57.687 | 0.023591 | 57.688 | 0.023267
3458A | ACV | 10 | 20 | 1 | V | mV | 1 | 20 | kHz | 577.35 | 0.0024695 | 577.14 | 0.0023387
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.2656 | 0.40924
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.266 | 0.40923
3458A | ACV | 10 | 20 | 0.001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.302 | 0.40822
3458A | ACV | 10 | 20 | 0.01 | V | mV | 20 | 50 | kHz | 5.2425 | 0.33547 | 5.2373 | 0.3351
3458A | ACV | 10 | 20 | 0.1 | V | mV | 20 | 50 | kHz | 57.496 | 0.059412 | 57.495 | 0.058915
3458A | ACV | 10 | 20 | 1 | V | mV | 20 | 50 | kHz | 577.33 | 0.0061608 | 576.8 | 0.0059565
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3767 | 1.3889
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3769 | 1.3889
3458A | ACV | 10 | 20 | 0.001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3922 | 1.3884
3458A | ACV | 10 | 20 | 0.01 | V | mV | 50 | 100 | kHz | 3.8217 | 1.3443 | 3.8184 | 1.3442
3458A | ACV | 10 | 20 | 0.1 | V | mV | 50 | 100 | kHz | 54.831 | 0.51571 | 54.83 | 0.51529
3458A | ACV | 10 | 20 | 1 | V | mV | 50 | 100 | kHz | 577.03 | 0.05599 | 572.43 | 0.055786
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 100 | 300 | kHz | 13 | 4.6 | 11.573 | 4.6268
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 100 | 300 | kHz | 13 | 4.6109 | 11.574 | 4.6268
3458A | ACV | 10 | 20 | 0.001 | V | mV | 100 | 300 | kHz | 13 | 4.6111 | 11.577 | 4.6267
3458A | ACV | 10 | 20 | 0.01 | V | mV | 100 | 300 | kHz | 13 | 4.6141 | 11.988 | 4.6141
3458A | ACV | 10 | 20 | 0.1 | V | mV | 100 | 300 | kHz | 44.406 | 3.7323 | 44.403 | 3.7322
3458A | ACV | 10 | 20 | 1 | V | mV | 100 | 300 | kHz | 573.82 | 0.64234 | 542.92 | 0.64214
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 300 | 1000 | kHz | 19 | 17 | 12.388 | 17.324
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 300 | 1000 | kHz | 19 | 17.087 | 12.389 | 17.324
3458A | ACV | 10 | 20 | 0.001 | V | mV | 300 | 1000 | kHz | 19 | 17.087 | 12.389 | 17.324
3458A | ACV | 10 | 20 | 0.01 | V | mV | 300 | 1000 | kHz | 19 | 17.092 | 12.517 | 17.32
3458A | ACV | 10 | 20 | 0.1 | V | mV | 300 | 1000 | kHz | 25.325 | 16.908 | 25.323 | 16.908
3458A | ACV | 10 | 20 | 1 | V | mV | 300 | 1000 | kHz | 533.13 | 7.3338 | 437.38 | 7.3336
3458A | DCV | 10 | 100 | 1e-06 | V | µV |  |  |  | 34 | 7.8 | 33.132 | 7.7866
3458A | DCV | 10 | 100 | 1e-05 | V | µV |  |  |  | 34 | 7.7848 | 33.3 | 7.7851
3458A | DCV | 10 | 100 | 0.0001 | V | µV |  |  |  | 48.99 | 7.6505 | 48.723 | 7.6514
3458A | DCV | 10 | 100 | 0.001 | V | µV |  |  |  | 545.46 | 4.5283 | 545.28 | 4.5242
3458A | DCV | 10 | 100 | 0.01 | V | µV |  |  |  | 5795 | 0.61987 | 5794.9 | 0.61637
3458A | Resistance | 10 | 100 | 1e-05 | Ω | µΩ |  |  |  | 580 | 14 | 577.27 | 0
3458A | Resistance | 10 | 100 | 0.0001 | Ω | µΩ |  |  |  | 580 | 13.975 | 579.78 | 13.975
3458A | Resistance | 10 | 100 | 0.001 | Ω | µΩ |  |  |  | 794.35 | 12.648 | 794.24 | 12.648
3458A | Resistance | 10 | 100 | 0.01 | Ω | µΩ |  |  |  | 5786.4 | 3.1626 | 5786.3 | 3.1617
3458A | Resistance | 10 | 100 | 0.1 | Ω | µΩ |  |  |  | 57737 | 0.32669 | 57737 | 0.32628
3458A | Resistance | 10 | 100 | 1 | Ω | µΩ |  |  |  | 577360 | 0.032802 | 577350 | 0.032639
34420A | DCVmeas | 10 | 100 | 0.0001 | V | mV |  |  |  | 0.58 | 0.041 | 0.57895 | 0.040741
34420A | DCVmeas | 10 | 100 | 0.001 | V | mV |  |  |  | 0.7495 | 0.039399 | 0.74879 | 0.039399
34420A | DCVmeas | 10 | 100 | 0.01 | V | mV |  |  |  | 5.7109 | 0.017264 | 5.7105 | 0.017251
34420A | DCVmeas | 10 | 100 | 0.1 | V | mV |  |  |  | 57.989 | 0.0019846 | 57.989 | 0.0019773
34420A | DCVmeas | 10 | 100 | 1 | V | mV |  |  |  | 580 | 0.000201 | 580 | 0.00019806
Transmille | ACV | 10 | 100 | 0.001 | V | mV | 0.04 | 1 | KHz | 35 | 0.95 | 34.943 | 0.948
Transmille | ACV | 10 | 100 | 0.01 | V | mV | 0.04 | 1 | KHz | 35.349 | 0.946 | 35.282 | 0.945
Transmille | ACV | 10 | 100 | 0.1 | V | mV | 0.04 | 1 | KHz | 65.076 | 0.769 | 64.997 | 0.768
Transmille | ACV | 10 | 100 | 1 | V | mV | 0.04 | 1 | KHz | 577.66 | 0.143 | 577.63 | 0.141
Transmille | ACV | 10 | 100 | 10 | V | mV | 0.04 | 1 | KHz | 5773.6 | 0.0151 | 5773.6 | 0.0143
Transmille | DCA 10 turn | 10 | 102 | 0.0001 | A | mA |  |  |  | 16 | 2 | 15.052 | 2.01
Transmille | DCA 10 turn | 10 | 102 | 0.001 | A | mA |  |  |  | 16 | 2 | 15.566 | 2
Transmille | DCA 10 turn | 10 | 102 | 0.01 | A | mA |  |  |  | 50.243 | 1.74 | 50.216 | 1.74
Transmille | DCA 10 turn | 10 | 102 | 0.1 | A | mA |  |  |  | 574.15 | 0.429 | 574.14 | 0.428
Transmille | DCA 10 turn | 10 | 102 | 1 | A | mA |  |  |  | 5773.2 | 0.0445 | 5773.2 | 0.0444
Transmille | DCA 10 turn | 10 | 102 | 10 | A | mA |  |  |  | 57735 | 0.00449 | 57735 | 0.00444
Transmille | DCV | 10 | 100 | 0.001 | V | mV |  |  |  | 6.8 | 0.082 | 6.7338 | 0.0825
Transmille | DCV | 10 | 100 | 0.01 | V | mV |  |  |  | 8.774 | 0.0728 | 8.766 | 0.0728
Transmille | DCV | 10 | 100 | 0.1 | V | mV |  |  |  | 58.072 | 0.0158 | 58.068 | 0.0158
Transmille | DCV | 10 | 100 | 1 | V | mV |  |  |  | 577.39 | 0.00164 | 577.39 | 0.00161
Transmille | DCV | 10 | 100 | 10 | V | mV |  |  |  | 5773.6 | 0.000174 | 5773.6 | 0.000161
5502A | ACW | 10.9 | 109.0001 | 0.001 | WAC | WAC |  |  |  | 0 | 0.00088 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0010256 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0053997 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.053268 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.0053323 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.00057887 | 0 | 0
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 0.045 | 0.10001 | kHz | 5.8 | 1.4 | 5.7745 | 1.386
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 0.045 | 0.10001 | kHz | 5.8 | 1.4 | 5.7855 | 1.3857
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 6.9249 | 1.3567 | 6.9028 | 1.3548
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 55.153 | 0.59996 | 55.168 | 0.5931
5502A | ACA | 11 | 20.501 | 1 | A | mA | 0.045 | 0.10001 | kHz | 579.65 | 0.069317 | 579.66 | 0.065887
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7754 | 1.7324
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7852 | 1.7321
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 0.1 | 1 | kHz | 6.7671 | 1.7062 | 6.7466 | 1.7047
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 0.1 | 1 | kHz | 53.699 | 0.85925 | 53.709 | 0.85296
5502A | ACA | 11 | 20.501 | 1 | A | mA | 0.1 | 1 | kHz | 579.44 | 0.102 | 579.45 | 0.098572
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7742 | 34.641
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7745 | 34.641
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.8395 | 34.639
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 1 | 5 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
5502A | ACA | 11 | 20.501 | 1 | A | mA | 1 | 5 | kHz | 437.55 | 23.603 | 437.54 | 23.602
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.7742 | 34.641
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.7745 | 34.641
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.8395 | 34.639
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 5 | 5.01 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
5502A | ACA | 11 | 20.501 | 1 | A | mA | 5 | 5.01 | kHz | 437.55 | 23.603 | 437.54 | 23.602
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.87 | 2.9 | 0.85389 | 2.8839
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 1.4281 | 2.8814 | 1.5429 | 2.8673
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 44.394 | 2.0188 | 44.317 | 2.0184
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 577.68 | 0.29774 | 577.66 | 0.2963
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 5799.8 | 0.030408 | 5799.8 | 0.029819
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.9 | 2.1 | 0.42297 | 2.1155
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 1.3908 | 2.0927 | 1.3648 | 2.0928
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 48.986 | 1.2457 | 48.967 | 1.244
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 0.045 | 1 | kHz | 578.74 | 0.1605 | 578.73 | 0.15918
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 0.045 | 1 | kHz | 5799.9 | 0.016507 | 5799.9 | 0.015972
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.174 | 7.2092
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.422 | 7.2034
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 33.827 | 6.6978 | 33.792 | 6.6979
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 1 | 5.01 | kHz | 566.19 | 1.8955 | 566.18 | 1.8942
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 1 | 5.01 | kHz | 5798.5 | 0.19808 | 5798.5 | 0.1975
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.627 | 28.936
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.682 | 28.935
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 5 | 10 | kHz | 63.353 | 28.822 | 63.613 | 28.797
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 5 | 10 | kHz | 458.92 | 21.118 | 458.89 | 21.117
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 5 | 10 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.627 | 28.936
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.682 | 28.935
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 63.353 | 28.822 | 63.613 | 28.797
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 10 | 10.01 | kHz | 458.92 | 21.118 | 458.89 | 21.117
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 10 | 10.01 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
5502A | DCA | 11 | 20.5 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
5502A | DCA | 11 | 20.5 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
5502A | DCA | 11 | 20.5 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
5502A | DCA | 11 | 20.5 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
5502A | DCA 10 turn | 11 | 30.001 | 0.001 | A | mA |  |  |  | 1.6 | 2 | 0.12764 | 2.0471
5502A | DCA 10 turn | 11 | 30.001 | 0.01 | A | mA |  |  |  | 1.6 | 2.0127 | 1.1038 | 2.0235
5502A | DCA 10 turn | 11 | 30.001 | 0.1 | A | mA |  |  |  | 49.094 | 1.1796 | 49.066 | 1.1772
5502A | DCA 10 turn | 11 | 30.001 | 1 | A | mA |  |  |  | 576.17 | 0.15058 | 576.16 | 0.14877
5502A | DCA 10 turn | 11 | 30.001 | 10 | A | mA |  |  |  | 5773.4 | 0.015656 | 5773.4 | 0.014925
5502A | DCA 10 turn | 11 | 30.001 | 100 | A | mA |  |  |  | 57736 | 0.0017851 | 57735 | 0.0014925
5502A | Resistance | 11 | 33.001 | 0.001 | Ω | mΩ |  |  |  | 18 | 0.14 | 17.339 | 0.13885
5502A | Resistance | 11 | 33.001 | 0.01 | Ω | mΩ |  |  |  | 18.263 | 0.13372 | 18.258 | 0.13357
5502A | Resistance | 11 | 33.001 | 0.1 | Ω | mΩ |  |  |  | 60.487 | 0.046304 | 60.481 | 0.04604
5502A | Resistance | 11 | 33.001 | 1 | Ω | mΩ |  |  |  | 580.26 | 0.0049938 | 580.25 | 0.0048783
5502A | DCA 20 turn | 20 | 60.01 | 0.001 | A | mA |  |  |  | 3.9 | 2.1 | 0.23928 | 2.0473
5502A | DCA 20 turn | 20 | 60.01 | 0.01 | A | mA |  |  |  | 3.9 | 2.1 | 0.76998 | 2.0407
5502A | DCA 20 turn | 20 | 60.01 | 0.1 | A | mA |  |  |  | 38.447 | 1.626 | 38.404 | 1.6257
5502A | DCA 20 turn | 20 | 60.01 | 1 | A | mA |  |  |  | 573.08 | 0.28848 | 573.07 | 0.28764
5502A | DCA 20 turn | 20 | 60.01 | 10 | A | mA |  |  |  | 5773.1 | 0.029468 | 5773.1 | 0.02912
5502A | DCA 20 turn | 20 | 60.01 | 100 | A | mA |  |  |  | 57735 | 0.0030514 | 57735 | 0.0029124
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 0 | 0 | kHz | 5 | 0.23 | 4.6009 | 0.23351
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 0 | 0 | kHz | 5 | 0.23167 | 4.6011 | 0.23351
3458A | ACV | 20 | 100 | 0.001 | V | mV | 0 | 0 | kHz | 5 | 0.23175 | 4.6195 | 0.23336
3458A | ACV | 20 | 100 | 0.01 | V | mV | 0 | 0 | kHz | 6.5042 | 0.22022 | 6.5015 | 0.22022
3458A | ACV | 20 | 100 | 0.1 | V | mV | 0 | 0 | kHz | 57.057 | 0.070854 | 57.056 | 0.070819
3458A | ACV | 20 | 100 | 1 | V | mV | 0 | 0 | kHz | 577.28 | 0.0075312 | 576.01 | 0.0075149
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 0 | 1 | kHz | 2.7 | 0.23 | 2.2758 | 0.23386
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 0 | 1 | kHz | 2.7 | 0.2318 | 2.276 | 0.23386
3458A | ACV | 20 | 100 | 0.001 | V | mV | 0 | 1 | kHz | 2.7 | 0.23188 | 2.2998 | 0.23364
3458A | ACV | 20 | 100 | 0.01 | V | mV | 0 | 1 | kHz | 4.7115 | 0.21577 | 4.7084 | 0.21577
3458A | ACV | 20 | 100 | 0.1 | V | mV | 0 | 1 | kHz | 56.896 | 0.062832 | 56.895 | 0.062796
3458A | ACV | 20 | 100 | 1 | V | mV | 0 | 1 | kHz | 577.27 | 0.0066108 | 576.14 | 0.0065945
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8564 | 0.20787
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8565 | 0.20787
3458A | ACV | 20 | 100 | 0.001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.882 | 0.20768
3458A | ACV | 20 | 100 | 0.01 | V | mV | 1 | 20 | kHz | 7.1093 | 0.19179 | 7.1056 | 0.1918
3458A | ACV | 20 | 100 | 0.1 | V | mV | 1 | 20 | kHz | 57.594 | 0.05585 | 57.592 | 0.055819
3458A | ACV | 20 | 100 | 1 | V | mV | 1 | 20 | kHz | 577.34 | 0.0058763 | 576.84 | 0.0058618
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 20 | 50 | kHz | 6.8 | 0.37 | 6.7323 | 0.36645
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36642 | 6.7324 | 0.36645
3458A | ACV | 20 | 100 | 0.001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36631 | 6.7495 | 0.36632
3458A | ACV | 20 | 100 | 0.01 | V | mV | 20 | 50 | kHz | 8.3578 | 0.35409 | 8.3518 | 0.35408
3458A | ACV | 20 | 100 | 0.1 | V | mV | 20 | 50 | kHz | 57.163 | 0.15069 | 57.159 | 0.15056
3458A | ACV | 20 | 100 | 1 | V | mV | 20 | 50 | kHz | 577.28 | 0.017106 | 575.94 | 0.01704
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 50 | 100 | kHz | 22 | 1.2 | 17.722 | 1.2358
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 50 | 100 | kHz | 22 | 1.2034 | 17.723 | 1.2358
3458A | ACV | 20 | 100 | 0.001 | V | mV | 50 | 100 | kHz | 22 | 1.2034 | 17.724 | 1.2358
3458A | ACV | 20 | 100 | 0.01 | V | mV | 50 | 100 | kHz | 22 | 1.2047 | 18.315 | 1.2311
3458A | ACV | 20 | 100 | 0.1 | V | mV | 50 | 100 | kHz | 55.391 | 0.9727 | 55.373 | 0.9727
3458A | ACV | 20 | 100 | 1 | V | mV | 50 | 100 | kHz | 576.34 | 0.18075 | 568.41 | 0.18063
Transmille | DCA 20 turn | 20 | 204 | 0.001 | A | mA |  |  |  | 32 | 2 | 30.087 | 2.01
Transmille | DCA 20 turn | 20 | 204 | 0.01 | A | mA |  |  |  | 32 | 2.01 | 30.344 | 2.01
Transmille | DCA 20 turn | 20 | 204 | 0.1 | A | mA |  |  |  | 52.695 | 1.92 | 52.601 | 1.92
Transmille | DCA 20 turn | 20 | 204 | 1 | A | mA |  |  |  | 565.98 | 0.784 | 565.94 | 0.784
Transmille | DCA 20 turn | 20 | 204 | 10 | A | mA |  |  |  | 5772.2 | 0.0889 | 5772.2 | 0.0886
5502A | DCA | 20.5 | 20.6 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
5502A | DCA | 20.5 | 20.6 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
5502A | DCA | 20.5 | 20.6 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
5502A | DCA | 20.5 | 20.6 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | -5.2 | 3 | -6.3515 | 3.0358
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | -5.2 | 3.0309 | -6.0166 | 3.031
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 24.504 | 2.6609 | 24.465 | 2.6608
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 569.88 | 0.59872 | 569.88 | 0.59787
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 5799 | 0.061619 | 5799 | 0.061255
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -5.078 | 2.227
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -4.615 | 2.2203
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 25.632 | 2.1228 | 32.901 | 1.7903
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 0.045 | 1 | kHz | 574.48 | 0.32468 | 574.48 | 0.3238
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 0.045 | 1 | kHz | 5799.4 | 0.033173 | 5799.4 | 0.032809
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 2.3 | 7.6 | 2.2215 | 7.5886
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 2.3482 | 7.5869 | 2.3433 | 7.5869
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 14.11 | 7.4273 | 14.091 | 7.4273
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 1 | 5.01 | kHz | 526.05 | 3.5184 | 526.05 | 3.5178
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 1 | 5.01 | kHz | 5793.5 | 0.40619 | 5793.5 | 0.40583
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 5 | 10 | kHz | 52 | 30 | 34.529 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 5 | 10 | kHz | 52 | 30.458 | 34.545 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 5 | 10 | kHz | 52 | 30.419 | 37.46 | 30.419
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 5 | 10 | kHz | 304.91 | 27.255 | 304.84 | 27.255
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 5 | 10 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 52 | 30 | 34.529 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 52 | 30.458 | 34.545 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 52 | 30.419 | 37.46 | 30.419
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 10 | 10.01 | kHz | 304.91 | 27.255 | 304.84 | 27.255
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 10 | 10.01 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 12 | 2.1 | 11.984 | 2.086
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 36.579 | 1.9251 | 36.551 | 1.9245
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 568.43 | 0.54491 | 568.44 | 0.54347
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5798.8 | 0.057436 | 5798.8 | 0.056801
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 58000 | 0.0059371 | 58000 | 0.0056828
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 15 | 2.3 | 14.872 | 2.2866
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 37.348 | 2.1392 | 37.319 | 2.1387
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 0.1 | 1 | kHz | 566.33 | 0.65492 | 566.34 | 0.65351
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 0.1 | 1 | kHz | 5798.5 | 0.069561 | 5798.5 | 0.068926
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0071509 | 58000 | 0.0068966
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 1 | 5 | kHz | 24 | 35 | 23.064 | 34.699
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 1 | 5 | kHz | 24 | 35 | 25.112 | 34.684
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 1 | 5 | kHz | 210.04 | 33.396 | 210.02 | 33.395
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 1 | 5 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 1 | 5 | kHz | 57966 | 1.4661 | 57966 | 1.4652
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 24 | 35 | 23.064 | 34.699
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 24 | 35 | 25.112 | 34.684
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 5 | 5.01 | kHz | 210.04 | 33.396 | 210.02 | 33.395
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 5 | 5.01 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 5 | 5.01 | kHz | 57966 | 1.4661 | 57966 | 1.4652
5502A | DCA 10 turn | 30 | 110.01 | 0.01 | A | mA |  |  |  | 4 | 2.1 | 2.509 | 2.1122
5502A | DCA 10 turn | 30 | 110.01 | 0.1 | A | mA |  |  |  | 29.551 | 1.93 | 29.511 | 1.9297
5502A | DCA 10 turn | 30 | 110.01 | 1 | A | mA |  |  |  | 565.25 | 0.52854 | 565.25 | 0.52743
5502A | DCA 10 turn | 30 | 110.01 | 10 | A | mA |  |  |  | 5772.3 | 0.055475 | 5772.3 | 0.054985
5502A | DCA 10 turn | 30 | 110.01 | 100 | A | mA |  |  |  | 57735 | 0.0056968 | 57735 | 0.005501
5502A | FDPT100 | 32 | 212.1 | 0.1 | °F | °F |  |  |  | 0.28537 | 0 | 0.26385 | 0
5502A | FDPT100 | 32 | 212.1 | 1 | °F | °F |  |  |  | 0.61101 | 0 | 0.59895 | 0
5502A | FDPT1000 | 32 | 212.1 | 0.1 | °F | °F |  |  |  | 0.19201 | 0 | 0.15827 | 0
5502A | FDR | 32 | 482.1 | 0.1 | °F | °F |  |  |  | 1.1966 | 0 | 1.19 | 0
5502A | FDR | 32 | 482.1 | 1 | °F | °F |  |  |  | 1.3448 | 0 | 1.3214 | 0
5502A | FDS | 32 | 482.1 | 0.1 | °F | °F |  |  |  | 0.99102 | 0 | 0.98329 | 0
5502A | FDS | 32 | 482.1 | 1 | °F | °F |  |  |  | 1.1653 | 0 | 1.1388 | 0
5502A | FDT | 32 | 248.1 | 0.1 | °F | °F |  |  |  | 0.20015 | 0 | 0.19651 | 0
5502A | FDT | 32 | 248.1 | 1 | °F | °F |  |  |  | 0.61188 | 0 | 0.60714 | 0
3458A | FDR | 32 | 1100 | 1 | °F | °F |  |  |  | 0.64166 | 0 | 0.59103 | 0
3458A | FDR | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13898 | 0
3458A | FDS | 32 | 1100 | 1 | °F | °F |  |  |  | 0.63802 | 0 | 0.58826 | 0
3458A | FDS | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12667 | 0
Transmille | TcindicatorR | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
Transmille | TcindicatorR | 32 | 482 | 1 | °F | °F |  |  |  | 4.8335 | 0 | 4.8271 | 0
Transmille | TcindicatorS | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
Transmille | TcindicatorS | 32 | 482 | 1 | °F | °F |  |  |  | 4.8334 | 0 | 4.8271 | 0
3458A | FDPT100 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014506 | 0
3458A | FDPT100 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57812 | 0 | 0.5775 | 0
3458A | FDPT100 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.060734 | 0 | 0.059249 | 0
3458A | FDPT1000 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014466 | 0
3458A | FDPT1000 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57781 | 0 | 0.5775 | 0
3458A | FDPT1000 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.059981 | 0 | 0.059239 | 0
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 3.5 | 0.58 | 3.4705 | 0.57868
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 4.2872 | 0.57647 | 4.2606 | 0.57652
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 46.641 | 0.47359 | 46.602 | 0.47361
5502A | ACV | 33 | 330.01 | 1 | V | mV | 0.01 | 0.04501 | Hz | 576.98 | 0.10542 | 576.96 | 0.10533
5502A | ACV | 33 | 330.01 | 10 | V | mV | 0.01 | 0.04501 | Hz | 5799.7 | 0.010862 | 5799.7 | 0.010822
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 0.045 | 10 | Hz | 12 | 0.92 | 10.393 | 0.92459
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 0.045 | 10 | Hz | 12 | 0.92331 | 10.834 | 0.9234
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 0.045 | 10 | Hz | 43.309 | 0.841 | 43.21 | 0.84115
5502A | ACV | 33 | 330.01 | 1 | V | mV | 0.045 | 10 | Hz | 572.72 | 0.26556 | 572.68 | 0.26538
5502A | ACV | 33 | 330.01 | 10 | V | mV | 0.045 | 10 | Hz | 5799.2 | 0.028474 | 5799.2 | 0.028386
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 10 | 20.01 | kHz | 24 | 1 | 10.388 | 1.04
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 10 | 20.01 | kHz | 24 | 1.0332 | 10.796 | 1.0389
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 10 | 20.01 | kHz | 41.651 | 0.95969 | 41.554 | 0.95985
5502A | ACV | 33 | 330.01 | 1 | V | mV | 10 | 20.01 | kHz | 570.92 | 0.32864 | 570.88 | 0.32848
5502A | ACV | 33 | 330.01 | 10 | V | mV | 10 | 20.01 | kHz | 5799 | 0.035762 | 5799 | 0.035674
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 20 | 50.01 | kHz | 22 | 1.4 | 10.266 | 1.3912
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 20 | 50.01 | kHz | 22 | 1.4 | 10.577 | 1.3904
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 20 | 50.01 | kHz | 36.02 | 1.3575 | 37.167 | 1.3205
5502A | ACV | 33 | 330.01 | 1 | V | mV | 20 | 50.01 | kHz | 564.6 | 0.55034 | 564.56 | 0.55021
5502A | ACV | 33 | 330.01 | 10 | V | mV | 20 | 50.01 | kHz | 5798.2 | 0.063012 | 5798.2 | 0.062925
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 50 | 100 | kHz | 92 | 2.8 | 91.937 | 2.7856
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 50 | 100 | kHz | 92.063 | 2.7853 | 92.033 | 2.7854
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 50 | 100 | kHz | 101.77 | 2.761 | 101.69 | 2.7611
5502A | ACV | 33 | 330.01 | 1 | V | mV | 50 | 100 | kHz | 546.63 | 1.8765 | 546.54 | 1.8763
5502A | ACV | 33 | 330.01 | 10 | V | mV | 50 | 100 | kHz | 5793.6 | 0.28499 | 5793.5 | 0.28476
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 100 | 100.1 | kHz | 92 | 2.8 | 91.937 | 2.7856
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 100 | 100.1 | kHz | 92.063 | 2.7853 | 92.033 | 2.7854
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 100 | 100.1 | kHz | 101.77 | 2.761 | 101.69 | 2.7611
5502A | ACV | 33 | 330.01 | 1 | V | mV | 100 | 100.1 | kHz | 546.63 | 1.8765 | 546.54 | 1.8763
5502A | ACV | 33 | 330.01 | 10 | V | mV | 100 | 100.1 | kHz | 5793.6 | 0.28499 | 5793.5 | 0.28476
5502A | DCV | 33 | 330.001 | 0.0001 | V | mV |  |  |  | 0.59 | 0.064 | 0.57834 | 0.063644
5502A | DCV | 33 | 330.001 | 0.001 | V | mV |  |  |  | 0.67471 | 0.063403 | 0.64578 | 0.063463
5502A | DCV | 33 | 330.001 | 0.01 | V | mV |  |  |  | 4.6629 | 0.053693 | 4.6153 | 0.053731
5502A | DCV | 33 | 330.001 | 0.1 | V | mV |  |  |  | 57.654 | 0.012974 | 57.637 | 0.012872
5502A | DCV | 33 | 330.001 | 1 | V | mV |  |  |  | 579.97 | 0.0013757 | 579.96 | 0.0013305
5502A | DCV | 33 | 330.001 | 10 | V | mV |  |  |  | 5800 | 0.00015118 | 5800 | 0.0001331
5502A | Resistance | 33 | 110.01 | 0.001 | Ω | mΩ |  |  |  | 18 | 0.1 | 17.342 | 0.10406
5502A | Resistance | 33 | 110.01 | 0.01 | Ω | mΩ |  |  |  | 18.224 | 0.10151 | 18.221 | 0.10128
5502A | Resistance | 33 | 110.01 | 0.1 | Ω | mΩ |  |  |  | 60.259 | 0.041284 | 60.259 | 0.04082
5502A | Resistance | 33 | 110.01 | 1 | Ω | mΩ |  |  |  | 580.23 | 0.0046533 | 580.23 | 0.0044422
3458A | Frequency | 40 | 1000 | 0.0001 | Hz | Hz |  |  |  | 2.9e-07 | 1.3e-07 | 2.53e-07 | 1.2897e-07
3458A | Frequency | 40 | 1000 | 0.001 | Hz | Hz |  |  |  | 6.57e-07 | 1.2894e-07 | 5.98e-07 | 1.2907e-07
3458A | Frequency | 40 | 1000 | 0.01 | Hz | Hz |  |  |  | 2.609e-05 | 1.3885e-07 | 2.5942e-05 | 1.3918e-07
3458A | Frequency | 40 | 1000 | 0.1 | Hz | Hz |  |  |  | 0.00039689 | 5.557e-07 | 0.00039681 | 5.5588e-07
3458A | Frequency | 40 | 1000 | 1 | Hz | Hz |  |  |  | 0.0039872 | 5.4141e-06 | 0.0039872 | 5.4141e-06
3458A | Frequency | 40 | 1000 | 10 | Hz | Hz |  |  |  | 0.055949 | 5.4127e-05 | 0.055949 | 5.4127e-05
3458A | Frequency | 40 | 10000000 | 0 | Hz | Hz |  |  |  | 0 | 1.2e-07 | 0 | 1.1561e-07
Transmille | DCA 50 turn | 50 | 510 | 0.001 | A | mA |  |  |  | 80 | 2 | 75.208 | 2.01
Transmille | DCA 50 turn | 50 | 510 | 0.01 | A | mA |  |  |  | 80 | 2.01 | 75.318 | 2.01
Transmille | DCA 50 turn | 50 | 510 | 0.1 | A | mA |  |  |  | 85.569 | 1.99 | 85.292 | 1.99
Transmille | DCA 50 turn | 50 | 510 | 1 | A | mA |  |  |  | 534.37 | 1.39 | 534.07 | 1.39
Transmille | DCA 50 turn | 50 | 510 | 10 | A | mA |  |  |  | 5765.3 | 0.22 | 5765.2 | 0.22
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8 | 4.2301 | 2.8841
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.7673 | 4.3733 | 2.8834
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8076 | 17.829 | 2.819
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.01 | 0.045 | kHz | 530.85 | 1.3025 | 530.83 | 1.3019
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.01 | 0.045 | kHz | 5794.1 | 0.14924 | 5794.1 | 0.14886
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.01506 | 57992 | 0.01491
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 20.257 | 2.0293
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.045 | 0.1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.045 | 0.1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 20.257 | 2.0293
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.1 | 1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.1 | 1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 1 | 5.01 | kHz | 58 | 7.2 | 55.853 | 7.2093
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 1 | 5.01 | kHz | 58 | 7.209 | 55.901 | 7.2091
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 1 | 5.01 | kHz | 60.896 | 7.1858 | 60.843 | 7.1859
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 1 | 5.01 | kHz | 422.41 | 5.6956 | 422.3 | 5.6954
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 1 | 5.01 | kHz | 5763.9 | 0.97779 | 5763.9 | 0.97715
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 1 | 5.01 | kHz | 57996 | 0.099047 | 57948 | 0.098783
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 5 | 10 | kHz | 290 | 29 | 288.13 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 5 | 10 | kHz | 290 | 28.936 | 288.14 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 5 | 10 | kHz | 290 | 28.931 | 289.4 | 28.93
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 5 | 10 | kHz | 405.34 | 28.398 | 405.23 | 28.397
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 5 | 10 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 5 | 10 | kHz | 57941 | 1.6237 | 57264 | 1.6205
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 10 | 10.01 | kHz | 290 | 29 | 288.13 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 10 | 10.01 | kHz | 290 | 28.936 | 288.14 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 10 | 10.01 | kHz | 290 | 28.931 | 289.4 | 28.93
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 10 | 10.01 | kHz | 405.34 | 28.398 | 405.23 | 28.397
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 10 | 10.01 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 10 | 10.01 | kHz | 57941 | 1.6237 | 57264 | 1.6205
5502A | DCA 50 turn | 55 | 150.01 | 0.001 | A | mA |  |  |  | 7.7 | 2 | 0.58265 | 2.0474
5502A | DCA 50 turn | 55 | 150.01 | 0.01 | A | mA |  |  |  | 7.7 | 2.0006 | 0.78384 | 2.0464
5502A | DCA 50 turn | 55 | 150.01 | 0.1 | A | mA |  |  |  | 19.451 | 1.9578 | 19.379 | 1.9579
5502A | DCA 50 turn | 55 | 150.01 | 1 | A | mA |  |  |  | 550.23 | 0.69408 | 550.2 | 0.69349
5502A | DCA 50 turn | 55 | 150.01 | 10 | A | mA |  |  |  | 5770.6 | 0.074862 | 5770.6 | 0.074565
5502A | DCA 50 turn | 55 | 150.01 | 100 | A | mA |  |  |  | 57735 | 0.0075815 | 57732 | 0.0074625
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 24 | 2.1 | 23.534 | 2.0875
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 37.146 | 2.0415 | 37.093 | 2.0415
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 540.32 | 0.97471 | 540.3 | 0.97417
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5795.1 | 0.11376 | 5795.1 | 0.11344
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 58000 | 0.011493 | 58000 | 0.011365
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 30 | 2.3 | 29.364 | 2.2879
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 41.544 | 2.2475 | 41.547 | 2.2467
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 0.1 | 1 | kHz | 534.45 | 1.1508 | 534.42 | 1.1503
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 0.1 | 1 | kHz | 5794.1 | 0.13793 | 5794.1 | 0.13761
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.01392 | 58000 | 0.013793
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 1 | 5 | kHz | 47 | 35 | 46.103 | 34.699
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 1 | 5 | kHz | 47 | 35 | 47.146 | 34.695
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 1 | 5 | kHz | 127.73 | 34.633 | 144.7 | 34.35
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 1 | 5 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 1 | 5 | kHz | 57864 | 2.9205 | 57864 | 2.9201
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 47 | 35 | 46.103 | 34.699
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 47 | 35 | 47.146 | 34.695
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 5 | 5.01 | kHz | 127.73 | 34.633 | 144.7 | 34.35
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 5 | 5.01 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 5 | 5.01 | kHz | 57864 | 2.9205 | 57864 | 2.9201
5502A | DCA 20 turn | 60 | 220.01 | 0.001 | A | mA |  |  |  | 19 | 2.2 | 4.5221 | 2.114
5502A | DCA 20 turn | 60 | 220.01 | 0.01 | A | mA |  |  |  | 19 | 2.2 | 19.726 | 2.0608
5502A | DCA 20 turn | 60 | 220.01 | 0.1 | A | mA |  |  |  | 535.04 | 0.95226 | 535.04 | 0.95071
5502A | DCA 20 turn | 60 | 220.01 | 1 | A | mA |  |  |  | 5768.4 | 0.11068 | 5768.5 | 0.10982
5502A | DCA 20 turn | 60 | 220.01 | 10 | A | mA |  |  |  | 57735 | 0.011347 | 57735 | 0.011002
5502A | DCA 20 turn | 60 | 220.01 | 100 | A | mA |  |  |  | 577360 | 0.0012382 | 577350 | 0.0011002
3458A | CDT | 93 | 400 | 1 | °C | °C |  |  |  | 0.58328 | 0 | 0.57831 | 0
3458A | CDT | 93 | 400 | 0.1 | °C | °C |  |  |  | 0.076722 | 0 | 0.066645 | 0
5502A | CDPT100 | 100 | 300.1 | 0.1 | °C | °C |  |  |  | 0.16571 | 0 | 0.15431 | 0
5502A | CDPT100 | 100 | 300.1 | 1 | °C | °C |  |  |  | 0.60946 | 0 | 0.59736 | 0
5502A | CDPT1000 | 100 | 260.1 | 0.1 | °C | °C |  |  |  | 0.11707 | 0 | 0.10029 | 0
5502A | CDPT1000 | 100 | 260.1 | 1 | °C | °C |  |  |  | 0.59806 | 0 | 0.58574 | 0
3458A | ACV | 100 | 1000 | 0.0001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.055 | 0.69387
3458A | ACV | 100 | 1000 | 0.001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.063 | 0.69386
3458A | ACV | 100 | 1000 | 0.01 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.256 | 0.69369
3458A | ACV | 100 | 1000 | 0.1 | V | mV | 1 | 20 | kHz | 40.883 | 0.68212 | 41.185 | 0.67806
3458A | ACV | 100 | 1000 | 1 | V | mV | 1 | 20 | kHz | 547.47 | 0.37317 | 547.4 | 0.3731
3458A | ACV | 100 | 1000 | 0.0001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 22.895 | 1.3896
3458A | ACV | 100 | 1000 | 0.001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 22.896 | 1.3896
3458A | ACV | 100 | 1000 | 0.01 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.007 | 1.3895
3458A | ACV | 100 | 1000 | 0.1 | V | mV | 20 | 50 | kHz | 33.01 | 1.39 | 33.875 | 1.3798
3458A | ACV | 100 | 1000 | 1 | V | mV | 20 | 50 | kHz | 496.92 | 1.0292 | 496.68 | 1.0293
3458A | ACV | 100 | 1000 | 0.0001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.116 | 3.4656
3458A | ACV | 100 | 1000 | 0.001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.117 | 3.4656
3458A | ACV | 100 | 1000 | 0.01 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.153 | 3.4656
3458A | ACV | 100 | 1000 | 0.1 | V | mV | 50 | 100 | kHz | 24.483 | 3.4985 | 28.03 | 3.4612
3458A | ACV | 100 | 1000 | 1 | V | mV | 50 | 100 | kHz | 369.25 | 3.1671 | 368.87 | 3.1673
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 0 | 0 | kHz | 48 | 0.46 | 46.118 | 0.46325
3458A | ACV | 100 | 330 | 0.001 | V | mV | 0 | 0 | kHz | 48 | 0.46086 | 46.12 | 0.46325
3458A | ACV | 100 | 330 | 0.01 | V | mV | 0 | 0 | kHz | 48 | 0.46129 | 46.341 | 0.46283
3458A | ACV | 100 | 330 | 0.1 | V | mV | 0 | 0 | kHz | 66.332 | 0.42712 | 66.294 | 0.42698
3458A | ACV | 100 | 330 | 1 | V | mV | 0 | 0 | kHz | 573.42 | 0.11327 | 573.42 | 0.11294
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 0 | 1 | kHz | 25 | 0.46 | 23.034 | 0.46337
3458A | ACV | 100 | 330 | 0.001 | V | mV | 0 | 1 | kHz | 25 | 0.46067 | 23.037 | 0.46337
3458A | ACV | 100 | 330 | 0.01 | V | mV | 0 | 1 | kHz | 25 | 0.46116 | 23.338 | 0.46274
3458A | ACV | 100 | 330 | 0.1 | V | mV | 0 | 1 | kHz | 49.032 | 0.41285 | 48.983 | 0.41271
3458A | ACV | 100 | 330 | 1 | V | mV | 0 | 1 | kHz | 571.92 | 0.096252 | 571.92 | 0.095921
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.059 | 0.69384
3458A | ACV | 100 | 330 | 0.001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.061 | 0.69384
3458A | ACV | 100 | 330 | 0.01 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.291 | 0.69334
3458A | ACV | 100 | 330 | 0.1 | V | mV | 1 | 20 | kHz | 44.083 | 0.65012 | 43.963 | 0.65028
3458A | ACV | 100 | 330 | 1 | V | mV | 1 | 20 | kHz | 565.1 | 0.19691 | 565.03 | 0.19677
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.053 | 1.3879
3458A | ACV | 100 | 330 | 0.001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.054 | 1.3879
3458A | ACV | 100 | 330 | 0.01 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.188 | 1.3876
3458A | ACV | 100 | 330 | 0.1 | V | mV | 20 | 50 | kHz | 35.929 | 1.3608 | 35.878 | 1.3595
3458A | ACV | 100 | 330 | 1 | V | mV | 20 | 50 | kHz | 533.77 | 0.66047 | 533.57 | 0.66042
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.11 | 3.4656
3458A | ACV | 100 | 330 | 0.001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.111 | 3.4656
3458A | ACV | 100 | 330 | 0.01 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.164 | 3.4655
3458A | ACV | 100 | 330 | 0.1 | V | mV | 50 | 100 | kHz | 24.914 | 3.4942 | 28.927 | 3.4523
3458A | ACV | 100 | 330 | 1 | V | mV | 50 | 100 | kHz | 418.09 | 2.6787 | 417.63 | 2.6793
3458A | DCV | 100 | 1000 | 1e-05 | V | mV |  |  |  | 0.11 | 0.0087 | 0.10064 | 0.0086547
3458A | DCV | 100 | 1000 | 0.0001 | V | mV |  |  |  | 0.11 | 0.0086529 | 0.10253 | 0.008653
3458A | DCV | 100 | 1000 | 0.001 | V | mV |  |  |  | 0.27989 | 0.0084976 | 0.27707 | 0.0084975
3458A | DCV | 100 | 1000 | 0.01 | V | mV |  |  |  | 5.3677 | 0.0051444 | 5.3663 | 0.0051359
3458A | DCV | 100 | 1000 | 0.1 | V | mV |  |  |  | 57.936 | 0.00072803 | 57.936 | 0.00072118
3458A | DCV | 100 | 1000 | 1 | V | mV |  |  |  | 579.99 | 7.53e-05 | 579.93 | 7.2528e-05
3458A | Resistance | 100 | 1000 | 0.0001 | Ω | µΩ |  |  |  | 590 | 12 | 568.01 | 11.812
3458A | Resistance | 100 | 1000 | 0.001 | Ω | µΩ |  |  |  | 647.66 | 11.942 | 669.04 | 11.724
3458A | Resistance | 100 | 1000 | 0.01 | Ω | µΩ |  |  |  | 5185.8 | 8.4755 | 5185 | 8.4748
3458A | Resistance | 100 | 1000 | 0.1 | Ω | µΩ |  |  |  | 57619 | 1.4299 | 57619 | 1.4287
3458A | Resistance | 100 | 1000 | 1 | Ω | µΩ |  |  |  | 577340 | 0.14501 | 577340 | 0.14451
3458A | Resistance | 100 | 1000 | 10 | Ω | µΩ |  |  |  | 5773600 | 0.014653 | 5773500 | 0.014453
Transmille | ACV | 100 | 1020 | 0.01 | V | mV | 0.04 | 1 | KHz | 360 | 1.1 | 337.15 | 1.12
Transmille | ACV | 100 | 1020 | 0.1 | V | mV | 0.04 | 1 | KHz | 360 | 1.11 | 341.11 | 1.11
Transmille | ACV | 100 | 1020 | 1 | V | mV | 0.04 | 1 | KHz | 639.27 | 0.929 | 638.46 | 0.928
Transmille | ACV | 100 | 1020 | 10 | V | mV | 0.04 | 1 | KHz | 5773 | 0.185 | 5772.7 | 0.183
Transmille | ACV | 100 | 1020 | 100 | V | mV | 0.04 | 1 | KHz | 57736 | 0.0194 | 57735 | 0.0186
Transmille | DCV | 100 | 1020 | 0.01 | V | mV |  |  |  | 36 | 0.092 | 35.843 | 0.0918
Transmille | DCV | 100 | 1020 | 0.1 | V | mV |  |  |  | 65.632 | 0.0746 | 65.525 | 0.0746
Transmille | DCV | 100 | 1020 | 1 | V | mV |  |  |  | 577.76 | 0.0138 | 577.71 | 0.0137
Transmille | DCV | 100 | 1020 | 10 | V | mV |  |  |  | 5773.6 | 0.00142 | 5773.6 | 0.00139
Transmille | DCV | 100 | 1020 | 100 | V | mV |  |  |  | 57736 | 0.000151 | 57736 | 0.000139
Transmille | Insulation DCV | 100 | 250 | 1 | V | V |  |  |  | 1.4 | 0.011 | 1.3333 | 0.0112
Transmille | Insulation DCV | 100 | 250 | 10 | V | V |  |  |  | 5.6899 | 0.00577 | 5.6798 | 0.00565
Transmille | Insulation DCV | 100 | 250 | 100 | V | V |  |  |  | 57.719 | 0.00246 | 57.716 | 0.00239
5502A | ACW | 109 | 363.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0012 | 0 | 0
5502A | ACW | 109 | 363.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0012284 | 0 | 0
5502A | ACW | 109 | 363.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.0054419 | 0 | 0
5502A | ACW | 109 | 363.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.053273 | 0 | 0
5502A | DCW | 109 | 330.0001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00015 | 0 | 0
5502A | DCW | 109 | 330.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.00015255 | 0 | 0
5502A | DCW | 109 | 330.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.00055202 | 0 | 0
5502A | DCW | 109 | 330.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.0053282 | 0 | 0
5502A | DCW | 109 | 330.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.05326 | 0 | 0
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 42 | 2.4 | 38.553 | 2.4159
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 46.787 | 2.3924 | 46.65 | 2.3916
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 500.59 | 1.4072 | 500.63 | 1.403
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5788.7 | 0.17673 | 5788.8 | 0.17407
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57999 | 0.018526 | 57999 | 0.017455
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 42.202 | 2.6324
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 49.765 | 2.6108 | 49.649 | 2.61
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 0.1 | 1 | kHz | 490.55 | 1.6184 | 490.58 | 1.6144
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 0.1 | 1 | kHz | 5786.7 | 0.2093 | 5786.8 | 0.20664
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 0.1 | 1 | kHz | 57999 | 0.021803 | 57999 | 0.020732
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 1 | 5 | kHz | 58 | 35 | 57.612 | 34.699
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 1 | 5 | kHz | 58 | 35 | 58.262 | 34.697
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 1 | 5 | kHz | 79.07 | 34.897 | 123.66 | 34.491
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 1 | 5 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 1 | 5 | kHz | 57768 | 3.2889 | 57768 | 3.2878
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 58 | 35 | 57.612 | 34.699
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 58 | 35 | 58.262 | 34.697
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 5 | 5.01 | kHz | 79.07 | 34.897 | 123.66 | 34.491
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 5 | 5.01 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 5 | 5.01 | kHz | 57768 | 3.2889 | 57768 | 3.2878
5502A | DCA 10 turn | 110 | 205 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
5502A | DCA 10 turn | 110 | 205 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
5502A | DCA 10 turn | 110 | 205 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
5502A | DCA 10 turn | 110 | 205 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
5502A | DCA 10 turn | 110 | 205 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
5502A | Resistance | 110 | 330.01 | 0.001 | Ω | mΩ |  |  |  | 25 | 0.3 | 24.812 | 0.29687
5502A | Resistance | 110 | 330.01 | 0.01 | Ω | mΩ |  |  |  | 25.358 | 0.29479 | 25.35 | 0.29441
5502A | Resistance | 110 | 330.01 | 0.1 | Ω | mΩ |  |  |  | 61.466 | 0.18435 | 61.46 | 0.18305
5502A | Resistance | 110 | 330.01 | 1 | Ω | mΩ |  |  |  | 580.25 | 0.024237 | 580.25 | 0.023449
5502A | CDK | 120 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.31026 | 0 | 0.30758 | 0
5502A | CDK | 120 | 1000.1 | 1 | °C | °C |  |  |  | 0.65669 | 0 | 0.65162 | 0
5502A | CDN | 120 | 410.1 | 0.1 | °C | °C |  |  |  | 0.22303 | 0 | 0.21855 | 0
5502A | CDN | 120 | 410.1 | 1 | °C | °C |  |  |  | 0.62103 | 0 | 0.61463 | 0
5502A | CDT | 120 | 400 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | CDT | 120 | 400 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0
5502A | Frequency | 120 | 1200.01 | 0.01 | Hz | Hz |  |  |  | 0.058 | 9.5e-06 | 0.0041865 | 2.6727e-05
5502A | Frequency | 120 | 1200.01 | 0.1 | Hz | Hz |  |  |  | 0.058519 | 1.0244e-05 | 0.056808 | 9.2702e-06
5502A | Frequency | 120 | 1200.01 | 1 | Hz | Hz |  |  |  | 0.57795 | 1.5973e-06 | 0.57725 | 1.0094e-06
5502A | Frequency | 120 | 1200.01 | 10 | Hz | Hz |  |  |  | 5.7739 | 3.3757e-07 | 5.7735 | 1.0104e-07
Transmille | TcindicatorK | 120 | 1370 | 0.1 | °C | °C |  |  |  | 1.4121 | 0 | 1.4113 | 0
Transmille | TcindicatorK | 120 | 1370 | 1 | °C | °C |  |  |  | 1.5267 | 0 | 1.5237 | 0
3458A | CDE | 149.001 | 593 | 1 | °C | °C |  |  |  | 0.57916 | 0 | 0.57842 | 0
3458A | CDE | 149.001 | 593 | 0.1 | °C | °C |  |  |  | 0.069131 | 0 | 0.067566 | 0
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 59 | 2.1 | 58.542 | 2.0879
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 62.205 | 2.0942 | 64.156 | 2.0802
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 445.36 | 1.6257 | 445.1 | 1.6253
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5769.8 | 0.28152 | 5769.8 | 0.28084
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57997 | 0.02869 | 57997 | 0.028411
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 74 | 2.3 | 73.171 | 2.2882
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 76.076 | 2.2962 | 78.162 | 2.2814
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 0.1 | 1 | kHz | 437.31 | 1.8466 | 437.03 | 1.8463
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 0.1 | 1 | kHz | 5764 | 0.34064 | 5763.9 | 0.33997
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 0.1 | 1 | kHz | 57996 | 0.034758 | 57996 | 0.034478
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.24 | 34.699
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.76 | 34.698
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 155.32 | 34.642
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 1 | 5 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 1 | 5 | kHz | 57174 | 7.1298 | 57174 | 7.1293
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.24 | 34.699
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.76 | 34.698
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 155.32 | 34.642
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 5 | 5.01 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 5 | 5.01 | kHz | 57174 | 7.1298 | 57174 | 7.1293
5502A | CDJ | 150 | 760.1 | 0.1 | °C | °C |  |  |  | 0.2089 | 0 | 0.20722 | 0
5502A | CDJ | 150 | 760.1 | 1 | °C | °C |  |  |  | 0.61296 | 0 | 0.61069 | 0
5502A | DCA 50 turn | 150 | 550.01 | 0.01 | A | mA |  |  |  | 19 | 2.2 | -17.053 | 2.1654
5502A | DCA 50 turn | 150 | 550.01 | 0.1 | A | mA |  |  |  | 19 | 2.2 | -10.261 | 2.1556
5502A | DCA 50 turn | 150 | 550.01 | 1 | A | mA |  |  |  | 409.49 | 1.6357 | 409.01 | 1.6349
5502A | DCA 50 turn | 150 | 550.01 | 10 | A | mA |  |  |  | 5740.6 | 0.27618 | 5740.5 | 0.27485
5502A | DCA 50 turn | 150 | 550.01 | 100 | A | mA |  |  |  | 57732 | 0.028329 | 57732 | 0.027783
Transmille | TcindicatorJ | 150 | 760 | 0.1 | °C | °C |  |  |  | 1.1342 | 0 | 1.1339 | 0
Transmille | TcindicatorJ | 150 | 760 | 1 | °C | °C |  |  |  | 1.2722 | 0 | 1.2711 | 0
3458A | FDT | 200 | 752 | 1 | °F | °F |  |  |  | 0.59635 | 0 | 0.58043 | 0
3458A | FDT | 200 | 752 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.083094 | 0
3458A | CDPT100 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.012329 | 0 | 0.011526 | 0
3458A | CDPT100 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.5777 | 0 | 0.57744 | 0
3458A | CDPT100 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.059241 | 0 | 0.058591 | 0
3458A | CDPT1000 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.011768 | 0 | 0.011474 | 0
3458A | CDPT1000 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.57753 | 0 | 0.57744 | 0
3458A | CDPT1000 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.058813 | 0 | 0.05858 | 0
5502A | DCA 10 turn | 205 | 206 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
5502A | DCA 10 turn | 205 | 206 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
5502A | DCA 10 turn | 205 | 206 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
5502A | DCA 10 turn | 205 | 206 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
5502A | DCA 10 turn | 205 | 206 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
5502A | FDPT100 | 212 | 572.1 | 0.1 | °F | °F |  |  |  | 0.28537 | 0 | 0.26385 | 0
5502A | FDPT100 | 212 | 572.1 | 1 | °F | °F |  |  |  | 0.61101 | 0 | 0.59895 | 0
5502A | FDPT1000 | 212 | 500.1 | 0.1 | °F | °F |  |  |  | 0.19201 | 0 | 0.15827 | 0
5502A | FDPT1000 | 212 | 500.1 | 1 | °F | °F |  |  |  | 0.59861 | 0 | 0.58629 | 0
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 84 | 2.4 | 76.978 | 2.4161
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 84 | 2.4069 | 81.077 | 2.4099
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 407.38 | 1.9712 | 407.21 | 1.9701
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5755.8 | 0.34673 | 5755.9 | 0.34526
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57996 | 0.03551 | 57996 | 0.034907
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 98 | 2.6 | 84.28 | 2.6326
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 98 | 2.6065 | 88.046 | 2.6269
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 0.1 | 1 | kHz | 396.27 | 2.2061 | 396.12 | 2.205
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 0.1 | 1 | kHz | 5747.7 | 0.41071 | 5747.8 | 0.40925
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 0.1 | 1 | kHz | 57995 | 0.042063 | 57995 | 0.04146
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.22 | 34.699
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.69 | 34.698
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 148.62 | 34.646
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 1 | 5 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 1 | 5 | kHz | 57090 | 6.4831 | 57090 | 6.4816
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.22 | 34.699
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.69 | 34.698
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 148.62 | 34.646
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 5 | 5.01 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 5 | 5.01 | kHz | 57090 | 6.4831 | 57090 | 6.4816
5502A | DCA 20 turn | 220 | 410 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
5502A | DCA 20 turn | 220 | 410 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
5502A | DCA 20 turn | 220 | 410 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
5502A | DCA 20 turn | 220 | 410 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
5502A | DCA 20 turn | 220 | 410 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
5502A | FDK | 248 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.55155 | 0 | 0.54665 | 0
5502A | FDK | 248 | 1832.1 | 1 | °F | °F |  |  |  | 0.80643 | 0 | 0.79299 | 0
5502A | FDN | 248 | 770.1 | 0.1 | °F | °F |  |  |  | 0.3915 | 0 | 0.38322 | 0
5502A | FDN | 248 | 770.1 | 1 | °F | °F |  |  |  | 0.70888 | 0 | 0.69055 | 0
5502A | FDT | 248 | 752 | 0.1 | °F | °F |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | FDT | 248 | 752 | 1 | °F | °F |  |  |  | 0.60528 | 0 | 0.60048 | 0
Transmille | TcindicatorK | 248 | 2500 | 0.1 | °F | °F |  |  |  | 2.5391 | 0 | 2.5376 | 0
Transmille | TcindicatorK | 248 | 2500 | 1 | °F | °F |  |  |  | 2.6075 | 0 | 2.6018 | 0
5502A | CDR | 250 | 400.1 | 0.1 | °C | °C |  |  |  | 0.41654 | 0 | 0.41163 | 0
5502A | CDR | 250 | 400.1 | 1 | °C | °C |  |  |  | 0.71813 | 0 | 0.70671 | 0
5502A | CDS | 250 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.42876 | 0 | 0.42399 | 0
5502A | CDS | 250 | 1000.1 | 1 | °C | °C |  |  |  | 0.72528 | 0 | 0.71398 | 0
Transmille | Insulation DCV | 250 | 500 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.2696 | 0.0114
Transmille | Insulation DCV | 250 | 500 | 10 | V | V |  |  |  | 5.1032 | 0.00803 | 5.0881 | 0.00793
Transmille | Insulation DCV | 250 | 500 | 100 | V | V |  |  |  | 57.497 | 0.00222 | 57.502 | 0.00214
Transmille | TcindicatorR | 250 | 1760 | 0.1 | °C | °C |  |  |  | 1.9965 | 0 | 1.9955 | 0
Transmille | TcindicatorR | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0804 | 0 | 2.0765 | 0
Transmille | TcindicatorS | 250 | 1760 | 0.1 | °C | °C |  |  |  | 2.0006 | 0 | 1.9996 | 0
Transmille | TcindicatorS | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0844 | 0 | 2.0805 | 0
5502A | CDPT1000 | 260 | 300.1 | 0.1 | °C | °C |  |  |  | 0.15328 | 0 | 0.14088 | 0
5502A | CDPT1000 | 260 | 300.1 | 1 | °C | °C |  |  |  | 0.60619 | 0 | 0.59404 | 0
3458A | CDJ | 260 | 732 | 1 | °C | °C |  |  |  | 0.58166 | 0 | 0.57849 | 0
3458A | CDJ | 260 | 732 | 0.1 | °C | °C |  |  |  | 0.073715 | 0 | 0.068215 | 0
3458A | CDK | 260 | 816 | 1 | °C | °C |  |  |  | 0.58441 | 0 | 0.5787 | 0
3458A | CDK | 260 | 816 | 0.1 | °C | °C |  |  |  | 0.08098 | 0 | 0.069993 | 0
5502A | CDPT100 | 300 | 400.1 | 0.1 | °C | °C |  |  |  | 0.22203 | 0 | 0.21366 | 0
5502A | CDPT100 | 300 | 400.1 | 1 | °C | °C |  |  |  | 0.62712 | 0 | 0.61537 | 0
5502A | CDPT1000 | 300 | 400.1 | 0.1 | °C | °C |  |  |  | 0.17 | 0 | 0.1589 | 0
5502A | CDPT1000 | 300 | 400.1 | 1 | °C | °C |  |  |  | 0.61063 | 0 | 0.59857 | 0
3458A | FDE | 300.001 | 1100 | 1 | °F | °F |  |  |  | 0.58299 | 0 | 0.5806 | 0
3458A | FDE | 300.001 | 1100 | 0.1 | °F | °F |  |  |  | 0.09 | 0 | 0.084258 | 0
5502A | FDJ | 302 | 1400.1 | 0.1 | °F | °F |  |  |  | 0.36584 | 0 | 0.36275 | 0
5502A | FDJ | 302 | 1400.1 | 1 | °F | °F |  |  |  | 0.686 | 0 | 0.6794 | 0
Transmille | TcindicatorJ | 302 | 1400 | 0.1 | °F | °F |  |  |  | 2.0393 | 0 | 2.0387 | 0
Transmille | TcindicatorJ | 302 | 1400 | 1 | °F | °F |  |  |  | 2.1202 | 0 | 2.1181 | 0
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 0.045 | 1 | Hz | 24 | 0.58 | 23.335 | 0.57834
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 0.045 | 1 | Hz | 33.354 | 0.57134 | 33.321 | 0.57119
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 0.045 | 1 | Hz | 510.18 | 0.32785 | 510.26 | 0.32724
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 0.045 | 1 | Hz | 5790.3 | 0.041506 | 5790.4 | 0.041123
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 0.045 | 1 | Hz | 57999 | 0.0042793 | 57999 | 0.004125
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 1 | 5.01 | kHz | 28 | 0.92 | 23.262 | 0.92437
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 1 | 5.01 | kHz | 29.895 | 0.91967 | 29.873 | 0.91957
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 1 | 5.01 | kHz | 446.23 | 0.66766 | 446.26 | 0.66723
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 1 | 5.01 | kHz | 5775.4 | 0.10274 | 5775.5 | 0.10236
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 1 | 5.01 | kHz | 57998 | 0.010468 | 57998 | 0.010314
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 5 | 10.01 | kHz | 64 | 1 | 23.23 | 1.0398
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 5 | 10.01 | kHz | 64 | 1.0355 | 29.165 | 1.0355
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 5 | 10.01 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 5 | 10.01 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 5 | 10.01 | kHz | 57997 | 0.013152 | 57997 | 0.012998
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 10 | 10.01 | kHz | 64 | 1 | 23.23 | 1.0398
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 10 | 10.01 | kHz | 64 | 1.0355 | 29.165 | 1.0355
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 10 | 10.01 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 10 | 10.01 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 10 | 10.01 | kHz | 57997 | 0.013152 | 57997 | 0.012998
5502A | DCV | 330 | 1000.01 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
5502A | DCV | 330 | 1000.01 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
5502A | DCV | 330 | 1000.01 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
5502A | DCV | 330 | 1000.01 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
5502A | DCV | 330 | 1000.01 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
5502A | DCW | 330 | 3000.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.0006 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.00062246 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.0018565 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.017586 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 100 | WDC | WDC |  |  |  | 0 | 0.17576 | 0 | 0
5502A | Resistance | 330 | 1100.01 | 0.01 | Ω | mΩ |  |  |  | 28 | 0.1 | 23.818 | 0.10366
5502A | Resistance | 330 | 1100.01 | 0.1 | Ω | mΩ |  |  |  | 52.968 | 0.087892 | 52.867 | 0.087789
5502A | Resistance | 330 | 1100.01 | 1 | Ω | mΩ |  |  |  | 577.2 | 0.017401 | 577.18 | 0.017221
5502A | Resistance | 330 | 1100.01 | 10 | Ω | mΩ |  |  |  | 5799.7 | 0.0018248 | 5799.7 | 0.00175
3458A | ACV | 330 | 1000 | 0.0001 | V | mV | 0 | 0 | kHz | 50 | 0.46 | 46.107 | 0.46338
3458A | ACV | 330 | 1000 | 0.001 | V | mV | 0 | 0 | kHz | 50 | 0.46054 | 46.108 | 0.46338
3458A | ACV | 330 | 1000 | 0.01 | V | mV | 0 | 0 | kHz | 50 | 0.46059 | 46.301 | 0.46322
3458A | ACV | 330 | 1000 | 0.1 | V | mV | 0 | 0 | kHz | 64.184 | 0.4486 | 64.129 | 0.44862
3458A | ACV | 330 | 1000 | 1 | V | mV | 0 | 0 | kHz | 564.15 | 0.20595 | 564.12 | 0.20589
3458A | ACV | 330 | 1000 | 0.0001 | V | mV | 0 | 1 | kHz | 27 | 0.46 | 23.031 | 0.46343
3458A | ACV | 330 | 1000 | 0.001 | V | mV | 0 | 1 | kHz | 27 | 0.46051 | 23.032 | 0.46343
3458A | ACV | 330 | 1000 | 0.01 | V | mV | 0 | 1 | kHz | 27 | 0.46056 | 23.29 | 0.46321
3458A | ACV | 330 | 1000 | 0.1 | V | mV | 0 | 1 | kHz | 45.921 | 0.44399 | 45.855 | 0.44402
3458A | ACV | 330 | 1000 | 1 | V | mV | 0 | 1 | kHz | 562.27 | 0.1928 | 562.23 | 0.19274
3458A | ACV | 330 | 1000 | 0.0001 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.056 | 0.69387
3458A | ACV | 330 | 1000 | 0.001 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.063 | 0.69386
3458A | ACV | 330 | 1000 | 0.01 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.256 | 0.69369
3458A | ACV | 330 | 1000 | 0.1 | V | mV | 1 | 10 | kHz | 40.883 | 0.68212 | 41.185 | 0.67806
3458A | ACV | 330 | 1000 | 1 | V | mV | 1 | 10 | kHz | 547.47 | 0.37317 | 547.4 | 0.3731
5502A | CDE | 350 | 650.1 | 0.1 | °C | °C |  |  |  | 0.19683 | 0 | 0.19628 | 0
5502A | CDE | 350 | 650.1 | 1 | °C | °C |  |  |  | 0.60777 | 0 | 0.60706 | 0
5502A | ACW | 363 | 990.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.001 | 0 | 0
5502A | ACW | 363 | 990.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.001007 | 0 | 0
5502A | ACW | 363 | 990.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.0018832 | 0 | 0
5502A | ACW | 363 | 990.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.01601 | 0 | 0
5502A | ACW | 363 | 990.0001 | 100 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
3458A | CDN | 371.001 | 816 | 1 | °C | °C |  |  |  | 0.58489 | 0 | 0.57864 | 0
3458A | CDN | 371.001 | 816 | 0.1 | °C | °C |  |  |  | 0.081513 | 0 | 0.069464 | 0
5502A | CDPT100 | 400 | 630.1 | 0.1 | °C | °C |  |  |  | 0.26 | 0 | 0.25288 | 0
5502A | CDPT100 | 400 | 630.1 | 1 | °C | °C |  |  |  | 0.64154 | 0 | 0.63006 | 0
5502A | CDPT1000 | 400 | 600.1 | 0.1 | °C | °C |  |  |  | 0.17204 | 0 | 0.16109 | 0
5502A | CDPT1000 | 400 | 600.1 | 1 | °C | °C |  |  |  | 0.61121 | 0 | 0.59915 | 0
5502A | CDR | 400 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.39417 | 0 | 0.38897 | 0
5502A | CDR | 400 | 1000.1 | 1 | °C | °C |  |  |  | 0.70538 | 0 | 0.69376 | 0
5502A | CDT | 400 | 400.1 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | CDT | 400 | 400.1 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0
3458A | FDPT100 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.021 | 0 | 0.018862 | 0
3458A | FDPT100 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57849 | 0 | 0.57763 | 0
3458A | FDPT100 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.062482 | 0 | 0.060463 | 0
3458A | FDPT1000 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.02 | 0 | 0.018759 | 0
3458A | FDPT1000 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57793 | 0 | 0.57763 | 0
3458A | FDPT1000 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.061159 | 0 | 0.060431 | 0
5502A | CDN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
5502A | CDN | 410 | 1300 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
5502A | DCA 20 turn | 410 | 411 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
5502A | DCA 20 turn | 410 | 411 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
5502A | DCA 20 turn | 410 | 411 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
5502A | DCA 20 turn | 410 | 411 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
5502A | DCA 20 turn | 410 | 411 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
Transmille | TcindicatorN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 1.3658 | 0 | 1.3651 | 0
Transmille | TcindicatorN | 410 | 1300 | 1 | °C | °C |  |  |  | 1.4837 | 0 | 1.481 | 0
3458A | CDPT100 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015887 | 0 | 0.015272 | 0
3458A | CDPT100 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57752 | 0
3458A | CDPT100 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060082 | 0 | 0.059441 | 0
3458A | CDPT1000 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015792 | 0 | 0.015173 | 0
3458A | CDPT1000 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57752 | 0
3458A | CDPT1000 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060057 | 0 | 0.059416 | 0
5502A | FDR | 482 | 752.1 | 0.1 | °F | °F |  |  |  | 0.74217 | 0 | 0.73324 | 0
5502A | FDR | 482 | 752.1 | 1 | °F | °F |  |  |  | 0.95934 | 0 | 0.93147 | 0
5502A | FDS | 482 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.76311 | 0 | 0.75443 | 0
5502A | FDS | 482 | 1832.1 | 1 | °F | °F |  |  |  | 0.97706 | 0 | 0.94824 | 0
Transmille | TcindicatorR | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5823 | 0 | 3.5805 | 0
Transmille | TcindicatorR | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6336 | 0 | 3.6263 | 0
Transmille | TcindicatorS | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5846 | 0 | 3.5828 | 0
Transmille | TcindicatorS | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6358 | 0 | 3.6285 | 0
5502A | FDPT1000 | 500 | 572.1 | 0.1 | °F | °F |  |  |  | 0.2619 | 0 | 0.23826 | 0
5502A | FDPT1000 | 500 | 572.1 | 1 | °F | °F |  |  |  | 0.61003 | 0 | 0.59795 | 0
3458A | FDJ | 500 | 1350 | 1 | °F | °F |  |  |  | 0.59113 | 0 | 0.58077 | 0
3458A | FDJ | 500 | 1350 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.085419 | 0
3458A | FDK | 500 | 1500 | 1 | °F | °F |  |  |  | 0.59947 | 0 | 0.58127 | 0
3458A | FDK | 500 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.088716 | 0
Transmille | Insulation DCV | 500 | 1000 | 1 | V | V |  |  |  | 1.8 | 0.011 | 1.2242 | 0.0115
Transmille | Insulation DCV | 500 | 1000 | 10 | V | V |  |  |  | 4.1693 | 0.00987 | 4.0987 | 0.00988
Transmille | Insulation DCV | 500 | 1000 | 100 | V | V |  |  |  | 57.066 | 0.00249 | 57.041 | 0.00245
5502A | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 210 | 2.4 | 192.41 | 2.4161
5502A | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 210 | 2.4 | 193.98 | 2.4152
5502A | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 350.89 | 2.3236 | 350.74 | 2.3224
5502A | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5545.6 | 0.82078 | 5546.4 | 0.81772
5502A | ACA 50 turn | 550 | 1025.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57972 | 0.088607 | 57972 | 0.087215
5502A | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
5502A | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
5502A | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 0.1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
5502A | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 0.1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
5502A | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
5502A | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
5502A | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
5502A | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
5502A | DCA 50 turn | 550 | 1025 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
5502A | DCA 50 turn | 550 | 1025 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
5502A | DCA 50 turn | 550 | 1025 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
5502A | DCA 50 turn | 550 | 1025 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
5502A | DCA 50 turn | 550 | 1025 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
5502A | FDPT100 | 572 | 752.1 | 0.1 | °F | °F |  |  |  | 0.39012 | 0 | 0.37466 | 0
5502A | FDPT100 | 572 | 752.1 | 1 | °F | °F |  |  |  | 0.63118 | 0 | 0.61951 | 0
5502A | FDPT1000 | 572 | 752.1 | 0.1 | °F | °F |  |  |  | 0.29342 | 0 | 0.27254 | 0
5502A | FDPT1000 | 572 | 752.1 | 1 | °F | °F |  |  |  | 0.61481 | 0 | 0.60282 | 0
3458A | CDE | 593.001 | 1000 | 1 | °C | °C |  |  |  | 0.57945 | 0 | 0.57871 | 0
3458A | CDE | 593.001 | 1000 | 0.1 | °C | °C |  |  |  | 0.071536 | 0 | 0.070024 | 0
3458A | CDR | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59357 | 0 | 0.57971 | 0
3458A | CDR | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.10063 | 0 | 0.077849 | 0
3458A | CDS | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59384 | 0 | 0.57998 | 0
3458A | CDS | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.10063 | 0 | 0.079868 | 0
5502A | CDPT1000 | 600 | 630 | 0.1 | °C | °C |  |  |  | 0.47301 | 0 | 0.46914 | 0
5502A | CDPT1000 | 600 | 630 | 1 | °C | °C |  |  |  | 0.75347 | 0 | 0.74372 | 0
5502A | CDPT100 | 630 | 800 | 0.1 | °C | °C |  |  |  | 0.43497 | 0 | 0.43075 | 0
5502A | CDPT100 | 630 | 800 | 1 | °C | °C |  |  |  | 0.73019 | 0 | 0.72012 | 0
5502A | CDPT1000 | 630 | 630.1 | 0.1 | °C | °C |  |  |  | 0.47301 | 0 | 0.46914 | 0
5502A | CDPT1000 | 630 | 630.1 | 1 | °C | °C |  |  |  | 0.75347 | 0 | 0.74372 | 0
3458A | CDPT100 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019744 | 0 | 0.01888 | 0
3458A | CDPT100 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57809 | 0 | 0.57763 | 0
3458A | CDPT100 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.061563 | 0 | 0.060469 | 0
3458A | CDPT1000 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019042 | 0 | 0.018733 | 0
3458A | CDPT1000 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57763 | 0
3458A | CDPT1000 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.060808 | 0 | 0.060423 | 0
5502A | CDE | 650 | 1000 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
5502A | CDE | 650 | 1000 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
5502A | FDE | 662 | 1202.1 | 0.1 | °F | °F |  |  |  | 0.34352 | 0 | 0.34251 | 0
5502A | FDE | 662 | 1202.1 | 1 | °F | °F |  |  |  | 0.67089 | 0 | 0.66882 | 0
3458A | FDN | 700.001 | 1500 | 1 | °F | °F |  |  |  | 0.60109 | 0 | 0.58115 | 0
3458A | FDN | 700.001 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.087954 | 0
3458A | CDJ | 732.001 | 1200 | 1 | °C | °C |  |  |  | 0.58132 | 0 | 0.57892 | 0
3458A | CDJ | 732.001 | 1200 | 0.1 | °C | °C |  |  |  | 0.076447 | 0 | 0.071753 | 0
5502A | FDPT100 | 752 | 1166.1 | 0.1 | °F | °F |  |  |  | 0.45987 | 0 | 0.44683 | 0
5502A | FDPT100 | 752 | 1166.1 | 1 | °F | °F |  |  |  | 0.64673 | 0 | 0.63534 | 0
5502A | FDPT1000 | 752 | 1112.1 | 0.1 | °F | °F |  |  |  | 0.29725 | 0 | 0.27666 | 0
5502A | FDPT1000 | 752 | 1112.1 | 1 | °F | °F |  |  |  | 0.61665 | 0 | 0.6047 | 0
5502A | FDR | 752 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.70146 | 0 | 0.69201 | 0
5502A | FDR | 752 | 1832.1 | 1 | °F | °F |  |  |  | 0.92821 | 0 | 0.89938 | 0
5502A | FDT | 752 | 752.1 | 0.1 | °F | °F |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | FDT | 752 | 752.1 | 1 | °F | °F |  |  |  | 0.60528 | 0 | 0.60048 | 0
5502A | CDJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
5502A | CDJ | 760 | 1200 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
Transmille | TcindicatorJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 1.2495 | 0 | 1.2492 | 0
Transmille | TcindicatorJ | 760 | 1200 | 1 | °C | °C |  |  |  | 1.376 | 0 | 1.375 | 0
5502A | FDN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
5502A | FDN | 770 | 2372 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
Transmille | TcindicatorN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 2.4558 | 0 | 2.4545 | 0
Transmille | TcindicatorN | 770 | 2372 | 1 | °F | °F |  |  |  | 2.5259 | 0 | 2.5208 | 0
5502A | CDPT100 | 800 | 800.1 | 0.1 | °C | °C |  |  |  | 0.43497 | 0 | 0.43075 | 0
5502A | CDPT100 | 800 | 800.1 | 1 | °C | °C |  |  |  | 0.73019 | 0 | 0.72012 | 0
3458A | FDPT100 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.027261 | 0 | 0.026097 | 0
3458A | FDPT100 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57877 | 0 | 0.57791 | 0
3458A | FDPT100 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.065034 | 0 | 0.063096 | 0
3458A | FDPT1000 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.02708 | 0 | 0.025909 | 0
3458A | FDPT1000 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57876 | 0 | 0.5779 | 0
3458A | FDPT1000 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.064958 | 0 | 0.063018 | 0
3458A | CDK | 816.001 | 1371 | 1 | °C | °C |  |  |  | 0.58744 | 0 | 0.57956 | 0
3458A | CDK | 816.001 | 1371 | 0.1 | °C | °C |  |  |  | 0.090471 | 0 | 0.076709 | 0
3458A | CDN | 816.001 | 1300 | 1 | °C | °C |  |  |  | 0.58599 | 0 | 0.57919 | 0
3458A | CDN | 816.001 | 1300 | 0.1 | °C | °C |  |  |  | 0.086274 | 0 | 0.073915 | 0
5502A | ACW | 990 | 3630.1 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 990 | 3630.1 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 990 | 3630.1 | 10 | WAC | WAC |  |  |  | 0 | 0.0061979 | 0 | 0
5502A | ACW | 990 | 3630.1 | 100 | WAC | WAC |  |  |  | 0 | 0.058622 | 0 | 0
5502A | ACW | 990 | 3630.1 | 1000 | WAC | WAC |  |  |  | 0 | 0.58586 | 0 | 0
5502A | CDE | 1000 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
5502A | CDE | 1000 | 1000.1 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
5502A | CDK | 1000 | 1372 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
5502A | CDK | 1000 | 1372 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
5502A | CDR | 1000 | 1767 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
5502A | CDR | 1000 | 1767 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
5502A | CDS | 1000 | 1400.1 | 0.1 | °C | °C |  |  |  | 0.4386 | 0 | 0.43394 | 0
5502A | CDS | 1000 | 1400.1 | 1 | °C | °C |  |  |  | 0.73114 | 0 | 0.71993 | 0
5502A | DCV | 1000 | 1001 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
5502A | DCV | 1000 | 1001 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
5502A | DCV | 1000 | 1001 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
5502A | DCV | 1000 | 1001 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
5502A | DCV | 1000 | 1001 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
3458A | Frequency | 1000 | 10000 | 0.001 | Hz | Hz |  |  |  | 4.4e-09 | 1.3e-07 | 7.41e-10 | 1.2922e-07
3458A | Frequency | 1000 | 10000 | 0.01 | Hz | Hz |  |  |  | 1.95e-07 | 1.292e-07 | 1.38e-07 | 1.2921e-07
3458A | Frequency | 1000 | 10000 | 0.1 | Hz | Hz |  |  |  | 1.3574e-05 | 1.2799e-07 | 1.3571e-05 | 1.2799e-07
3458A | Frequency | 1000 | 10000 | 1 | Hz | Hz |  |  |  | 0.00068566 | 9.1517e-08 | 0.00068565 | 9.1518e-08
3458A | Frequency | 1000 | 10000 | 10 | Hz | Hz |  |  |  | 0.0057844 | 1.571e-08 | 0.0057844 | 1.571e-08
3458A | Frequency | 1000 | 10000 | 100 | Hz | Hz |  |  |  | 0.057734 | 1.59e-09 | 0.05772 | 1.59e-09
3458A | Resistance | 1000 | 10000 | 0.001 | Ω | mΩ |  |  |  | 5.9 | 0.012 | 5.8778 | 0.011642
3458A | Resistance | 1000 | 10000 | 0.01 | Ω | mΩ |  |  |  | 6.5111 | 0.011939 | 6.8827 | 0.011555
3458A | Resistance | 1000 | 10000 | 0.1 | Ω | mΩ |  |  |  | 52.024 | 0.0083239 | 52.009 | 0.0083231
3458A | Resistance | 1000 | 10000 | 1 | Ω | mΩ |  |  |  | 576.23 | 0.0013955 | 576.23 | 0.0013939
3458A | Resistance | 1000 | 10000 | 10 | Ω | mΩ |  |  |  | 5773.4 | 0.0001416 | 5773.4 | 0.00014095
3458A | Resistance | 1000 | 10000 | 100 | Ω | mΩ |  |  |  | 57736 | 1.4357e-05 | 57735 | 1.4097e-05
34420A | Resistance | 1000 | 10000 | 1e-05 | Ohms | µΩ |  |  |  | 24.0351 | 0.069214 | 23.2955 | 0.069264
34420A | Resistance | 1000 | 10000 | 0.0001 | Ohms | µΩ |  |  |  | 66.5586 | 0.066056 | 41.3894 | 0.067687
34420A | Resistance | 1000 | 10000 | 0.001 | Ohms | µΩ |  |  |  | 643.2147 | 0.054751 | 550.1922 | 0.037118
34420A | Resistance | 1000 | 10000 | 0.01 | Ohms | mΩ |  |  |  | 6.0044 | 9.5076e-05 | 5.7959 | 4.809e-06
34420A | Resistance | 1000 | 10000 | 0.1 | Ohms | mΩ |  |  |  | 58.5135 | 0.00024789 | 57.9996 | 4.8274e-07
34420A | Resistance | 1000 | 10000 | 10 | Ohm | mΩ |  |  |  | 24.0351 | 6.9214e-05 | 23.2955 | 6.9264e-05
34420A | Resistance | 1000 | 10000 | 100 | Ohm | mΩ |  |  |  | 66.5586 | 6.6056e-05 | 41.3894 | 6.7688e-05
34420A | Resistance | 1000 | 10000 | 1000 | Ohm | mΩ |  |  |  | 643.2147 | 5.4751e-05 | 550.192 | 3.7118e-05
34420A | Resistance | 1000 | 10000 | 10000 | Ohm | Ω |  |  |  | 6.0044 | 9.5076e-08 | 5.7959 | 4.809e-09
34420A | Resistance | 1000 | 10000 | 100000 | Ohm | Ω |  |  |  | 58.5135 | 2.4789e-07 | 57.9996 | 4.83e-10
5502A | DCA 50 turn | 1025 | 1026 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
5502A | DCA 50 turn | 1025 | 1026 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
5502A | DCA 50 turn | 1025 | 1026 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
5502A | DCA 50 turn | 1025 | 1026 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
5502A | DCA 50 turn | 1025 | 1026 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
5502A | Resistance | 1100 | 3300.1 | 0.01 | Ω | Ω |  |  |  | 0.25 | 0.0001 | 0.23096 | 0.00010411
5502A | Resistance | 1100 | 3300.1 | 0.1 | Ω | Ω |  |  |  | 0.25 | 0.00010329 | 0.23666 | 0.00010325
5502A | Resistance | 1100 | 3300.1 | 1 | Ω | Ω |  |  |  | 0.60487 | 6.4429e-05 | 0.60441 | 6.4221e-05
5502A | Resistance | 1100 | 3300.1 | 10 | Ω | Ω |  |  |  | 5.8014 | 8.3761e-06 | 5.8012 | 8.2294e-06
3458A | FDE | 1100.001 | 1832 | 1 | °F | °F |  |  |  | 0.58352 | 0 | 0.58114 | 0
3458A | FDE | 1100.001 | 1832 | 0.1 | °F | °F |  |  |  | 0.091751 | 0 | 0.08789 | 0
3458A | FDR | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.6276 | 0 | 0.5841 | 0
3458A | FDR | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.16 | 0 | 0.10571 | 0
3458A | FDS | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.62833 | 0 | 0.5849 | 0
3458A | FDS | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.10571 | 0
5502A | FDPT1000 | 1112 | 1166 | 0.1 | °F | °F |  |  |  | 0.84699 | 0 | 0.83998 | 0
5502A | FDPT1000 | 1112 | 1166 | 1 | °F | °F |  |  |  | 1.0262 | 0 | 1.0191 | 0
5502A | FDPT100 | 1166 | 1472 | 0.1 | °F | °F |  |  |  | 0.77811 | 0 | 0.77048 | 0
5502A | FDPT100 | 1166 | 1472 | 1 | °F | °F |  |  |  | 0.739 | 0 | 0.72906 | 0
5502A | FDPT1000 | 1166 | 1166.1 | 0.1 | °F | °F |  |  |  | 0.84699 | 0 | 0.83998 | 0
5502A | FDPT1000 | 1166 | 1166.1 | 1 | °F | °F |  |  |  | 1.0262 | 0 | 1.0191 | 0
5502A | CDJ | 1200 | 1200.1 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
5502A | CDJ | 1200 | 1200.1 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
5502A | Frequency | 1200 | 12001 | 0.1 | Hz | Hz |  |  |  | 0.58 | 9.1e-06 | 0.036333 | 2.6333e-05
5502A | Frequency | 1200 | 12001 | 0.1 | Hz | Hz |  |  |  | 0.58497 | 9.8286e-06 | 0.56786 | 8.8368e-06
5502A | Frequency | 1200 | 12001 | 1 | Hz | Hz |  |  |  | 5.7795 | 1.5456e-06 | 5.7725 | 9.5754e-07
5502A | Frequency | 1200 | 12001 | 10 | Hz | Hz |  |  |  | 57.739 | 3.3237e-07 | 57.735 | 9.584e-08
3458A | FDPT100 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.034473 | 0 | 0.032867 | 0
3458A | FDPT100 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57975 | 0 | 0.57826 | 0
3458A | FDPT100 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.069375 | 0 | 0.066183 | 0
3458A | FDPT1000 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.033169 | 0 | 0.032594 | 0
3458A | FDPT1000 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57876 | 0 | 0.57824 | 0
3458A | FDPT1000 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.067183 | 0 | 0.066048 | 0
5502A | FDE | 1202 | 1832 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
5502A | FDE | 1202 | 1832 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
3458A | CDR | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.595 | 0 | 0.58117 | 0
3458A | CDR | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.10872 | 0 | 0.088053 | 0
3458A | CDS | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.59549 | 0 | 0.58168 | 0
3458A | CDS | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.11141 | 0 | 0.091359 | 0
5502A | CDN | 1300 | 1300.1 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
5502A | CDN | 1300 | 1300.1 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
3458A | FDJ | 1350.001 | 2192 | 1 | °F | °F |  |  |  | 0.58927 | 0 | 0.58157 | 0
3458A | FDJ | 1350.001 | 2192 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.090692 | 0
5502A | CDK | 1372 | 1372.1 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
5502A | CDK | 1372 | 1372.1 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
5502A | CDS | 1400 | 1767 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
5502A | CDS | 1400 | 1767 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
5502A | FDJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
5502A | FDJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
Transmille | TcindicatorJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 2.2469 | 0 | 2.2464 | 0
Transmille | TcindicatorJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 2.3207 | 0 | 2.3187 | 0
5502A | FDPT100 | 1472 | 1472.1 | 0.1 | °F | °F |  |  |  | 0.77811 | 0 | 0.77048 | 0
5502A | FDPT100 | 1472 | 1472.1 | 1 | °F | °F |  |  |  | 0.739 | 0 | 0.72906 | 0
3458A | FDK | 1500.001 | 2500 | 1 | °F | °F |  |  |  | 0.60796 | 0 | 0.58292 | 0
3458A | FDK | 1500.001 | 2500 | 0.1 | °F | °F |  |  |  | 0.13135 | 0 | 0.098982 | 0
3458A | FDN | 1500.001 | 2372 | 1 | °F | °F |  |  |  | 0.60386 | 0 | 0.58223 | 0
3458A | FDN | 1500.001 | 2372 | 0.1 | °F | °F |  |  |  | 0.13 | 0 | 0.094806 | 0
5502A | CDR | 1767 | 1767.1 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
5502A | CDR | 1767 | 1767.1 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
5502A | CDS | 1767 | 1767.1 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
5502A | CDS | 1767 | 1767.1 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
5502A | FDE | 1832 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
5502A | FDE | 1832 | 1832.1 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
5502A | FDK | 1832 | 2500 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
5502A | FDK | 1832 | 2500 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
5502A | FDR | 1832 | 3213 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
5502A | FDR | 1832 | 3213 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
5502A | FDS | 1832 | 2552.1 | 0.1 | °F | °F |  |  |  | 0.78274 | 0 | 0.77428 | 0
5502A | FDS | 1832 | 2552.1 | 1 | °F | °F |  |  |  | 0.99106 | 0 | 0.96411 | 0
5502A | FDJ | 2192 | 2192.1 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
5502A | FDJ | 2192 | 2192.1 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
3458A | FDR | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.63063 | 0 | 0.58736 | 0
3458A | FDR | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.12242 | 0
3458A | FDS | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.63063 | 0 | 0.58878 | 0
3458A | FDS | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.12242 | 0
5502A | FDN | 2372 | 2372.1 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
5502A | FDN | 2372 | 2372.1 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
5502A | FDK | 2500 | 2500.1 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
5502A | FDK | 2500 | 2500.1 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
5502A | FDS | 2552 | 3213 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
5502A | FDS | 2552 | 3213 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
5502A | DCW | 3000 | 11000.1 | 10 | WDC | WDC |  |  |  | 0 | 0.00091 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 100 | WDC | WDC |  |  |  | 0 | 0.00091 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 1000 | WDC | WDC |  |  |  | 0 | 0.00092998 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 10000 | WDC | WDC |  |  |  | 0 | 0.0021367 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 100000 | WDC | WDC |  |  |  | 0 | 0.019355 | 0 | 0
5502A | FDR | 3213 | 3213.1 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
5502A | FDR | 3213 | 3213.1 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
5502A | FDS | 3213 | 3213.1 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
5502A | FDS | 3213 | 3213.1 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
5502A | Resistance | 3300 | 11001 | 0.1 | Ω | Ω |  |  |  | 0.17 | 0.0001 | 0.12474 | 0.00010353
5502A | Resistance | 3300 | 11001 | 1 | Ω | Ω |  |  |  | 0.46536 | 8.4151e-05 | 0.46464 | 8.4044e-05
5502A | Resistance | 3300 | 11001 | 10 | Ω | Ω |  |  |  | 5.7682 | 1.5398e-05 | 5.7682 | 1.5238e-05
5502A | Resistance | 3300 | 11001 | 100 | Ω | Ω |  |  |  | 57.997 | 1.6098e-06 | 57.997 | 1.5438e-06
5502A | ACW | 3630 | 11200.1 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0016 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 1 | WAC | WAC |  |  |  | 0 | 0.0016 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 10 | WAC | WAC |  |  |  | 0 | 0.0022012 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 100 | WAC | WAC |  |  |  | 0 | 0.016051 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 1000 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
3458A | Frequency | 10000 | 100000 | 0.01 | Hz | Hz |  |  |  | 4.4e-08 | 7.41e-12 | 7.41e-09 | 1.2922e-07
3458A | Frequency | 10000 | 100000 | 0.1 | Hz | Hz |  |  |  | 1.95e-06 | 1.38e-09 | 1.38e-06 | 1.2921e-07
3458A | Frequency | 10000 | 100000 | 1 | Hz | Hz |  |  |  | 0.00013574 | 1.3571e-07 | 0.00013571 | 1.2799e-07
3458A | Frequency | 10000 | 100000 | 10 | Hz | Hz |  |  |  | 0.0068566 | 6.8565e-06 | 0.0068565 | 9.1518e-08
3458A | Frequency | 10000 | 100000 | 100 | Hz | Hz |  |  |  | 0.057844 | 5.7844e-05 | 0.057844 | 1.571e-08
3458A | Frequency | 10000 | 100000 | 1000 | Hz | Hz |  |  |  | 0.57734 | 0.0005772 | 0.5772 | 1.59e-09
3458A | Frequency | 10000 | 100000 | 10000 | Hz | Hz |  |  |  | 5.7736 | 0.0057734 | 5.7734 | 1.59e-10
3458A | Resistance | 10000 | 100000 | 0.01 | Ω | mΩ |  |  |  | 58 | 0.012 | 53.873 | 0.012027
3458A | Resistance | 10000 | 100000 | 0.1 | Ω | mΩ |  |  |  | 64.108 | 0.011939 | 63.986 | 0.011939
3458A | Resistance | 10000 | 100000 | 1 | Ω | mΩ |  |  |  | 516.51 | 0.0086656 | 516.37 | 0.0086649
3458A | Resistance | 10000 | 100000 | 10 | Ω | mΩ |  |  |  | 5761.5 | 0.0014742 | 5761.5 | 0.0014726
3458A | Resistance | 10000 | 100000 | 100 | Ω | mΩ |  |  |  | 57734 | 0.00014965 | 57734 | 0.000149
3458A | Resistance | 10000 | 100000 | 1000 | Ω | mΩ |  |  |  | 577360 | 1.5162e-05 | 577350 | 1.4902e-05
34420A | Resistance | 10000 | 100000 | 0.0001 | Ohms | mΩ |  |  |  | 0.24035 | 6.9214e-05 | 0.23295 | 6.9264e-05
34420A | Resistance | 10000 | 100000 | 0.001 | Ohms | mΩ |  |  |  | 0.66559 | 6.6056e-05 | 0.41389 | 6.7687e-05
34420A | Resistance | 10000 | 100000 | 0.01 | Ohms | mΩ |  |  |  | 6.4321 | 5.4751e-05 | 5.5019 | 3.7118e-05
34420A | Resistance | 10000 | 100000 | 0.1 | Ohms | mΩ |  |  |  | 60.0444 | 9.5076e-05 | 57.9593 | 4.809e-06
34420A | Resistance | 10000 | 100000 | 1 | Ohms | mΩ |  |  |  | 585.1347 | 0.00024789 | 579.9959 | 4.83e-07
34420A | Resistance | 10000 | 100000 | 100 | Ohm | Ω |  |  |  | 0.46937 | 6.92e-08 | 0.46349 | 6.93e-08
34420A | Resistance | 10000 | 100000 | 1000 | Ohm | Ω |  |  |  | 0.82928 | 6.67e-08 | 0.61213 | 6.81e-08
34420A | Resistance | 10000 | 100000 | 10000 | Ohm | Ω |  |  |  | 6.4547 | 5.59e-08 | 5.5271 | 3.87e-08
34420A | Resistance | 10000 | 100000 | 100000 | Ohm | Ω |  |  |  | 60.0461 | 9.53e-08 | 57.9607 | 5.09e-09
34420A | Resistance | 10000 | 100000 | 1000000 | Ohm | Ω |  |  |  | 585.1349 | 2.48e-07 | 579.996 | 5.1e-10
5502A | DCW | 11000 | 24500 | 10 | WDC | WDC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 11000 | 24500 | 100 | WDC | WDC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 11000 | 24500 | 1000 | WDC | WDC |  |  |  | 0 | 0.0014429 | 0 | 0
5502A | DCW | 11000 | 24500 | 10000 | WDC | WDC |  |  |  | 0 | 0.0054411 | 0 | 0
5502A | DCW | 11000 | 24500 | 100000 | WDC | WDC |  |  |  | 0 | 0.052745 | 0 | 0
5502A | Resistance | 11000 | 33001 | 0.1 | Ω | Ω |  |  |  | 1.3 | 0.0001 | 1.1561 | 0.00010412
5502A | Resistance | 11000 | 33001 | 1 | Ω | Ω |  |  |  | 1.3 | 0.00010253 | 1.2442 | 0.00010252
5502A | Resistance | 11000 | 33001 | 10 | Ω | Ω |  |  |  | 5.6614 | 5.2695e-05 | 5.6605 | 5.2621e-05
5502A | Resistance | 11000 | 33001 | 100 | Ω | Ω |  |  |  | 57.978 | 6.2175e-06 | 57.978 | 6.1738e-06
5502A | ACW | 11200 | 20900 | 0.1 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
5502A | ACW | 11200 | 20900 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 11200 | 20900 | 10 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 11200 | 20900 | 100 | WAC | WAC |  |  |  | 0 | 0.0055466 | 0 | 0
5502A | ACW | 11200 | 20900 | 1000 | WAC | WAC |  |  |  | 0 | 0.051733 | 0 | 0
5502A | Frequency | 12000 | 120010 | 1 | Hz | Hz |  |  |  | 19 | 2.8e-05 | 17.33 | 2.8854e-05
5502A | Frequency | 12000 | 120010 | 10 | Hz | Hz |  |  |  | 19 | 2.8e-05 | 18.255 | 2.7641e-05
5502A | Frequency | 12000 | 120010 | 100 | Hz | Hz |  |  |  | 60.592 | 1.1287e-05 | 60.269 | 9.1178e-06
5502A | Frequency | 12000 | 120010 | 1000 | Hz | Hz |  |  |  | 577.75 | 1.9059e-06 | 577.61 | 9.6075e-07
5502A | ACW | 20900 | 20901 | 0.1 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
5502A | ACW | 20900 | 20901 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 20900 | 20901 | 10 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 20900 | 20901 | 100 | WAC | WAC |  |  |  | 0 | 0.0055466 | 0 | 0
5502A | ACW | 20900 | 20901 | 1000 | WAC | WAC |  |  |  | 0 | 0.051733 | 0 | 0
5502A | DCW | 24500 | 24501 | 10 | WAC | WAC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 24500 | 24501 | 100 | WAC | WAC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 24500 | 24501 | 1000 | WAC | WAC |  |  |  | 0 | 0.0014429 | 0 | 0
5502A | DCW | 24500 | 24501 | 10000 | WAC | WAC |  |  |  | 0 | 0.0054411 | 0 | 0
5502A | DCW | 24500 | 24501 | 100000 | WAC | WAC |  |  |  | 0 | 0.052745 | 0 | 0
5502A | Resistance | 33000 | 110010 | 1 | Ω | Ω |  |  |  | 1.3 | 0.00013 | 1.2347 | 0.00012665
5502A | Resistance | 33000 | 110010 | 10 | Ω | Ω |  |  |  | 4.353 | 0.00010805 | 4.3501 | 0.00010797
5502A | Resistance | 33000 | 110010 | 100 | Ω | Ω |  |  |  | 57.521 | 2.2173e-05 | 57.522 | 2.2051e-05
5502A | Resistance | 33000 | 110010 | 1000 | Ω | Ω |  |  |  | 579.95 | 2.2971e-06 | 579.95 | 2.2465e-06
3458A | Frequency | 100000 | 1000000 | 0.1 | Hz | Hz |  |  |  | 4.4e-07 | 1.3e-07 | 7.41e-08 | 1.2922e-07
3458A | Frequency | 100000 | 1000000 | 1 | Hz | Hz |  |  |  | 1.9518e-05 | 1.292e-07 | 1.3837e-05 | 1.2921e-07
3458A | Frequency | 100000 | 1000000 | 10 | Hz | Hz |  |  |  | 0.0013574 | 1.2799e-07 | 0.0013571 | 1.2799e-07
3458A | Frequency | 100000 | 1000000 | 100 | Hz | Hz |  |  |  | 0.068566 | 9.1517e-08 | 0.068565 | 9.1518e-08
3458A | Frequency | 100000 | 1000000 | 1000 | Hz | Hz |  |  |  | 0.57844 | 1.571e-08 | 0.57844 | 1.571e-08
3458A | Frequency | 100000 | 1000000 | 10000 | Hz | Hz |  |  |  | 5.7734 | 1.59e-09 | 5.772 | 1.59e-09
3458A | Resistance | 100000 | 1000000 | 0.1 | Ω | Ω |  |  |  | 2.3 | 1.8e-05 | 2.2687 | 1.7871e-05
3458A | Resistance | 100000 | 1000000 | 1 | Ω | Ω |  |  |  | 2.3 | 1.8e-05 | 2.3129 | 1.7835e-05
3458A | Resistance | 100000 | 1000000 | 10 | Ω | Ω |  |  |  | 5.5163 | 1.5438e-05 | 5.5116 | 1.5439e-05
3458A | Resistance | 100000 | 1000000 | 100 | Ω | Ω |  |  |  | 57.516 | 3.6361e-06 | 57.515 | 3.6327e-06
3458A | Resistance | 100000 | 1000000 | 1000 | Ω | Ω |  |  |  | 577.33 | 3.7584e-07 | 577.33 | 3.7433e-07
3458A | Resistance | 100000 | 1000000 | 10000 | Ω | Ω |  |  |  | 5773.6 | 3.8048e-08 | 5773.5 | 3.7445e-08
34420A | Resistance | 100000 | 1000 | 1 | Ohm | mΩ |  |  |  | 2.4035 | 6.9214e-05 | 2.3295 | 6.9264e-05
34420A | Resistance | 100000 | 1000 | 10 | Ohm | mΩ |  |  |  | 6.6559 | 6.6056e-05 | 4.1389 | 6.7687e-05
34420A | Resistance | 100000 | 1000 | 100 | Ohm | mΩ |  |  |  | 64.3215 | 5.4751e-05 | 55.0192 | 3.7118e-05
34420A | Resistance | 100000 | 1000 | 1000 | Ohm | mΩ |  |  |  | 600.444 | 9.5076e-05 | 579.5927 | 4.809e-06
34420A | Resistance | 100000 | 1000 | 10000 | Ohm | Ω |  |  |  | 5.8513 | 2.4789e-07 | 5.8 | 4.83e-10
34420A | Resistance | 100000 | 1000 | 1000000 | Ohm | Ω |  |  |  | 5 | 8.05e-08 | 4.6334 | 8.08e-08
34420A | Resistance | 100000 | 1000 | 10000000 | Ohm | Ω |  |  |  | 8.0349 | 7.84e-08 | 5.9988 | 7.96e-08
34420A | Resistance | 100000 | 1000 | 100000000 | Ohm | Ω |  |  |  | 63.945 | 6.39e-08 | 54.4969 | 4.88e-08
34420A | Resistance | 100000 | 1000 | 1000000000 | Ohm | Ω |  |  |  | 600.3396 | 9.67e-08 | 579.4589 | 6.8e-09
34420A | Resistance | 100000 | 1000 | 10000000000 | Ohm | kΩ |  |  |  | 5.8513 | 2.481e-10 | 5.7999 | 6.9e-13
5502A | Resistance | 110000 | 330010 | 1 | Ω | Ω |  |  |  | 12 | 0.00014 | 11.565 | 0.0001393
5502A | Resistance | 110000 | 330010 | 10 | Ω | Ω |  |  |  | 12.289 | 0.00013784 | 12.284 | 0.00013782
5502A | Resistance | 110000 | 330010 | 100 | Ω | Ω |  |  |  | 55.026 | 8.0748e-05 | 55.024 | 8.0667e-05
5502A | Resistance | 110000 | 330010 | 1000 | Ω | Ω |  |  |  | 579.51 | 1.0142e-05 | 579.51 | 1.0091e-05
5502A | Frequency | 120000 | 1200010 | 10 | Hz | Hz |  |  |  | 60 | 1.6e-05 | 18.159 | 2.8435e-05
5502A | Frequency | 120000 | 1200010 | 100 | Hz | Hz |  |  |  | 60.876 | 1.8461e-05 | 59.55 | 1.5104e-05
5502A | Frequency | 120000 | 1200010 | 1000 | Hz | Hz |  |  |  | 578.01 | 3.8671e-06 | 577.51 | 1.8144e-06
5502A | Frequency | 120000 | 1200010 | 10000 | Hz | Hz |  |  |  | 5773.8 | 1.0097e-06 | 5773.6 | 1.8186e-07
Transmille | Insulation Res | 250000 | 100000000 | 10000 | Ω | kΩ |  |  |  | 18 | 9.2e-06 | 13.773 | 9.24e-06
Transmille | Insulation Res | 250000 | 100000000 | 100000 | Ω | kΩ |  |  |  | 57.523 | 8.82e-06 | 57.453 | 8.82e-06
Transmille | Insulation Res | 250000 | 100000000 | 1000000 | Ω | kΩ |  |  |  | 576.26 | 5.27e-06 | 576.24 | 5.27e-06
Transmille | Insulation Res | 250000 | 100000000 | 10000000 | Ω | kΩ |  |  |  | 5775 | 1.02e-06 | 5775 | 1.02e-06
5502A | Resistance | 330000 | 1100010 | 10 | Ω | Ω |  |  |  | 16 | 0.00017 | 12.341 | 0.00017274
5502A | Resistance | 330000 | 1100010 | 100 | Ω | Ω |  |  |  | 38.744 | 0.00015613 | 38.702 | 0.00015611
5502A | Resistance | 330000 | 1100010 | 1000 | Ω | Ω |  |  |  | 571.18 | 3.9241e-05 | 571.18 | 3.9165e-05
5502A | Resistance | 330000 | 1100010 | 10000 | Ω | Ω |  |  |  | 5799.1 | 4.0776e-06 | 5799.1 | 4.0446e-06
3458A | Frequency | 1000000 | 10000000 | 1 | Hz | Hz |  |  |  | 4.4e-06 | 1.3e-07 | 7.41e-07 | 1.2922e-07
3458A | Frequency | 1000000 | 10000000 | 10 | Hz | Hz |  |  |  | 0.00019518 | 1.292e-07 | 0.00013837 | 1.2921e-07
3458A | Frequency | 1000000 | 10000000 | 100 | Hz | Hz |  |  |  | 0.013574 | 1.2799e-07 | 0.013571 | 1.2799e-07
3458A | Frequency | 1000000 | 10000000 | 1000 | Hz | Hz |  |  |  | 0.68566 | 9.1517e-08 | 0.68565 | 9.1518e-08
3458A | Frequency | 1000000 | 10000000 | 10000 | Hz | Hz |  |  |  | 5.7844 | 1.571e-08 | 5.7844 | 1.571e-08
3458A | Frequency | 1000000 | 10000000 | 100000 | Hz | Hz |  |  |  | 57.734 | 1.59e-09 | 57.72 | 1.59e-09
3458A | Resistance | 1000000 | 10000000 | 1 | Ω | Ω |  |  |  | 120 | 5.9e-05 | 115.5 | 5.7739e-05
3458A | Resistance | 1000000 | 10000000 | 10 | Ω | Ω |  |  |  | 120 | 5.9e-05 | 115.61 | 5.7731e-05
3458A | Resistance | 1000000 | 10000000 | 100 | Ω | Ω |  |  |  | 124.34 | 5.8567e-05 | 125.65 | 5.6965e-05
3458A | Resistance | 1000000 | 10000000 | 1000 | Ω | Ω |  |  |  | 569.84 | 3.3313e-05 | 569.55 | 3.3236e-05
3458A | Resistance | 1000000 | 10000000 | 10000 | Ω | Ω |  |  |  | 5771.9 | 4.3708e-06 | 5771.8 | 4.3145e-06
3458A | Resistance | 1000000 | 10000000 | 100000 | Ω | Ω |  |  |  | 57735 | 4.5579e-07 | 57731 | 4.3308e-07
5502A | Resistance | 1100000 | 3301000 | 10 | Ω | kΩ |  |  |  | 0.19 | 1.7e-07 | 0.17339 | 1.7318e-07
5502A | Resistance | 1100000 | 3301000 | 100 | Ω | kΩ |  |  |  | 0.19 | 1.7215e-07 | 0.1791 | 1.7213e-07
5502A | Resistance | 1100000 | 3301000 | 1000 | Ω | kΩ |  |  |  | 0.55516 | 1.1804e-07 | 0.55501 | 1.1789e-07
5502A | Resistance | 1100000 | 3301000 | 10000 | Ω | kΩ |  |  |  | 5.7933 | 1.659e-08 | 5.7933 | 1.6469e-08
5502A | Frequency | 1200000 | 2000000 | 100 | Hz | kHz |  |  |  | 0.43 | 0.00013 | -0.35824 | 0.00036326
5502A | Frequency | 1200000 | 2000000 | 1000 | Hz | kHz |  |  |  | 0.43 | 0.00014137 | 0.42566 | 0.00012836
5502A | Frequency | 1200000 | 2000000 | 10000 | Hz | kHz |  |  |  | 5.7549 | 2.1951e-05 | 5.757 | 1.4015e-05
5502A | Frequency | 1200000 | 2000000 | 100000 | Hz | kHz |  |  |  | 57.733 | 4.5961e-06 | 57.734 | 1.4029e-06
5502A | Frequency | 2000000 | 2001000 | 100 | Hz | kHz |  |  |  | 0.43 | 0.00013 | -0.35824 | 0.00036326
5502A | Frequency | 2000000 | 2001000 | 1000 | Hz | kHz |  |  |  | 0.43 | 0.00014137 | 0.42566 | 0.00012836
5502A | Frequency | 2000000 | 2001000 | 10000 | Hz | kHz |  |  |  | 5.7549 | 2.1951e-05 | 5.757 | 1.4015e-05
5502A | Frequency | 2000000 | 2001000 | 100000 | Hz | kHz |  |  |  | 57.733 | 4.5961e-06 | 57.734 | 1.4029e-06
5502A | Resistance | 3300000 | 11001000 | 100 | Ω | kΩ |  |  |  | 0.33 | 6.9e-07 | 0.29033 | 6.9272e-07
5502A | Resistance | 3300000 | 11001000 | 1000 | Ω | kΩ |  |  |  | 0.37289 | 6.8715e-07 | 0.37252 | 6.8716e-07
5502A | Resistance | 3300000 | 11001000 | 10000 | Ω | kΩ |  |  |  | 4.8626 | 4.4972e-07 | 4.8623 | 4.4966e-07
5502A | Resistance | 3300000 | 11001000 | 100000 | Ω | kΩ |  |  |  | 57.852 | 6.2356e-08 | 57.852 | 6.2304e-08
3458A | Resistance | 10000000 | 100000000 | 10 | Ω | kΩ |  |  |  | 1.2 | 5.8e-07 | 1.1508 | 5.778e-07
3458A | Resistance | 10000000 | 100000000 | 100 | Ω | kΩ |  |  |  | 1.2 | 5.8e-07 | 1.1511 | 5.778e-07
3458A | Resistance | 10000000 | 100000000 | 1000 | Ω | kΩ |  |  |  | 1.2 | 5.8e-07 | 1.1777 | 5.7756e-07
3458A | Resistance | 10000000 | 100000000 | 10000 | Ω | kΩ |  |  |  | 3.4529 | 5.5768e-07 | 3.442 | 5.5771e-07
3458A | Resistance | 10000000 | 100000000 | 100000 | Ω | kΩ |  |  |  | 55.449 | 2.7074e-07 | 55.444 | 2.7056e-07
3458A | Resistance | 10000000 | 100000000 | 10000000 | Ω | kΩ |  |  |  | 5773.5 | 3.31e-09 | 5770.9 | 3.3e-09
5502A | Resistance | 11000000 | 33001000 | 100 | Ω | kΩ |  |  |  | 2.9 | 1.2e-06 | 2.8873 | 1.1548e-06
5502A | Resistance | 11000000 | 33001000 | 1000 | Ω | kΩ |  |  |  | 2.9004 | 1.1545e-06 | 2.9 | 1.1545e-06
5502A | Resistance | 11000000 | 33001000 | 10000 | Ω | kΩ |  |  |  | 4.2484 | 1.1259e-06 | 4.2478 | 1.1259e-06
5502A | Resistance | 11000000 | 33001000 | 100000 | Ω | kΩ |  |  |  | 54.575 | 4.9852e-07 | 54.575 | 4.9846e-07
5502A | Resistance | 33000000 | 110010000 | 1000 | Ω | kΩ |  |  |  | 3.5 | 5.8e-06 | 3.4663 | 5.7735e-06
5502A | Resistance | 33000000 | 110010000 | 10000 | Ω | kΩ |  |  |  | 3.5803 | 5.7727e-06 | 3.5792 | 5.7727e-06
5502A | Resistance | 33000000 | 110010000 | 100000 | Ω | kΩ |  |  |  | 14.463 | 5.6975e-06 | 14.465 | 5.6974e-06
5502A | Resistance | 33000000 | 110010000 | 1000000 | Ω | kΩ |  |  |  | 503.99 | 3.2606e-06 | 503.99 | 3.2605e-06
3458A | Resistance | 100000000 | 1000000000 | 100 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 11.592 | 5.7735e-06
3458A | Resistance | 100000000 | 1000000000 | 1000 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 11.592 | 5.7735e-06
3458A | Resistance | 100000000 | 1000000000 | 10000 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 11.595 | 5.7707e-06
3458A | Resistance | 100000000 | 1000000000 | 100000 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 14.68 | 5.7707e-06
3458A | Resistance | 100000000 | 1000000000 | 10000000 | Ω | kΩ |  |  |  | 5540.2 | 2.6331e-06 | 5540.2 | 2.633e-06
3458A | Resistance | 100000000 | 1000000000 | 100000000 | Ω | kΩ |  |  |  | 57707 | 3.1791e-07 | 57475 | 3.179e-07
Transmille | Insulation Res | 100000000 | 250000000 | 10000 | Ω | kΩ |  |  |  | 14 | 1.1e-05 | 13.511 | 1.08e-05
Transmille | Insulation Res | 100000000 | 250000000 | 100000 | Ω | kΩ |  |  |  | 57.138 | 1.06e-05 | 57.125 | 1.06e-05
Transmille | Insulation Res | 100000000 | 250000000 | 1000000 | Ω | kΩ |  |  |  | 575.4 | 8.79e-06 | 575.35 | 8.79e-06
Transmille | Insulation Res | 100000000 | 250000000 | 10000000 | Ω | kΩ |  |  |  | 5773 | 2.52e-06 | 5772.9 | 2.52e-06
Transmille | Insulation Res | 100000000 | 250000000 | 100000000 | Ω | kΩ |  |  |  | 57896 | 1.31e-06 | 57896 | 1.31e-06
5502A | Resistance | 110000000 | 330010000 | 1000 | Ω | kΩ |  |  |  | 120 | 5.8e-06 | 115.47 | 5.7735e-06
5502A | Resistance | 110000000 | 330010000 | 10000 | Ω | kΩ |  |  |  | 120 | 5.7734e-06 | 115.52 | 5.7734e-06
5502A | Resistance | 110000000 | 330010000 | 100000 | Ω | kΩ |  |  |  | 120 | 5.7671e-06 | 118.42 | 5.7671e-06
5502A | Resistance | 110000000 | 330010000 | 1000000 | Ω | kΩ |  |  |  | 371.66 | 5.2445e-06 | 371.67 | 5.2444e-06
Transmille | Insulation Res | 250000000 | 500000000 | 10000 | Ω | kΩ |  |  |  | 18 | 1.1e-05 | 12.764 | 1.1e-05
Transmille | Insulation Res | 250000000 | 500000000 | 100000 | Ω | kΩ |  |  |  | 55.079 | 1.09e-05 | 54.77 | 1.09e-05
Transmille | Insulation Res | 250000000 | 500000000 | 1000000 | Ω | kΩ |  |  |  | 572.67 | 9.95e-06 | 572.64 | 9.95e-06
Transmille | Insulation Res | 250000000 | 500000000 | 10000000 | Ω | kΩ |  |  |  | 5771.4 | 4.47e-06 | 5771.3 | 4.47e-06
5502A | Resistance | 330000000 | 1100010000 | 10000 | Ω | kΩ |  |  |  | 930 | 1.7e-05 | 577.19 | 1.7321e-05
5502A | Resistance | 330000000 | 1100010000 | 100000 | Ω | kΩ |  |  |  | 930 | 1.732e-05 | 578 | 1.732e-05
5502A | Resistance | 330000000 | 1100010000 | 1000000 | Ω | kΩ |  |  |  | 930 | 1.7297e-05 | 611.79 | 1.7297e-05
5502A | Resistance | 330000000 | 1100010000 | 10000000 | Ω | kΩ |  |  |  | 3454.2 | 1.5468e-05 | 3454.1 | 1.5468e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 10000 | Ω | kΩ |  |  |  | 25 | 1.1e-05 | -76.51 | 1.12e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 100000 | Ω | kΩ |  |  |  | 25 | 1.12e-05 | -75.391 | 1.12e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 1000000 | Ω | kΩ |  |  |  | 25 | 1.1e-05 | 24.396 | 1.1e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 10000000 | Ω | kΩ |  |  |  | 5114.5 | 7.34e-06 | 5114.5 | 7.34e-06
Transmille | Insulation Res | 500000000 | 1000000000 | 100000000 | Ω | kΩ |  |  |  | 57638 | 1.44e-06 | 57638 | 1.44e-06
5502A | Resistance | 1100000000 | 1101000000 | 10000 | Ω | kΩ |  |  |  | 930 | 1.7e-05 | 577.19 | 1.7321e-05
5502A | Resistance | 1100000000 | 1101000000 | 100000 | Ω | kΩ |  |  |  | 930 | 1.732e-05 | 578 | 1.732e-05
5502A | Resistance | 1100000000 | 1101000000 | 1000000 | Ω | kΩ |  |  |  | 930 | 1.7297e-05 | 611.79 | 1.7297e-05
5502A | Resistance | 1100000000 | 1101000000 | 10000000 | Ω | kΩ |  |  |  | 3454.2 | 1.5468e-05 | 3454.1 | 1.5468e-05

## Worksheet
|  | Johnson Gage and Inspection, Inc. |  | Page 2 of 2 | [16 empty] | INSPECTOR NOTE: PRINT PAGE 1 ONLY | [6 empty] | µF | µΩ | °C
| Supplementary Data Sheet to Report # |  |  | 22007 | [16 empty] | This section is for Creation of Datasheet Only | [6 empty] | µV | @ | °F
| Fluke 179 |  | Date: | 2022-01-04 | [13 empty] | Secondary Unit for Range field |  |  |  |  | Note: You can now use "O" instead of Ω, "DGC" instead |  |  |  |  | µA |  |  | RESULT | Special Character Filter
[11 empty] | REPEAT READINGS | [11 empty] | of "°C", and "DGF" instead of "°F" if you want to. | [7 empty] | CONV. | [8 empty] | 89.74%
Function | Range | Signal Applied | Result | Min | Max | Uncertainty @k=2 | Hi Unc | Coil Used | Meter Used | Notes: | 1 | 2 | 3 | 4 | 5 | Special | Wire Type |  | Resolution | Range | Primary Unit | DC/AC | Nominal | tolerance | max | min | (leave blank) | Frequency | Secondary Unit |  | TO NUMBER | [7 empty] | Expected Unc | Unc. Is Correct | Found row | Correct Lookup Row
Volts AC | 600 mV | 300 mVAC @ 45Hz | 300.2 | 297.1mVAC | 303.1mVAC | 140 µV |  |  | 5502A | [9 empty] | 0.1 | 600 | mV | AC | 300 | 3.3 | 303.1 | 297.1 |  | 45 | Hz | @ 45Hz | 300.2 | mV | 300.2 | 303.1 | 297.1 | 1 | 1 | 1 | 140 µV | TRUE | 553 | 553
| 6 V | 5 VAC @ 500Hz | 4.998 | 4.951VAC | 5.051VAC | 2.5 mV |  |  | 5502A | [9 empty] | 0.001 | 6 | V | AC | 5 | 0.053 | 5.051 | 4.951 |  | 500 | Hz | @ 500Hz | 4.998 | V | 4.998 | 5.051 | 4.951 | 1 | 1 | 1 | 2.5 mV | TRUE | 622 | 622
| 6 V | 5 VAC @ 1kHz | 4.959 | 4.901VAC | 5.101VAC | 2.5 mV |  |  | 5502A | [9 empty] | 0.001 | 6 | V | AC | 5 | 0.103 | 5.101 | 4.901 |  | 1 | kHz | @ 1kHz | 4.959 | V | 4.959 | 5.101 | 4.901 | 1 | 1 | 1 | 2.5 mV | TRUE | 622 | 622
| 60 V | 50 VAC @ 45Hz | 49.98 | 49.51VAC | 50.51VAC | 58 mV |  |  | 5502A | [9 empty] | 0.01 | 60 | V | AC | 50 | 0.53 | 50.51 | 49.51 |  | 45 | Hz | @ 45Hz | 49.98 | V | 49.98 | 50.51 | 49.51 | 1 | 1 | 1 | 58 mV | TRUE | 657 | 657
| 60 V | 50 VAC @ 1kHz | 50.12 | 49.01VAC | 51.01VAC | 58 mV |  |  | 5502A | [9 empty] | 0.01 | 60 | V | AC | 50 | 1.03 | 51.01 | 49.01 |  | 1 | kHz | @ 1kHz | 50.12 | V | 50.12 | 51.01 | 49.01 | 1 | 1 | 1 | 58 mV | TRUE | 657 | 657
| 600 V | 300 VAC @ 45Hz | 300.2 | 297.1VAC | 303.1VAC | 300 mV |  |  | 5502A | [9 empty] | 0.1 | 600 | V | AC | 300 | 3.3 | 303.1 | 297.1 |  | 45 | Hz | @ 45Hz | 300.2 | V | 300.2 | 303.1 | 297.1 | 1 | 1 | 1 | 300 mV | TRUE | 658 | 658
| 600 V | 500 VAC @ 500Hz | 501.0 | 495.1VAC | 505.1VAC | 670 mV |  |  | 5502A | [9 empty] | 0.1 | 600 | V | AC | 500 | 5.3 | 505.1 | 495.1 |  | 500 | Hz | @ 500Hz | 501 | V | 501 | 505.1 | 495.1 | 1 | 1 | 1 | 670 mV | TRUE | 683 | 683
| 600 V | 500 VAC @ 1kHz | 501.0 | 490.1VAC | 510.1VAC | 670 mV |  |  | 5502A | [9 empty] | 0.1 | 600 | V | AC | 500 | 10.3 | 510.1 | 490.1 |  | 1 | kHz | @ 1kHz | 501 | V | 501 | 510.1 | 490.1 | 1 | 1 | 1 | 780 mV | FALSE | 688 | 683
| 1000 V | 1000 VAC @ 45Hz | 1002 | 1VAC | 1VAC | 5800 mV |  |  | 5502A | [9 empty] | 1 | 1000 | V | AC | 1000 | 13 | 1 | 1 |  | 45 | Hz | @ 45Hz | 1002 | V | 1002 | 1 | 1 | 0 | 1 | 0 | 5800 mV | TRUE | 684 | 684
Frequency | 99.99 Hz | 1V @ 45Hz | 45.00 | 44.91Hz | 45.11Hz | 0.0068 Hz |  |  | 5502A | [9 empty] | 0.01 | 99.99 | Hz |  | 45 | 0.06 | 45.11 | 44.91 |  | 1 | V | 1V | 45 | Hz | 45 | 45.11 | 44.91 | 1 | 1 | 1 | 0.0068 Hz | TRUE | 1267 | 1267
| 99.99 kHz | 5V @ 50kHz | 50.00 | 49.91kHz | 50.11kHz | 20 Hz |  |  | 5502A | [9 empty] | 0.01 | 99.99 | kHz |  | 50 | 0.06 | 50.11 | 49.91 |  | 5 | V | 5V | 50 | kHz | 50 | 50.11 | 49.91 | 1 | 1 | 1 | 20 Hz | TRUE | 1282 | 1282
Volts DC | 6 V | 5 VDC | 5.000 | 4.991VDC | 5.011VDC | 0.72 mV |  |  | 5502A | [9 empty] | 0.001 | 6 | V | DC | 5 | 0.007 | 5.011 | 4.991 |  |  |  |  | 5 | V | 5 | 5.011 | 4.991 | 1 | 1 | 1 | 0.72 mV | TRUE | 1075 | 1075
| 600 V | 300 VDC | 300.0 | 300.1VDC | 301.1VDC | 62 mV |  |  | 5502A | [9 empty] | 0.1 | 600 | V | DC | 300 | 0.5 | 301.1 | 300.1 |  |  |  |  | 300 | V | 300 | 301.1 | 300.1 | 0 | 0 | 1 | 62 mV | TRUE | 1082 | 1082
| 1000 V | 1000 VDC | 1000 | 1VDC | 1VDC | 580 mV |  |  | 5502A | [9 empty] | 1 | 1000 | V | DC | 1000 | 3 | 1 | 1 |  |  |  |  | 1000 | V | 1000 | 1 | 1 | 0 | 1 | 0 | 580 mV | TRUE | 1093 | 1093
| 1000 V | -1000 VDC | -1000 | -1VDC | -1VDC | 580 mV |  |  | 5502A | [9 empty] | 1 | 1000 | V | DC | -1000 | 3 | -1 | -1 |  |  |  |  | -1000 | V | -1000 | -1 | -1 | 0 | 0 | 1 | 580 mV | TRUE | 1093 | 1093
Volts DC Frequency | 99.99 Hz | 3V @ 45Hz | 45.00 | 44.91Hz | 45.11Hz | 0.0068 Hz |  |  | 5502A | [9 empty] | 0.01 | 99.99 | Hz |  | 45 | 0.06 | 45.11 | 44.91 |  | 3 | V | 3V | 45 | Hz | 45 | 45.11 | 44.91 | 1 | 1 | 1 | 0.0068 Hz | TRUE | 1267 | 1267
| 99.99 kHz | 30V @ 50kHz | 50.00 | 49.91kHz | 50.11kHz | 20 Hz |  |  | 5502A | [9 empty] | 0.01 | 99.99 | kHz |  | 50 | 0.06 | 50.11 | 49.91 |  | 30 | V | 30V | 50 | kHz | 50 | 50.11 | 49.91 | 1 | 1 | 1 | 20 Hz | TRUE | 1282 | 1282
Millivolts DC | 600 mV | 30 mVDC | 30.0 | 30.1mVDC | 30.1mVDC | 59 µV |  |  | 5502A | [9 empty] | 0.1 | 600 | mV | DC | 30 | 0.2 | 30.1 | 30.1 |  |  |  |  | 30 | mV | 30 | 30.1 | 30.1 | 0 | 0 | 1 | 59 µV | TRUE | 1061 | 1061
| 600 mV | -300 mVDC | -299.9 | -301.1mVDC | -300.1mVDC | 63 µV |  |  | 5502A | [9 empty] | 0.1 | 600 | mV | DC | -300 | 0.5 | -300.1 | -301.1 |  |  |  |  | -299.9 | mV | -299.9 | -300.1 | -301.1 | 0 | 1 | 0 | 63 µV | TRUE | 1061 | 1061
| 600 mV | 600 mVDC | 599.9 | 599.1mVDC | 601.1mVDC | 76 µV |  |  | 5502A | [9 empty] | 0.1 | 600 | mV | DC | 600 | 0.7 | 601.1 | 599.1 |  |  |  |  | 599.9 | mV | 599.9 | 601.1 | 599.1 | 1 | 1 | 1 | 76 µV | TRUE | 1069 | 1069
Resistance | 600 Ω | 500 Ω | 500.2 | 495.1Ω | 505.1Ω | 97 mΩ |  |  | 5502A | [9 empty] | 0.1 | 600 | Ω |  | 500 | 4.7 | 505.1 | 495.1 |  |  |  |  | 500.2 | Ω | 500.2 | 505.1 | 495.1 | 1 | 1 | 1 | 97 mΩ | TRUE | 1321 | 1321
| 50 MΩ | 19 MΩ | 19.00 | 18.71MΩ | 19.31MΩ | 26 kΩ |  |  | 5502A | [9 empty] | 0.01 | 50 | MΩ |  | 19 | 0.32 | 19.31 | 18.71 |  |  |  |  | 19 | MΩ | 19 | 19.31 | 18.71 | 1 | 1 | 1 | 26 kΩ | TRUE | 1358 | 1358
Capacitance | 1000 nF | 900 nF | 900 | 1nF | 1nF | 3.8 nF |  |  | 5502A | [9 empty] | 1 | 1000 | nF |  | 900 | 13 | 1 | 1 |  |  |  |  | 900 | nF | 900 | 1 | 1 | 0 | 1 | 0 | 3.8 nF | TRUE | 793 | 793
Continuity | 600 Ω | 25 Ω |  | 25.1Ω | 25.1Ω |  |  |  | 5502A | [9 empty] | 0.1 | 600 | Ω |  | 25 |  | 25.1 | 25.1 |  |  |  |  | 0 | Ω | 0 | 25.1 | 25.1 | 0 | 0 | 1 | #N/A | #N/A
| 600 Ω | 250 Ω |  | 250.1Ω | 250.1Ω |  |  |  | 5502A | [9 empty] | 0.1 | 600 | Ω |  | 250 |  | 250.1 | 250.1 |  |  |  |  | 0 | Ω | 0 | 250.1 | 250.1 | 0 | 0 | 1 | #N/A | #N/A
Diode | 2.400 V | 2.0 VDC | 1.999 | 1.981VDC | 2.021VDC | 600 µV |  |  | 5502A | [9 empty] | 0.001 | 2.400 | V | DC | 2.0 | 0.022 | 2.021 | 1.981 |  |  |  |  | 1.999 | V | 1.999 | 2.021 | 1.981 | 1 | 1 | 1 | 600 µV | TRUE | 1070 | 1070
Amps AC | 60 mA | 3 mAAC @ 45Hz | 3.00 | 2.91mAAC | 3.11mAAC | 6.9 µA |  |  | 5502A | [9 empty] | 0.01 | 60 | mA | AC | 3 | 0.08 | 3.11 | 2.91 |  | 45 | Hz | @ 45Hz | 3 | mA | 3 | 3.11 | 2.91 | 1 | 1 | 1 | 6.9 µA | TRUE | 14 | 14
| 60 mA | 50 mAAC @ 1kHz | 50.02 | 49.21mAAC | 50.81mAAC | 0.12 mA |  |  | 5502A | [9 empty] | 0.01 | 60 | mA | AC | 50 | 0.78 | 50.81 | 49.21 |  | 1 | kHz | @ 1kHz | 50.02 | mA | 50.02 | 50.81 | 49.21 | 1 | 1 | 1 | 0.12 mA | TRUE | 157 | 157
| 400 mA | 400 mAAC @ 1kHz | 399.3 | 394.1mAAC | 406.1mAAC | 4 mA |  |  | 5502A | [9 empty] | 0.1 | 400 | mA | AC | 400 | 6.3 | 406.1 | 394.1 |  | 1 | kHz | @ 1kHz | 399.3 | mA | 399.3 | 406.1 | 394.1 | 1 | 1 | 1 | 4 mA | TRUE | 187 | 187
| 6 A | 4 AAC @ 45Hz | 3.999 | 3.941AAC | 4.061AAC | 5.2 mA |  |  | 5502A | [9 empty] | 0.001 | 6 | A | AC | 4 | 0.063 | 4.061 | 3.941 |  | 45 | Hz | @ 45Hz | 3.999 | A | 3.999 | 4.061 | 3.941 | 1 | 1 | 1 | 5.2 mA | TRUE | 227 | 227
| 10 A | 9 AAC @ 1kHz | 9.01 | 8.81AAC | 9.21AAC | 14 mA | [12 empty] | 0.01 | 10 | A | AC | 9 | 0.17 | 9.21 | 8.81 |  | 1 | kHz | @ 1kHz | 9.01 | A | 9.01 | 9.21 | 8.81 | 1 | 1 | 1 | 320 mA | FALSE | 238
Amps DC | 60 mA | 3 mADC | 3.00 | 2.91mADC | 3.11mADC | 5.8 µA |  |  | 5502A | [9 empty] | 0.01 | 60 | mA | DC | 3 | 0.06 | 3.11 | 2.91 |  |  |  |  | 3 | mA | 3 | 3.11 | 2.91 | 1 | 1 | 1 | 5.8 µA | TRUE | 958 | 958
| 60 mA | 50 mADC | 49.98 | 49.51mADC | 50.51mADC | 11 µA |  |  | 5502A | [9 empty] | 0.01 | 60 | mA | DC | 50 | 0.53 | 50.51 | 49.51 |  |  |  |  | 49.98 | mA | 49.98 | 50.51 | 49.51 | 1 | 1 | 1 | 11 µA | TRUE | 966 | 966
| 400 mA | -400 mADC | -399.9 | -404.1mADC | -396.1mADC | 0.23 mA |  |  | 5502A | [9 empty] | 0.1 | 400 | mA | DC | -400 | 4.3 | -396.1 | -404.1 |  |  |  |  | -399.9 | mA | -399.9 | -396.1 | -404.1 | 1 | 1 | 1 | 0.23 mA | TRUE | 973 | 973
| 6 A | 4 ADC | 3.999 | 3.961ADC | 4.041ADC | 3.7 mA |  |  | 5502A | [9 empty] | 0.001 | 6 | A | DC | 4 | 0.043 | 4.041 | 3.961 |  |  |  |  | 3.999 | A | 3.999 | 4.041 | 3.961 | 1 | 1 | 1 | 3.7 mA | TRUE | 978 | 978
| 10 A | -9 ADC | -9.00 | -9.11ADC | -8.91ADC | 10 mA |  |  | 5502A | [9 empty] | 0.01 | 10 | A | DC | -9 | 0.12 | -8.91 | -9.11 |  |  |  |  | -9 | A | -9 | -8.91 | -9.11 | 1 | 1 | 1 | 10 mA | TRUE | 979 | 979
Temperature | 400 °C | 0 °C | 0.2 | -1.1°C | 1.1°C | 0.2 °C |  |  | 5502A | [7 empty] | K |  | 0.1 | 400 | °C |  | 0 | 1 | 1.1 | -1.1 |  |  |  |  | 0.2 | °C | 0.2 | 1.1 | -1.1 | 1 | 1 | 1 | 0.2 °C | TRUE | 866 | 866
| 400 °C | -40 °C | -39.8 | -41.1°C | -39.1°C | 0.23 °C |  |  | 5502A | [7 empty] | K |  | 0.1 | 400 | °C |  | -40 | 1.4 | -39.1 | -41.1 |  |  |  |  | -39.8 | °C | -39.8 | -39.1 | -41.1 | 1 | 1 | 1 | 0.23 °C | TRUE | 865 | 865
| 400 °C | 400 °C | 399.7 | 395.1°C | 405.1°C | 0.31 °C |  |  | 5502A | [7 empty] | K |  | 0.1 | 400 | °C |  | 400 | 5 | 405.1 | 395.1 |  |  |  |  | 399.7 | °C | 399.7 | 405.1 | 395.1 | 1 | 1 | 1 | 0.31 °C | TRUE | 867 | 867
Temperature |  | Temperature | 72 | [21 empty] | 0 | 0 |  |  |  |  | 72 | 0 | 72 | 0 | 0 | 0 | 1 | 0 | #N/A
Conditions |  | Humidity | 22% | [21 empty] | 0 | 0 |  |  |  |  | 0.22 | 0 | 0.22 | 0 | 0 | 0 | 1 | 0 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 |  |  |  | 3458A | [15 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
|  | 0 |  | 00 | 00 | [19 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | #N/A
Calculate Uncertainty

## Unc. Lookup
Standard | Item | Range1_Low | Range1_High | Res | Meas_Units | Uncert_Units | Range2_Low | Range2_High | Range2_Units | Base | Mult | Base2 | Mult2
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 0.23 | 2300 | 0.17326 | 2315.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 0.23 | 2300 | 0.17481 | 2314.9
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 0.35294 | 2267.3 | 0.35227 | 2267.5
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 5.4199 | 1301.8 | 5.4195 | 1301.8
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 57.686 | 174.69 | 57.685 | 174.66
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 0.32 | 1400 | 0.17327 | 1443.8
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 0.32 | 1400 | 0.17604 | 1443.1
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 0.41399 | 1381.1 | 0.41349 | 1381.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 5.6118 | 601.75 | 5.6114 | 601.72
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 57.716 | 69.766 | 57.716 | 69.735
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17375 | 1155.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17651 | 1154.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.43013 | 1124.2 | 0.44201 | 1086.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.6654 | 409.29 | 5.6651 | 409.25
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.723 | 45.397 | 57.723 | 45.366
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 0.27 | 2300 | 0.23112 | 2309.7
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 0.27 | 2300 | 0.2328 | 2309.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 0.40201 | 2264.3 | 0.40156 | 2264.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 5.4266 | 1309.3 | 5.4263 | 1309.3
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 57.686 | 176.15 | 57.686 | 176.12
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 0.35 | 5800 | 0.34654 | 5773.6
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 0.35 | 5800 | 0.34724 | 5773.4
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 0.41769 | 5779.5 | 0.4264 | 5752
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 4.6373 | 4728.3 | 4.6369 | 4728.3
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 57.432 | 1053.3 | 57.432 | 1053.2
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69332 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69339 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.7339 | 11536
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 10 | 30 | kHz | 3.6592 | 11103 | 3.7773 | 10742
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 10 | 30 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
5502A | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.69332 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.69339 | 11547
5502A | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.7339 | 11536
5502A | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 30 | 30.1 | kHz | 3.6592 | 11103 | 3.7773 | 10742
5502A | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 30 | 30.1 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 2.4 | 2100 | 2.3005 | 2087.6
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 2.4 | 2100 | 2.3157 | 2087.1
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 4.0897 | 2048.8 | 4.1185 | 2039.5
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 54.774 | 1117.8 | 54.773 | 1117.7
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 576.95 | 144.81 | 576.95 | 144.75
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 3.8 | 1000 | 2.2968 | 1043.6
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 3.8 | 1000 | 2.3282 | 1042.8
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 4.917 | 977.92 | 4.9145 | 977.93
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 56.865 | 350.66 | 56.863 | 350.55
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 577.26 | 38.407 | 577.26 | 38.349
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 2.4 | 470 | 2.2998 | 465.62
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 2.4 | 470 | 2.3464 | 464.49
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 5.6396 | 392.41 | 5.6369 | 392.36
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 57.584 | 84.81 | 57.583 | 84.672
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 577.34 | 8.7266 | 577.34 | 8.6681
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 2.4 | 930 | 2.3048 | 925.75
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 2.4 | 930 | 2.3386 | 924.88
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 5.0491 | 857.92 | 5.0466 | 857.92
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 57.047 | 284.47 | 57.045 | 284.35
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 577.28 | 30.67 | 577.28 | 30.612
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 3.8 | 2300 | 3.4644 | 2309.4
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 3.8 | 2300 | 3.4807 | 2309
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 5.0186 | 2268.7 | 5.0129 | 2268.8
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 54.397 | 1333.5 | 54.39 | 1333.5
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 576.87 | 180.75 | 576.87 | 180.65
5502A | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
5502A | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
5502A | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 10 | 30 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
5502A | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 10 | 30 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 10 | 30 | kHz | 575.41 | 695.3 | 575.4 | 695.06
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
5502A | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 575.41 | 695.3 | 575.4 | 695.06
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 0.13 | 2300 | 0.11565 | 2314.2
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 0.13 | 2300 | 0.12522 | 2290.9
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 0.56558 | 1474.9 | 0.56533 | 1474.5
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 5.7994 | 124.28 | 5.7993 | 123.73
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 58.025 | -857.45 | 58.025 | -857.67
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.01 | 0.02001 | kHz | 580.26 | -8787.4 | 580.03 | -8787.5
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.01 | 0.02001 | kHz | 5802.6 | -87896 | 5800 | -87896
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.01 | 0.02001 | kHz | 58026 | -878960 | 58000 | -878960
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 0.01 | 0.02001 | kHz | 580260 | -8789600 | 580000 | -8789600
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 0.01 | 0.02001 | kHz | 5802600 | -87896000 | 5800000 | -87896000
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 0.13 | 1700 | 0.11582 | 1731.6
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 0.13 | 1700 | 0.1262 | 1707.4
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 0.57237 | 986.73 | 0.57213 | 986.13
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 5.7723 | 128.02 | 5.7722 | 127.48
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 57.735 | 13.015 | 57.735 | 12.795
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.02 | 0.04501 | kHz | 577.36 | 1.3674 | 577.35 | 1.2795
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.02 | 0.04501 | kHz | 5773.6 | 0.16308 | 5773.6 | 0.12796
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.02 | 0.04501 | kHz | 57736 | 0.026847 | 57736 | 0.012796
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 0.02 | 0.04501 | kHz | 577360 | 0.0068998 | 577360 | 0.0012794
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 0.02 | 0.04501 | kHz | 5773600 | 0.0023794 | 5773600 | 0.00012995
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.14 | 1400 | 0.11585 | 1442.9
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.14 | 1400 | 0.12669 | 1418.5
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.5767 | 759.64 | 0.57647 | 758.92
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.7731 | 93.964 | 5.773 | 93.42
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.736 | 9.5875 | 57.735 | 9.3679
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.045 | 1 | kHz | 577.36 | 1.0246 | 577.36 | 0.93682
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.045 | 1 | kHz | 5773.6 | 0.12881 | 5773.6 | 0.093682
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.045 | 1 | kHz | 57736 | 0.02342 | 57736 | 0.0093682
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 0.045 | 1 | kHz | 577360 | 0.0065572 | 577360 | 0.00093674
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 0.045 | 1 | kHz | 5773600 | 0.0023422 | 5773600 | 9.5917e-05
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 0.18 | 3500 | 0.17344 | 3463.7
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 0.18 | 3500 | 0.17984 | 3448.1
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 0.56231 | 2652.9 | 0.56207 | 2652.8
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 5.7665 | 471.27 | 5.7664 | 470.74
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 57.735 | 47.918 | 57.735 | 47.699
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 1 | 5.01 | kHz | 577.36 | 4.8583 | 577.35 | 4.7705
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 1 | 5.01 | kHz | 5773.6 | 0.51218 | 5773.6 | 0.47705
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 1 | 5.01 | kHz | 57736 | 0.061756 | 57736 | 0.047705
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 1 | 5.01 | kHz | 577360 | 0.010391 | 577360 | 0.0047703
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 1 | 5.01 | kHz | 5773600 | 0.0027259 | 5773600 | 0.00047649
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 0.25 | 9200 | 0.23106 | 9237.4
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 0.25 | 9200 | 0.23463 | 9228.1
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 0.5163 | 8526.7 | 0.51576 | 8527.5
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 5.7139 | 2807.7 | 5.7137 | 2806.8
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 57.729 | 302.48 | 57.729 | 302.01
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 5 | 10.01 | kHz | 577.35 | 30.413 | 577.35 | 30.226
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 5 | 10.01 | kHz | 5773.6 | 3.0976 | 5773.5 | 3.0226
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 5 | 10.01 | kHz | 57736 | 0.33224 | 57736 | 0.30226
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 5 | 10.01 | kHz | 577360 | 0.04222 | 577360 | 0.030226
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 5 | 10.01 | kHz | 5773600 | 0.0078188 | 5773600 | 0.0030229
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46198 | 18475
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46389 | 18470
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 10 | 30 | kHz | 0.6298 | 18044 | 0.62958 | 18044
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 10 | 30 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 10 | 30 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 10 | 30 | kHz | 577.35 | 121.09 | 577.32 | 120.9
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 10 | 30 | kHz | 5773.6 | 12.165 | 5773.5 | 12.09
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 10 | 30 | kHz | 57736 | 1.239 | 57735 | 1.209
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 10 | 30 | kHz | 577360 | 0.1329 | 577360 | 0.1209
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 10 | 30 | kHz | 5773600 | 0.016888 | 5773600 | 0.012089
5502A | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 30 | 30.1 | kHz | 0.62 | 18000 | 0.46198 | 18475
5502A | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 30 | 30.1 | kHz | 0.62 | 18000 | 0.46389 | 18470
5502A | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 30 | 30.1 | kHz | 0.6298 | 18044 | 0.62958 | 18044
5502A | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 30 | 30.1 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
5502A | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 30 | 30.1 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
5502A | ACA | 0 | 0.00033001 | 0.001 | A | µA | 30 | 30.1 | kHz | 577.35 | 121.09 | 577.32 | 120.9
5502A | ACA | 0 | 0.00033001 | 0.01 | A | µA | 30 | 30.1 | kHz | 5773.6 | 12.165 | 5773.5 | 12.09
5502A | ACA | 0 | 0.00033001 | 0.1 | A | µA | 30 | 30.1 | kHz | 57736 | 1.239 | 57735 | 1.209
5502A | ACA | 0 | 0.00033001 | 1 | A | µA | 30 | 30.1 | kHz | 577360 | 0.1329 | 577360 | 0.1209
5502A | ACA | 0 | 0.00033001 | 10 | A | µA | 30 | 30.1 | kHz | 5773600 | 0.016888 | 5773600 | 0.012089
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.01 | 0.02001 | kHz | 0.023 | 2.1 | 0.022917 | 2.0907
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.01 | 0.02001 | kHz | 0.023 | 2.1 | 0.023129 | 2.0901
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.01 | 0.02001 | kHz | 0.041052 | 2.0453 | 0.041148 | 2.0425
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.01 | 0.02001 | kHz | 0.54767 | 1.1202 | 0.54766 | 1.1202
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.01 | 0.02001 | kHz | 5.7695 | 0.14517 | 5.7695 | 0.14515
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.02 | 0.04501 | kHz | 0.038 | 1 | 0.023094 | 1.0436
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.02 | 0.04501 | kHz | 0.038 | 1 | 0.023358 | 1.0428
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.02 | 0.04501 | kHz | 0.049203 | 0.97804 | 0.049192 | 0.97804
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.02 | 0.04501 | kHz | 0.56864 | 0.35073 | 0.56864 | 0.35068
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.02 | 0.04501 | kHz | 5.7726 | 0.03839 | 5.7726 | 0.038365
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.02282 | 0.47204
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.023285 | 0.47091
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.056197 | 0.39853 | 0.056185 | 0.39851
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.57578 | 0.086645 | 0.57577 | 0.086584
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7734 | 0.0088941 | 5.7734 | 0.0088682
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 1 | 5.01 | kHz | 0.058 | 1.2 | 0.057624 | 1.1588
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 1 | 5.01 | kHz | 0.058 | 1.2 | 0.057825 | 1.1583
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 1 | 5.01 | kHz | 0.073934 | 1.1517 | 0.075029 | 1.1175
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 1 | 5.01 | kHz | 0.56965 | 0.47368 | 0.56963 | 0.47361
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 1 | 5.01 | kHz | 5.7726 | 0.053742 | 5.7726 | 0.053697
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 5 | 10.01 | kHz | 0.12 | 2.3 | 0.11543 | 2.3114
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 5 | 10.01 | kHz | 0.12 | 2.3 | 0.11552 | 2.3112
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 5 | 10.01 | kHz | 0.12471 | 2.2891 | 0.12466 | 2.2892
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 5 | 10.01 | kHz | 0.55922 | 1.4904 | 0.55916 | 1.4904
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 5 | 10.01 | kHz | 5.7697 | 0.21298 | 5.7697 | 0.21288
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 10 | 30 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 10 | 30 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
5502A | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
5502A | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
5502A | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
5502A | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 30 | 30.01 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
5502A | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 30 | 30.01 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11617 | 2.0784
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11882 | 2.0766
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.35291 | 1.9255 | 0.35284 | 1.9254
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 5.6471 | 0.55145 | 5.6471 | 0.55121
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 57.722 | 0.057755 | 57.722 | 0.057649
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.12 | 0.58 | 0.11734 | 0.57739
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.12267 | 0.57758 | 0.124 | 0.57333
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.52825 | 0.38235 | 0.52823 | 0.382
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7643 | 0.053023 | 5.7643 | 0.05276
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.735 | 0.005408 | 57.735 | 0.0053018
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 1 | 5.01 | kHz | 1.2 | 6.9 | 1.1549 | 6.9282
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 1 | 5.01 | kHz | 1.2 | 6.9 | 1.1555 | 6.9278
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 1 | 5.01 | kHz | 1.2164 | 6.8903 | 1.2155 | 6.8904
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 1 | 5.01 | kHz | 5.1011 | 4.9149 | 5.0998 | 4.9136
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 1 | 5.01 | kHz | 57.598 | 0.72965 | 57.598 | 0.72819
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7734 | 28.868
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7742 | 28.867
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7874 | 28.859
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 5 | 10 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 5 | 10 | kHz | 55.816 | 11.862 | 55.815 | 11.86
5502A | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7734 | 28.868
5502A | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7742 | 28.867
5502A | ACA | 0.33 | 1.1 | 0.001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7874 | 28.859
5502A | ACA | 0.33 | 1.1 | 0.01 | A | mA | 10 | 10.01 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
5502A | ACA | 0.33 | 1.1 | 0.1 | A | mA | 10 | 10.01 | kHz | 55.816 | 11.862 | 55.815 | 11.86
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11589 | 2.0787
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11676 | 2.0785
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.18892 | 2.077 | 0.20868 | 2.0565
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 4.9071 | 1.2281 | 4.9043 | 1.2265
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 57.613 | 0.15838 | 57.613 | 0.15706
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11619 | 0.69369
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11868 | 0.6931
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.35192 | 0.64234 | 0.34551 | 0.6421
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.6463 | 0.18057 | 5.6448 | 0.17758
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.722 | 0.019812 | 57.722 | 0.018474
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1548 | 6.9283
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1552 | 6.9282
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1804 | 6.9223
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 3.4547 | 6.4116 | 3.4521 | 6.4115
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 56.451 | 1.7722 | 56.451 | 1.771
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.773 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7731 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.78 | 28.866
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 5 | 10 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 5 | 10 | kHz | 45.662 | 21.093 | 45.659 | 21.092
5502A | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.773 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7731 | 28.868
5502A | ACA | 1.1 | 3.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.78 | 28.866
5502A | ACA | 1.1 | 3.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
5502A | ACA | 1.1 | 3.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 45.662 | 21.093 | 45.659 | 21.092
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 0.045 | 0.10001 | kHz | 2.4 | 0.69 | 2.3298 | 0.69373
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 0.045 | 0.10001 | kHz | 2.4 | 0.69 | 2.3709 | 0.69112
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 5.6644 | 0.53041 | 5.6622 | 0.53026
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 57.647 | 0.085551 | 57.646 | 0.085343
5502A | ACA | 3 | 11.001 | 1 | A | mA | 0.045 | 0.10001 | kHz | 577.35 | 0.0086938 | 577.35 | 0.0086087
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.324 | 1.1553
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.3551 | 1.1531
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 0.1 | 1 | kHz | 5.1972 | 0.99041 | 5.1951 | 0.99036
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 0.1 | 1 | kHz | 57.412 | 0.20433 | 57.411 | 0.20413
5502A | ACA | 3 | 11.001 | 1 | A | mA | 0.1 | 1 | kHz | 577.32 | 0.020884 | 577.32 | 0.020799
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3103 | 34.641
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3112 | 34.641
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.5084 | 34.627
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 1 | 5 | kHz | 20.888 | 33.346 | 20.867 | 33.347
5502A | ACA | 3 | 11.001 | 1 | A | mA | 1 | 5 | kHz | 547.31 | 13.25 | 547.3 | 13.249
5502A | ACA | 3 | 11.001 | 0.0001 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.3103 | 34.641
5502A | ACA | 3 | 11.001 | 0.001 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.3112 | 34.641
5502A | ACA | 3 | 11.001 | 0.01 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.5084 | 34.627
5502A | ACA | 3 | 11.001 | 0.1 | A | mA | 5 | 5.01 | kHz | 20.888 | 33.346 | 20.867 | 33.347
5502A | ACA | 3 | 11.001 | 1 | A | mA | 5 | 5.01 | kHz | 547.31 | 13.25 | 547.3 | 13.249
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 0.045 | 0.10001 | kHz | 5.8 | 1.4 | 5.7745 | 1.386
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 0.045 | 0.10001 | kHz | 5.8 | 1.4 | 5.7855 | 1.3857
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 6.9249 | 1.3567 | 6.9028 | 1.3548
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 55.153 | 0.59996 | 55.168 | 0.5931
5502A | ACA | 11 | 20.501 | 1 | A | mA | 0.045 | 0.10001 | kHz | 579.65 | 0.069317 | 579.66 | 0.065887
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7754 | 1.7324
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7852 | 1.7321
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 0.1 | 1 | kHz | 6.7671 | 1.7062 | 6.7466 | 1.7047
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 0.1 | 1 | kHz | 53.699 | 0.85925 | 53.709 | 0.85296
5502A | ACA | 11 | 20.501 | 1 | A | mA | 0.1 | 1 | kHz | 579.44 | 0.102 | 579.45 | 0.098572
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7742 | 34.641
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7745 | 34.641
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.8395 | 34.639
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 1 | 5 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
5502A | ACA | 11 | 20.501 | 1 | A | mA | 1 | 5 | kHz | 437.55 | 23.603 | 437.54 | 23.602
5502A | ACA | 11 | 20.501 | 0.0001 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.7742 | 34.641
5502A | ACA | 11 | 20.501 | 0.001 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.7745 | 34.641
5502A | ACA | 11 | 20.501 | 0.01 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.8395 | 34.639
5502A | ACA | 11 | 20.501 | 0.1 | A | mA | 5 | 5.01 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
5502A | ACA | 11 | 20.501 | 1 | A | mA | 5 | 5.01 | kHz | 437.55 | 23.603 | 437.54 | 23.602
5502A | ACA 10 turn | 1 | 11.001 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 3.3 | 2.7 | 0.97318 | 2.873
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 4.3265 | 2.6149 | 4.3204 | 2.6148
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 57.287 | 0.84059 | 57.285 | 0.83952
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 579.92 | 0.090643 | 579.92 | 0.090136
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 5800 | 0.0092241 | 5800 | 0.0090207
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 58000 | 0.00098346 | 58000 | 0.00090208
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.66 | 2.1 | 0.65508 | 2.0543
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 4.6183 | 1.7587 | 4.6169 | 1.7583
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.618 | 0.44323 | 57.618 | 0.44244
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 0.045 | 1 | kHz | 579.96 | 0.046302 | 579.96 | 0.045959
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 0.045 | 1 | kHz | 5800 | 0.004735 | 5800 | 0.0045977
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.405 | 7.1882
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 12.361 | 7.1187 | 12.349 | 7.1191
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 56.262 | 4.6574 | 56.248 | 4.6564
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 1 | 5.01 | kHz | 579.63 | 0.67288 | 579.63 | 0.67163
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 1 | 5.01 | kHz | 5800 | 0.068091 | 5800 | 0.067582
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.703 | 28.929
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 28.914 | 57.905 | 28.915
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 5 | 10 | kHz | 76.668 | 27.611 | 76.638 | 27.612
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 5 | 10 | kHz | 575.98 | 10.476 | 575.96 | 10.474
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 5 | 10 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
5502A | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.703 | 28.929
5502A | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 58 | 28.914 | 57.905 | 28.915
5502A | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 76.668 | 27.611 | 76.638 | 27.612
5502A | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 10 | 10.01 | kHz | 575.98 | 10.476 | 575.96 | 10.474
5502A | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 10 | 10.01 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.87 | 2.9 | 0.85389 | 2.8839
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 1.4281 | 2.8814 | 1.5429 | 2.8673
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 44.394 | 2.0188 | 44.317 | 2.0184
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 577.68 | 0.29774 | 577.66 | 0.2963
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 5799.8 | 0.030408 | 5799.8 | 0.029819
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.9 | 2.1 | 0.42297 | 2.1155
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 1.3908 | 2.0927 | 1.3648 | 2.0928
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 48.986 | 1.2457 | 48.967 | 1.244
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 0.045 | 1 | kHz | 578.74 | 0.1605 | 578.73 | 0.15918
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 0.045 | 1 | kHz | 5799.9 | 0.016507 | 5799.9 | 0.015972
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.174 | 7.2092
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.422 | 7.2034
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 33.827 | 6.6978 | 33.792 | 6.6979
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 1 | 5.01 | kHz | 566.19 | 1.8955 | 566.18 | 1.8942
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 1 | 5.01 | kHz | 5798.5 | 0.19808 | 5798.5 | 0.1975
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.627 | 28.936
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.682 | 28.935
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 5 | 10 | kHz | 63.353 | 28.822 | 63.613 | 28.797
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 5 | 10 | kHz | 458.92 | 21.118 | 458.89 | 21.117
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 5 | 10 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
5502A | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.627 | 28.936
5502A | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.682 | 28.935
5502A | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 63.353 | 28.822 | 63.613 | 28.797
5502A | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 10 | 10.01 | kHz | 458.92 | 21.118 | 458.89 | 21.117
5502A | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 10 | 10.01 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 12 | 2.1 | 11.984 | 2.086
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 36.579 | 1.9251 | 36.551 | 1.9245
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 568.43 | 0.54491 | 568.44 | 0.54347
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5798.8 | 0.057436 | 5798.8 | 0.056801
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 58000 | 0.0059371 | 58000 | 0.0056828
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 15 | 2.3 | 14.872 | 2.2866
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 37.348 | 2.1392 | 37.319 | 2.1387
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 0.1 | 1 | kHz | 566.33 | 0.65492 | 566.34 | 0.65351
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 0.1 | 1 | kHz | 5798.5 | 0.069561 | 5798.5 | 0.068926
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0071509 | 58000 | 0.0068966
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 1 | 5 | kHz | 24 | 35 | 23.064 | 34.699
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 1 | 5 | kHz | 24 | 35 | 25.112 | 34.684
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 1 | 5 | kHz | 210.04 | 33.396 | 210.02 | 33.395
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 1 | 5 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 1 | 5 | kHz | 57966 | 1.4661 | 57966 | 1.4652
5502A | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 24 | 35 | 23.064 | 34.699
5502A | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 24 | 35 | 25.112 | 34.684
5502A | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 5 | 5.01 | kHz | 210.04 | 33.396 | 210.02 | 33.395
5502A | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 5 | 5.01 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
5502A | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 5 | 5.01 | kHz | 57966 | 1.4661 | 57966 | 1.4652
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 42 | 2.4 | 38.553 | 2.4159
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 46.787 | 2.3924 | 46.65 | 2.3916
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 500.59 | 1.4072 | 500.63 | 1.403
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5788.7 | 0.17673 | 5788.8 | 0.17407
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57999 | 0.018526 | 57999 | 0.017455
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 42.202 | 2.6324
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 49.765 | 2.6108 | 49.649 | 2.61
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 0.1 | 1 | kHz | 490.55 | 1.6184 | 490.58 | 1.6144
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 0.1 | 1 | kHz | 5786.7 | 0.2093 | 5786.8 | 0.20664
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 0.1 | 1 | kHz | 57999 | 0.021803 | 57999 | 0.020732
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 1 | 5 | kHz | 58 | 35 | 57.612 | 34.699
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 1 | 5 | kHz | 58 | 35 | 58.262 | 34.697
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 1 | 5 | kHz | 79.07 | 34.897 | 123.66 | 34.491
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 1 | 5 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 1 | 5 | kHz | 57768 | 3.2889 | 57768 | 3.2878
5502A | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 58 | 35 | 57.612 | 34.699
5502A | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 58 | 35 | 58.262 | 34.697
5502A | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 5 | 5.01 | kHz | 79.07 | 34.897 | 123.66 | 34.491
5502A | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 5 | 5.01 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
5502A | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 5 | 5.01 | kHz | 57768 | 3.2889 | 57768 | 3.2878
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | -5.2 | 3 | -6.3515 | 3.0358
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | -5.2 | 3.0309 | -6.0166 | 3.031
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 24.504 | 2.6609 | 24.465 | 2.6608
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 569.88 | 0.59872 | 569.88 | 0.59787
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 5799 | 0.061619 | 5799 | 0.061255
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -5.078 | 2.227
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -4.615 | 2.2203
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 25.632 | 2.1228 | 32.901 | 1.7903
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 0.045 | 1 | kHz | 574.48 | 0.32468 | 574.48 | 0.3238
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 0.045 | 1 | kHz | 5799.4 | 0.033173 | 5799.4 | 0.032809
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 2.3 | 7.6 | 2.2215 | 7.5886
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 2.3482 | 7.5869 | 2.3433 | 7.5869
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 14.11 | 7.4273 | 14.091 | 7.4273
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 1 | 5.01 | kHz | 526.05 | 3.5184 | 526.05 | 3.5178
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 1 | 5.01 | kHz | 5793.5 | 0.40619 | 5793.5 | 0.40583
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 5 | 10 | kHz | 52 | 30 | 34.529 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 5 | 10 | kHz | 52 | 30.458 | 34.545 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 5 | 10 | kHz | 52 | 30.419 | 37.46 | 30.419
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 5 | 10 | kHz | 304.91 | 27.255 | 304.84 | 27.255
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 5 | 10 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
5502A | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 52 | 30 | 34.529 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 52 | 30.458 | 34.545 | 30.459
5502A | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 52 | 30.419 | 37.46 | 30.419
5502A | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 10 | 10.01 | kHz | 304.91 | 27.255 | 304.84 | 27.255
5502A | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 10 | 10.01 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 24 | 2.1 | 23.534 | 2.0875
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 37.146 | 2.0415 | 37.093 | 2.0415
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 540.32 | 0.97471 | 540.3 | 0.97417
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5795.1 | 0.11376 | 5795.1 | 0.11344
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 58000 | 0.011493 | 58000 | 0.011365
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 30 | 2.3 | 29.364 | 2.2879
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 41.544 | 2.2475 | 41.547 | 2.2467
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 0.1 | 1 | kHz | 534.45 | 1.1508 | 534.42 | 1.1503
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 0.1 | 1 | kHz | 5794.1 | 0.13793 | 5794.1 | 0.13761
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.01392 | 58000 | 0.013793
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 1 | 5 | kHz | 47 | 35 | 46.103 | 34.699
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 1 | 5 | kHz | 47 | 35 | 47.146 | 34.695
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 1 | 5 | kHz | 127.73 | 34.633 | 144.7 | 34.35
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 1 | 5 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 1 | 5 | kHz | 57864 | 2.9205 | 57864 | 2.9201
5502A | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 47 | 35 | 46.103 | 34.699
5502A | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 47 | 35 | 47.146 | 34.695
5502A | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 5 | 5.01 | kHz | 127.73 | 34.633 | 144.7 | 34.35
5502A | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 5 | 5.01 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
5502A | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 5 | 5.01 | kHz | 57864 | 2.9205 | 57864 | 2.9201
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 84 | 2.4 | 76.978 | 2.4161
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 84 | 2.4069 | 81.077 | 2.4099
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 407.38 | 1.9712 | 407.21 | 1.9701
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5755.8 | 0.34673 | 5755.9 | 0.34526
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57996 | 0.03551 | 57996 | 0.034907
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 98 | 2.6 | 84.28 | 2.6326
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 98 | 2.6065 | 88.046 | 2.6269
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 0.1 | 1 | kHz | 396.27 | 2.2061 | 396.12 | 2.205
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 0.1 | 1 | kHz | 5747.7 | 0.41071 | 5747.8 | 0.40925
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 0.1 | 1 | kHz | 57995 | 0.042063 | 57995 | 0.04146
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.22 | 34.699
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.69 | 34.698
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 148.62 | 34.646
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 1 | 5 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 1 | 5 | kHz | 57090 | 6.4831 | 57090 | 6.4816
5502A | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.22 | 34.699
5502A | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.69 | 34.698
5502A | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 148.62 | 34.646
5502A | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 5 | 5.01 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
5502A | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 5 | 5.01 | kHz | 57090 | 6.4831 | 57090 | 6.4816
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8 | 4.2301 | 2.8841
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.7673 | 4.3733 | 2.8834
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8076 | 17.829 | 2.819
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.01 | 0.045 | kHz | 530.85 | 1.3025 | 530.83 | 1.3019
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.01 | 0.045 | kHz | 5794.1 | 0.14924 | 5794.1 | 0.14886
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.01506 | 57992 | 0.01491
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 20.257 | 2.0293
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.045 | 0.1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.045 | 0.1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 20.257 | 2.0293
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.1 | 1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.1 | 1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 1 | 5.01 | kHz | 58 | 7.2 | 55.853 | 7.2093
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 1 | 5.01 | kHz | 58 | 7.209 | 55.901 | 7.2091
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 1 | 5.01 | kHz | 60.896 | 7.1858 | 60.843 | 7.1859
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 1 | 5.01 | kHz | 422.41 | 5.6956 | 422.3 | 5.6954
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 1 | 5.01 | kHz | 5763.9 | 0.97779 | 5763.9 | 0.97715
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 1 | 5.01 | kHz | 57996 | 0.099047 | 57948 | 0.098783
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 5 | 10 | kHz | 290 | 29 | 288.13 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 5 | 10 | kHz | 290 | 28.936 | 288.14 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 5 | 10 | kHz | 290 | 28.931 | 289.4 | 28.93
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 5 | 10 | kHz | 405.34 | 28.398 | 405.23 | 28.397
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 5 | 10 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 5 | 10 | kHz | 57941 | 1.6237 | 57264 | 1.6205
5502A | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 10 | 10.01 | kHz | 290 | 29 | 288.13 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 10 | 10.01 | kHz | 290 | 28.936 | 288.14 | 28.936
5502A | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 10 | 10.01 | kHz | 290 | 28.931 | 289.4 | 28.93
5502A | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 10 | 10.01 | kHz | 405.34 | 28.398 | 405.23 | 28.397
5502A | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 10 | 10.01 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
5502A | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 10 | 10.01 | kHz | 57941 | 1.6237 | 57264 | 1.6205
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 59 | 2.1 | 58.542 | 2.0879
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 62.205 | 2.0942 | 64.156 | 2.0802
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 445.36 | 1.6257 | 445.1 | 1.6253
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5769.8 | 0.28152 | 5769.8 | 0.28084
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57997 | 0.02869 | 57997 | 0.028411
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 74 | 2.3 | 73.171 | 2.2882
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 76.076 | 2.2962 | 78.162 | 2.2814
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 0.1 | 1 | kHz | 437.31 | 1.8466 | 437.03 | 1.8463
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 0.1 | 1 | kHz | 5764 | 0.34064 | 5763.9 | 0.33997
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 0.1 | 1 | kHz | 57996 | 0.034758 | 57996 | 0.034478
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.24 | 34.699
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.76 | 34.698
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 155.32 | 34.642
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 1 | 5 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 1 | 5 | kHz | 57174 | 7.1298 | 57174 | 7.1293
5502A | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.24 | 34.699
5502A | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.76 | 34.698
5502A | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 155.32 | 34.642
5502A | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 5 | 5.01 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
5502A | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 5 | 5.01 | kHz | 57174 | 7.1298 | 57174 | 7.1293
5502A | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 210 | 2.4 | 192.41 | 2.4161
5502A | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 210 | 2.4 | 193.98 | 2.4152
5502A | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 350.89 | 2.3236 | 350.74 | 2.3224
5502A | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5545.6 | 0.82078 | 5546.4 | 0.81772
5502A | ACA 50 turn | 550 | 1025.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57972 | 0.088607 | 57972 | 0.087215
5502A | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
5502A | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
5502A | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 0.1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
5502A | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 0.1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
5502A | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
5502A | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
5502A | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
5502A | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 0.01 | 0.04501 | Hz | 25 | 1700 | 23.181 | 1732
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 0.01 | 0.04501 | Hz | 25 | 1718.2 | 23.854 | 1717.9
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 0.01 | 0.04501 | Hz | 62.01 | 1125.5 | 61.999 | 1123.8
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 0.01 | 0.04501 | Hz | 580.38 | 157.54 | 580.38 | 156.33
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 0.01 | 0.04501 | Hz | 5800 | 16.202 | 5800 | 15.714
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 0.01 | 0.04501 | Hz | 58000 | 1.7665 | 58000 | 1.5715
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 0.01 | 0.04501 | Hz | 580000 | 0.23514 | 580000 | 0.15715
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 0.01 | 0.04501 | Hz | 5800000 | 0.046912 | 5800000 | 0.015715
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 1200 | 23.185 | 1155.4
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 1143.4 | 23.87 | 1142.8
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 0.045 | 10 | Hz | 62.237 | 673.84 | 62.225 | 671.86
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 0.045 | 10 | Hz | 580.43 | 86.247 | 580.43 | 85.036
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800.1 | 9.0181 | 5800.1 | 8.5307
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 0.045 | 10 | Hz | 58000 | 1.0481 | 58000 | 0.8531
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 0.045 | 10 | Hz | 580000 | 0.1633 | 580000 | 0.08531
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 0.045 | 10 | Hz | 5800000 | 0.039728 | 5800000 | 0.008531
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 10 | 20.01 | kHz | 25 | 1700 | 23.181 | 1732
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 10 | 20.01 | kHz | 25 | 1718.2 | 23.854 | 1717.9
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 10 | 20.01 | kHz | 62.01 | 1125.5 | 61.999 | 1123.8
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 10 | 20.01 | kHz | 580.38 | 157.54 | 580.38 | 156.33
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 10 | 20.01 | kHz | 5800 | 16.202 | 5800 | 15.714
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 10 | 20.01 | kHz | 58000 | 1.7665 | 58000 | 1.5715
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 10 | 20.01 | kHz | 580000 | 0.23514 | 580000 | 0.15715
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 10 | 20.01 | kHz | 5800000 | 0.046912 | 5800000 | 0.015715
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 20 | 50.01 | kHz | 24 | 2300 | 23.565 | 2300.9
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 20 | 50.01 | kHz | 24.224 | 2286.5 | 24.216 | 2286.3
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 20 | 50.01 | kHz | 61.904 | 1615.7 | 61.89 | 1614.4
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 20 | 50.01 | kHz | 580.34 | 247.93 | 580.33 | 246.73
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 20 | 50.01 | kHz | 5800 | 25.351 | 5800 | 24.864
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 20 | 50.01 | kHz | 58000 | 2.6816 | 58000 | 2.4866
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 20 | 50.01 | kHz | 580000 | 0.32665 | 580000 | 0.24866
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 20 | 50.01 | kHz | 5800000 | 0.056063 | 5800000 | 0.024866
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 50 | 100 | kHz | 40 | 4000 | 38.937 | 4027
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 50 | 100.01 | kHz | 40 | 4017.5 | 39.332 | 4018
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 50 | 100.01 | kHz | 68.858 | 3411.4 | 68.766 | 3411.7
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 50 | 100.01 | kHz | 580.9 | 730.69 | 580.86 | 729
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 50 | 100.01 | kHz | 5800.1 | 75.275 | 5800.1 | 74.549
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 50 | 100.01 | kHz | 58000 | 7.7474 | 58000 | 7.4566
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 50 | 100.01 | kHz | 580000 | 0.86197 | 580000 | 0.74566
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 50 | 100.01 | kHz | 5800000 | 0.12109 | 5800000 | 0.074566
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 100 | 500 | kHz | 89 | 11000 | 72.86 | 11473
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 100 | 500 | kHz | 89 | 11468 | 73.063 | 11468
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 100 | 500 | kHz | 91.464 | 11024 | 91.322 | 11026
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 100 | 500 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 100 | 500 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 100 | 500 | kHz | 58000 | 53.659 | 58000 | 52.994
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 100 | 500 | kHz | 580000 | 5.5657 | 580000 | 5.2995
5502A | ACV | 0 | 0.03301 | 1e-06 | V | µV | 500 | 501 | kHz | 89 | 11000 | 72.86 | 11473
5502A | ACV | 0 | 0.03301 | 1e-05 | V | µV | 500 | 501 | kHz | 89 | 11468 | 73.063 | 11468
5502A | ACV | 0 | 0.03301 | 0.0001 | V | µV | 500 | 501 | kHz | 91.464 | 11024 | 91.322 | 11026
5502A | ACV | 0 | 0.03301 | 0.001 | V | µV | 500 | 501 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
5502A | ACV | 0 | 0.03301 | 0.01 | V | µV | 500 | 501 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
5502A | ACV | 0 | 0.03301 | 0.1 | V | µV | 500 | 501 | kHz | 58000 | 53.659 | 58000 | 52.994
5502A | ACV | 0 | 0.03301 | 1 | V | µV | 500 | 501 | kHz | 580000 | 5.5657 | 580000 | 5.2995
5502A | ACV | 0 | 0.03301 | 10 | V | µV | 100 | 500.01 | kHz | 5800000 | 0.63642 | 5800000 | 0.52995
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 0.01 | 0.04501 | Hz | 33 | 550 | 23.236 | 577.95
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 0.01 | 0.04501 | Hz | 33 | 550.18 | 23.663 | 576.89
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 0.01 | 0.04501 | Hz | 55.234 | 504.47 | 55.135 | 504.68
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 0.01 | 0.04501 | Hz | 577.51 | 123.44 | 577.47 | 123.45
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 0.01 | 0.04501 | Hz | 5799.8 | 12.763 | 577.47 | 12.763
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 350 | 23.254 | 347.52
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 346.2 | 23.771 | 346.32
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 0.045 | 10 | Hz | 58.586 | 276.03 | 58.483 | 276.21
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 0.045 | 10 | Hz | 579.4 | 50.965 | 579.36 | 50.969
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800 | 5.1718 | 5799.9 | 5.1718
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 10 | 20.01 | kHz | 24 | 810 | 23.218 | 808.67
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 10 | 20.01 | kHz | 24 | 807.65 | 23.583 | 807.74
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 10 | 20.01 | kHz | 52.238 | 738.2 | 52.145 | 738.41
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 10 | 20.01 | kHz | 574.82 | 223.39 | 574.77 | 223.41
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 10 | 20.01 | kHz | 5799.5 | 23.686 | 5799.4 | 23.686
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 20 | 50.01 | kHz | 47 | 1200 | 46.389 | 1155.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 20 | 50.01 | kHz | 47 | 1154.8 | 46.618 | 1154.8
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 20 | 50.01 | kHz | 66.007 | 1107.9 | 65.939 | 1108
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 20 | 50.01 | kHz | 571.22 | 452.98 | 571.18 | 452.85
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 20 | 50.01 | kHz | 5799 | 51.02 | 5798.9 | 50.942
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 50 | 100.01 | kHz | 200 | 2700 | 196.5 | 2656.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 50 | 100.01 | kHz | 200 | 2656.2 | 196.56 | 2656.3
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 50 | 100.01 | kHz | 202.89 | 2641.9 | 202.83 | 2642
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 50 | 100.01 | kHz | 582.19 | 1932.6 | 582.1 | 1932.5
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 50 | 100.01 | kHz | 5796.8 | 308.19 | 5796.8 | 308.02
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 100 | 500 | kHz | 390 | 5800 | 381.75 | 5779.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 100 | 500 | kHz | 390 | 5779.3 | 381.8 | 5779.3
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 100 | 500 | kHz | 390 | 5771.9 | 384.93 | 5772
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 100 | 500 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 100 | 500 | kHz | 5783 | 1371.2 | 5783 | 1370.9
5502A | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 500 | 501 | kHz | 390 | 5800 | 381.75 | 5779.4
5502A | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 500 | 501 | kHz | 390 | 5779.3 | 381.8 | 5779.3
5502A | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 500 | 501 | kHz | 390 | 5771.9 | 384.93 | 5772
5502A | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 500 | 501 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
5502A | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 500 | 501 | kHz | 5783 | 1371.2 | 5783 | 1370.9
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 0.01 | 0.04501 | Hz | 0.11 | 0.56 | 0.10391 | 0.56221
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 0.01 | 0.04501 | Hz | 0.11 | 0.56033 | 0.10921 | 0.56037
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 0.48274 | 0.46448 | 0.4824 | 0.4645
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 5.7719 | 0.10338 | 5.7717 | 0.10332
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 57.997 | 0.010643 | 57.997 | 0.010612
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 0.045 | 10 | Hz | 0.07 | 0.35 | 0.069477 | 0.34756
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.079443 | 0.34499 | 0.079136 | 0.34505
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.52675 | 0.24887 | 0.5264 | 0.24885
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.7894 | 0.041576 | 5.7892 | 0.041503
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.999 | 0.0042263 | 57.999 | 0.004196
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 10 | 20.01 | kHz | 0.07 | 0.81 | 0.069435 | 0.80878
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 10 | 20.01 | kHz | 0.075008 | 0.80729 | 0.074829 | 0.80733
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 10 | 20.01 | kHz | 0.43427 | 0.71669 | 0.43391 | 0.71674
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 10 | 20.01 | kHz | 5.7428 | 0.2035 | 5.7426 | 0.20344
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 10 | 20.01 | kHz | 57.994 | 0.021456 | 57.994 | 0.021426
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 20 | 50.01 | kHz | 0.07 | 1.2 | 0.069506 | 1.1559
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 20 | 50.01 | kHz | 0.073711 | 1.1548 | 0.073371 | 1.1549
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 20 | 50.01 | kHz | 0.38065 | 1.0748 | 0.37987 | 1.0749
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 20 | 50.01 | kHz | 5.6886 | 0.39166 | 5.6883 | 0.39154
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 20 | 50.01 | kHz | 57.988 | 0.043219 | 57.988 | 0.043148
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 50 | 100.01 | kHz | 0.24 | 2.7 | 0.23108 | 2.6571
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 50 | 100.01 | kHz | 0.24 | 2.6566 | 0.23256 | 2.6567
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 50 | 100.01 | kHz | 0.38846 | 2.6152 | 0.38733 | 2.6154
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 50 | 100.01 | kHz | 5.3723 | 1.6168 | 5.3715 | 1.6167
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 50 | 100.01 | kHz | 57.935 | 0.23029 | 57.935 | 0.23012
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 100 | 500 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 100 | 500 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 100 | 500 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 100 | 500 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 100 | 500 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
5502A | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 500 | 501 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
5502A | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 500 | 501 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
5502A | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 500 | 501 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
5502A | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 500 | 501 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
5502A | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 500 | 501 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 0.01 | 0.04501 | Hz | 0.93 | 0.58 | 0.92292 | 0.5783
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 0.98726 | 0.57661 | 0.98668 | 0.57662
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 4.8574 | 0.48407 | 4.8565 | 0.48406
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 57.704 | 0.11064 | 57.704 | 0.11059
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 0.01 | 0.04501 | Hz | 579.97 | 0.011401 | 579.97 | 0.011382
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.7 | 0.35 | 0.69218 | 0.34795
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.78957 | 0.34543 | 0.78891 | 0.34543
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.2633 | 0.24921 | 5.2624 | 0.24918
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.892 | 0.04162 | 57.892 | 0.041575
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 0.045 | 10 | Hz | 579.99 | 0.0042219 | 579.99 | 0.0042034
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 10 | 20.01 | kHz | 0.7 | 0.81 | 0.69304 | 0.80896
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 10 | 20.01 | kHz | 0.74744 | 0.8075 | 0.74699 | 0.80751
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 10 | 20.01 | kHz | 4.3392 | 0.7169 | 4.3383 | 0.7169
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 10 | 20.01 | kHz | 57.426 | 0.20355 | 57.426 | 0.20351
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 10 | 20.01 | kHz | 579.94 | 0.021452 | 579.94 | 0.021433
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 20 | 50.01 | kHz | 0.7 | 1.2 | 0.69122 | 1.1565
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 20 | 50.01 | kHz | 0.73466 | 1.1553 | 0.73165 | 1.1554
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 20 | 50.01 | kHz | 3.8002 | 1.0754 | 3.7971 | 1.0754
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 20 | 50.01 | kHz | 56.883 | 0.3919 | 56.882 | 0.39182
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 20 | 50.01 | kHz | 579.88 | 0.043223 | 579.88 | 0.043182
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 50 | 100.01 | kHz | 2.3 | 2.7 | 2.3079 | 2.6577
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 50 | 100.01 | kHz | 2.3265 | 2.6572 | 2.3234 | 2.6573
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 50 | 100.01 | kHz | 3.8797 | 2.6158 | 3.8711 | 2.616
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 50 | 100.01 | kHz | 53.719 | 1.6172 | 53.713 | 1.6172
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 50 | 100.01 | kHz | 579.35 | 0.23028 | 579.35 | 0.23021
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3364 | 2.6995
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3501 | 2.699
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 3.831 | 2.6581
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 100 | 500 | kHz | 47.329 | 3.5597 | 53.608 | 1.6549
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 100 | 500 | kHz | 579.33 | 0.23723 | 579.33 | 0.23717
5502A | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 2.3364 | 2.6995
5502A | ACV | 3.3 | 33.001 | 0.001 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 2.3501 | 2.699
5502A | ACV | 3.3 | 33.001 | 0.01 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 3.831 | 2.6581
5502A | ACV | 3.3 | 33.001 | 0.1 | V | mV | 500 | 501 | kHz | 47.329 | 3.5597 | 53.608 | 1.6549
5502A | ACV | 3.3 | 33.001 | 1 | V | mV | 500 | 501 | kHz | 579.33 | 0.23723 | 579.33 | 0.23717
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 3.5 | 0.58 | 3.4705 | 0.57868
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 4.2872 | 0.57647 | 4.2606 | 0.57652
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 46.641 | 0.47359 | 46.602 | 0.47361
5502A | ACV | 33 | 330.01 | 1 | V | mV | 0.01 | 0.04501 | Hz | 576.98 | 0.10542 | 576.96 | 0.10533
5502A | ACV | 33 | 330.01 | 10 | V | mV | 0.01 | 0.04501 | Hz | 5799.7 | 0.010862 | 5799.7 | 0.010822
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 0.045 | 10 | Hz | 12 | 0.92 | 10.393 | 0.92459
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 0.045 | 10 | Hz | 12 | 0.92331 | 10.834 | 0.9234
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 0.045 | 10 | Hz | 43.309 | 0.841 | 43.21 | 0.84115
5502A | ACV | 33 | 330.01 | 1 | V | mV | 0.045 | 10 | Hz | 572.72 | 0.26556 | 572.68 | 0.26538
5502A | ACV | 33 | 330.01 | 10 | V | mV | 0.045 | 10 | Hz | 5799.2 | 0.028474 | 5799.2 | 0.028386
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 10 | 20.01 | kHz | 24 | 1 | 10.388 | 1.04
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 10 | 20.01 | kHz | 24 | 1.0332 | 10.796 | 1.0389
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 10 | 20.01 | kHz | 41.651 | 0.95969 | 41.554 | 0.95985
5502A | ACV | 33 | 330.01 | 1 | V | mV | 10 | 20.01 | kHz | 570.92 | 0.32864 | 570.88 | 0.32848
5502A | ACV | 33 | 330.01 | 10 | V | mV | 10 | 20.01 | kHz | 5799 | 0.035762 | 5799 | 0.035674
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 20 | 50.01 | kHz | 22 | 1.4 | 10.266 | 1.3912
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 20 | 50.01 | kHz | 22 | 1.4 | 10.577 | 1.3904
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 20 | 50.01 | kHz | 36.02 | 1.3575 | 37.167 | 1.3205
5502A | ACV | 33 | 330.01 | 1 | V | mV | 20 | 50.01 | kHz | 564.6 | 0.55034 | 564.56 | 0.55021
5502A | ACV | 33 | 330.01 | 10 | V | mV | 20 | 50.01 | kHz | 5798.2 | 0.063012 | 5798.2 | 0.062925
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 50 | 100 | kHz | 92 | 2.8 | 91.937 | 2.7856
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 50 | 100 | kHz | 92.063 | 2.7853 | 92.033 | 2.7854
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 50 | 100 | kHz | 101.77 | 2.761 | 101.69 | 2.7611
5502A | ACV | 33 | 330.01 | 1 | V | mV | 50 | 100 | kHz | 546.63 | 1.8765 | 546.54 | 1.8763
5502A | ACV | 33 | 330.01 | 10 | V | mV | 50 | 100 | kHz | 5793.6 | 0.28499 | 5793.5 | 0.28476
5502A | ACV | 33 | 330.01 | 0.001 | V | mV | 100 | 100.1 | kHz | 92 | 2.8 | 91.937 | 2.7856
5502A | ACV | 33 | 330.01 | 0.01 | V | mV | 100 | 100.1 | kHz | 92.063 | 2.7853 | 92.033 | 2.7854
5502A | ACV | 33 | 330.01 | 0.1 | V | mV | 100 | 100.1 | kHz | 101.77 | 2.761 | 101.69 | 2.7611
5502A | ACV | 33 | 330.01 | 1 | V | mV | 100 | 100.1 | kHz | 546.63 | 1.8765 | 546.54 | 1.8763
5502A | ACV | 33 | 330.01 | 10 | V | mV | 100 | 100.1 | kHz | 5793.6 | 0.28499 | 5793.5 | 0.28476
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 0.045 | 1 | Hz | 24 | 0.58 | 23.335 | 0.57834
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 0.045 | 1 | Hz | 33.354 | 0.57134 | 33.321 | 0.57119
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 0.045 | 1 | Hz | 510.18 | 0.32785 | 510.26 | 0.32724
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 0.045 | 1 | Hz | 5790.3 | 0.041506 | 5790.4 | 0.041123
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 0.045 | 1 | Hz | 57999 | 0.0042793 | 57999 | 0.004125
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 1 | 5.01 | kHz | 28 | 0.92 | 23.262 | 0.92437
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 1 | 5.01 | kHz | 29.895 | 0.91967 | 29.873 | 0.91957
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 1 | 5.01 | kHz | 446.23 | 0.66766 | 446.26 | 0.66723
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 1 | 5.01 | kHz | 5775.4 | 0.10274 | 5775.5 | 0.10236
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 1 | 5.01 | kHz | 57998 | 0.010468 | 57998 | 0.010314
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 5 | 10.01 | kHz | 64 | 1 | 23.23 | 1.0398
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 5 | 10.01 | kHz | 64 | 1.0355 | 29.165 | 1.0355
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 5 | 10.01 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 5 | 10.01 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 5 | 10.01 | kHz | 57997 | 0.013152 | 57997 | 0.012998
5502A | ACV | 330 | 1020.01 | 0.001 | V | mV | 10 | 10.01 | kHz | 64 | 1 | 23.23 | 1.0398
5502A | ACV | 330 | 1020.01 | 0.01 | V | mV | 10 | 10.01 | kHz | 64 | 1.0355 | 29.165 | 1.0355
5502A | ACV | 330 | 1020.01 | 0.1 | V | mV | 10 | 10.01 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
5502A | ACV | 330 | 1020.01 | 1 | V | mV | 10 | 10.01 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
5502A | ACV | 330 | 1020.01 | 10 | V | mV | 10 | 10.01 | kHz | 57997 | 0.013152 | 57997 | 0.012998
5502A | ACW | 0 | 1.09e-05 | 1e-10 | WAC | WAC |  |  |  | 0 | 0.025 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-09 | WAC | WAC |  |  |  | 0 | 0.025178 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.062989 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.58053 | 0 | 0
5502A | ACW | 0 | 1.09e-05 | 1e-06 | WAC | WAC |  |  |  | 0 | 5.8001 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-09 | WAC | WAC |  |  |  | 0 | 0.0026 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.0026 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0058884 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.05332 | 0 | 0
5502A | ACW | 1.09e-05 | 0.000109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.53261 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.0017 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0017 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.0055645 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.053285 | 0 | 0
5502A | ACW | 0.000109 | 0.0010901 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0013 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.0013483 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.0054702 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.053275 | 0 | 0
5502A | ACW | 0.00109 | 0.0109 | 0.001 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.00078 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.00094316 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.0053847 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 0.001 | WAC | WAC |  |  |  | 0 | 0.053267 | 0 | 0
5502A | ACW | 0.0109 | 0.109 | 0.01 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.00078 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.00094283 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.001 | WAC | WAC |  |  |  | 0 | 0.0053846 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.01 | WAC | WAC |  |  |  | 0 | 0.053267 | 0 | 0
5502A | ACW | 0.109 | 1.0901 | 0.1 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.0013 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.001 | WAC | WAC |  |  |  | 0 | 0.0013347 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0054669 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.053275 | 0 | 0
5502A | ACW | 1.09 | 10.9001 | 1 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 0.001 | WAC | WAC |  |  |  | 0 | 0.00088 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0010256 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0053997 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.053268 | 0 | 0
5502A | ACW | 10.9 | 109.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | ACW | 109 | 363.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0012 | 0 | 0
5502A | ACW | 109 | 363.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0012284 | 0 | 0
5502A | ACW | 109 | 363.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.0054419 | 0 | 0
5502A | ACW | 109 | 363.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.053273 | 0 | 0
5502A | ACW | 363 | 990.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.001 | 0 | 0
5502A | ACW | 363 | 990.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.001007 | 0 | 0
5502A | ACW | 363 | 990.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.0018832 | 0 | 0
5502A | ACW | 363 | 990.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.01601 | 0 | 0
5502A | ACW | 363 | 990.0001 | 100 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
5502A | ACW | 990 | 3630.1 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 990 | 3630.1 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 990 | 3630.1 | 10 | WAC | WAC |  |  |  | 0 | 0.0061979 | 0 | 0
5502A | ACW | 990 | 3630.1 | 100 | WAC | WAC |  |  |  | 0 | 0.058622 | 0 | 0
5502A | ACW | 990 | 3630.1 | 1000 | WAC | WAC |  |  |  | 0 | 0.58586 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0016 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 1 | WAC | WAC |  |  |  | 0 | 0.0016 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 10 | WAC | WAC |  |  |  | 0 | 0.0022012 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 100 | WAC | WAC |  |  |  | 0 | 0.016051 | 0 | 0
5502A | ACW | 3630 | 11200.1 | 1000 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
5502A | ACW | 11200 | 20900 | 0.1 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
5502A | ACW | 11200 | 20900 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 11200 | 20900 | 10 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 11200 | 20900 | 100 | WAC | WAC |  |  |  | 0 | 0.0055466 | 0 | 0
5502A | ACW | 11200 | 20900 | 1000 | WAC | WAC |  |  |  | 0 | 0.051733 | 0 | 0
5502A | ACW | 20900 | 20901 | 0.1 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
5502A | ACW | 20900 | 20901 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 20900 | 20901 | 10 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
5502A | ACW | 20900 | 20901 | 100 | WAC | WAC |  |  |  | 0 | 0.0055466 | 0 | 0
5502A | ACW | 20900 | 20901 | 1000 | WAC | WAC |  |  |  | 0 | 0.051733 | 0 | 0
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-13 | F | nF |  |  |  | 0.012 | 5800000 | 0.01155 | 5775200
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-12 | F | nF |  |  |  | 0.012 | 5800000 | 0.011563 | 5772800
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-11 | F | nF |  |  |  | 0.012845 | 5547800 | 0.012838 | 5547800
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-10 | F | nF |  |  |  | 0.058705 | 2016900 | 0.058698 | 2014500
5502A | Capacitance | 2.2e-10 | 3.3e-09 | 1e-09 | F | nF |  |  |  | 0.57745 | 218190 | 0.57745 | 217000
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-12 | F | nF |  |  |  | 0.012 | 2900000 | 0.011576 | 2885200
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-11 | F | nF |  |  |  | 0.012482 | 2856200 | 0.012511 | 2834700
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-10 | F | nF |  |  |  | 0.056911 | 1398700 | 0.056878 | 1390600
5502A | Capacitance | 3.3e-09 | 1.1001e-08 | 1e-09 | F | nF |  |  |  | 0.57722 | 165280 | 0.57721 | 160660
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-12 | F | nF |  |  |  | 0.12 | 2900000 | 0.11547 | 2886700
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-11 | F | nF |  |  |  | 0.12 | 2900000 | 0.11561 | 2885200
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-10 | F | nF |  |  |  | 0.12801 | 2745100 | 0.12797 | 2743500
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-09 | F | nF |  |  |  | 0.58647 | 856770 | 0.58644 | 853690
5502A | Capacitance | 1.1e-08 | 3.3001e-08 | 1e-08 | F | nF |  |  |  | 5.7745 | 853270 | 5.7744 | 89445
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-11 | F | nF |  |  |  | 0.12 | 2900000 | 0.11565 | 2886200
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-10 | F | nF |  |  |  | 0.12459 | 2858300 | 0.125 | 2835700
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-09 | F | nF |  |  |  | 0.56898 | 1398000 | 0.56874 | 1390900
5502A | Capacitance | 3.3e-08 | 1.1001e-07 | 1e-08 | F | nF |  |  |  | 5.7722 | 164700 | 5.7721 | 160690
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-11 | F | nF |  |  |  | 0.35 | 2900000 | 0.34649 | 2886700
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-10 | F | nF |  |  |  | 0.35 | 2900000 | 0.34958 | 2881200
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-09 | F | nF |  |  |  | 0.60931 | 2463100 | 0.60906 | 2462300
5502A | Capacitance | 1.1e-07 | 3.3001e-07 | 1e-08 | F | nF |  |  |  | 5.7585 | 484500 | 5.7585 | 483140
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-10 | F | nF |  |  |  | 1.2 | 2900000 | 1.1566 | 2886200
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-09 | F | nF |  |  |  | 1.2438 | 2860200 | 1.2501 | 2835600
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-08 | F | nF |  |  |  | 5.6891 | 1394000 | 5.6874 | 1390900
5502A | Capacitance | 3.3e-07 | 1.1001e-06 | 1e-07 | F | nF |  |  |  | 57.722 | 162480 | 57.721 | 160690
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-10 | F | nF |  |  |  | 3.5 | 2900000 | 3.4649 | 2886700
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-09 | F | nF |  |  |  | 3.5 | 2900000 | 3.4958 | 2881200
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-08 | F | nF |  |  |  | 6.0937 | 2463200 | 6.0906 | 2462300
5502A | Capacitance | 1.1e-06 | 3.3001e-06 | 1e-07 | F | nF |  |  |  | 57.585 | 484680 | 57.585 | 483140
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-09 | F | nF |  |  |  | 12 | 2900000 | 11.566 | 2886200
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-08 | F | nF |  |  |  | 12.429 | 2861100 | 12.501 | 2835600
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-07 | F | nF |  |  |  | 56.88 | 1394300 | 56.874 | 1390900
5502A | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-06 | F | nF |  |  |  | 577.21 | 162540 | 577.21 | 160690
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-09 | F | nF |  |  |  | 36 | 4600000 | 34.648 | 4618700
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-08 | F | nF |  |  |  | 36 | 4600000 | 34.894 | 4613900
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-07 | F | nF |  |  |  | 56.814 | 4211300 | 56.808 | 4211000
5502A | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-06 | F | nF |  |  |  | 572.02 | 1057800 | 572.02 | 1057200
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-08 | F | nF |  |  |  | 120 | 5200000 | 115.61 | 5195600
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-07 | F | nF |  |  |  | 121.85 | 5183200 | 122.72 | 5152800
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-06 | F | nF |  |  |  | 536.63 | 3284200 | 536.45 | 3282000
5502A | Capacitance | 3.3e-05 | 0.00011001 | 1e-05 | F | nF |  |  |  | 5766.3 | 438500 | 5766.3 | 436530
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-08 | F | µF |  |  |  | 0.35 | 5200 | 0.34613 | 5199.8
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-07 | F | µF |  |  |  | 0.35 | 5200 | 0.34836 | 5195.3
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-06 | F | µF |  |  |  | 0.55623 | 4803.8 | 0.55608 | 4803.7
5502A | Capacitance | 0.00011 | 0.00033001 | 1e-05 | F | µF |  |  |  | 5.7038 | 1294.4 | 5.7038 | 1293.8
5502A | Capacitance | 0.00033 | 0.0011001 | 1e-07 | F | µF |  |  |  | 1.2 | 5200 | 1.1558 | 5198.3
5502A | Capacitance | 0.00033 | 0.0011001 | 1e-06 | F | µF |  |  |  | 1.2193 | 5182.5 | 1.2268 | 5155.5
5502A | Capacitance | 0.00033 | 0.0011001 | 1e-05 | F | µF |  |  |  | 5.3658 | 3286.5 | 5.3641 | 3284.2
5502A | Capacitance | 0.00033 | 0.0011001 | 0.0001 | F | µF |  |  |  | 57.663 | 438.86 | 57.663 | 436.89
5502A | Capacitance | 0.0011 | 0.0033001 | 1e-07 | F | µF |  |  |  | 3.5 | 5200 | 3.4636 | 5198.9
5502A | Capacitance | 0.0011 | 0.0033001 | 1e-06 | F | µF |  |  |  | 3.5 | 5200 | 3.4866 | 5194.4
5502A | Capacitance | 0.0011 | 0.0033001 | 1e-05 | F | µF |  |  |  | 5.5761 | 4803.9 | 5.5634 | 4802.9
5502A | Capacitance | 0.0011 | 0.0033001 | 0.0001 | F | µF |  |  |  | 57.043 | 1299.1 | 57.038 | 1293.7
5502A | Capacitance | 0.0033 | 0.011001 | 1e-06 | F | µF |  |  |  | 13 | 5200 | 12.076 | 5198.3
5502A | Capacitance | 0.0033 | 0.011001 | 1e-05 | F | µF |  |  |  | 13 | 5200 | 12.713 | 5155.6
5502A | Capacitance | 0.0033 | 0.011001 | 0.0001 | F | µF |  |  |  | 53.666 | 3287.9 | 53.64 | 3284.3
5502A | Capacitance | 0.0033 | 0.011001 | 0.001 | F | µF |  |  |  | 576.63 | 439.99 | 576.63 | 436.9
5502A | Capacitance | 0.011 | 0.033001 | 1e-06 | F | µF |  |  |  | 35 | 8700 | 34.642 | 8661.9
5502A | Capacitance | 0.011 | 0.033001 | 1e-05 | F | µF |  |  |  | 35 | 8700 | 34.806 | 8658.5
5502A | Capacitance | 0.011 | 0.033001 | 0.0001 | F | µF |  |  |  | 50.508 | 8339.7 | 50.437 | 8339.6
5502A | Capacitance | 0.011 | 0.033001 | 0.001 | F | µF |  |  |  | 557.55 | 3119.3 | 557.52 | 3115.8
5502A | Capacitance | 0.033 | 0.11001 | 1e-05 | F | µF* |  |  |  | 120 | 13000 | 115.63 | 12701
5502A | Capacitance | 0.033 | 0.11001 | 0.0001 | F | µF* |  |  |  | 120 | 13000 | 119.5 | 12676
5502A | Capacitance | 0.033 | 0.11001 | 0.001 | F | µF* |  |  |  | 430.99 | 10809 | 430.28 | 10808
5502A | Capacitance | 0.033 | 0.11001 | 0.01 | F | µF* |  |  |  | 5725.6 | 2215.6 | 5725.3 | 2210.1
5502A | Capacitance | 0.11 | 0.1101 | 1e-05 | F | µF |  |  |  | 120 | 13000 | 115.63 | 12701
5502A | Capacitance | 0.11 | 0.1101 | 0.0001 | F | µF |  |  |  | 120 | 13000 | 119.5 | 12676
5502A | Capacitance | 0.11 | 0.1101 | 0.001 | F | µF |  |  |  | 430.99 | 10809 | 430.28 | 10808
5502A | Capacitance | 0.11 | 0.1101 | 0.01 | F | µF |  |  |  | 5725.6 | 2215.6 | 5725.3 | 2210.1
5502A | CDE | -250 | -99.9 | 0.1 | °C | °C |  |  |  | 0.58423 | 0 | 0.58258 | 0
5502A | CDE | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.19859 | 0 | 0.19765 | 0
5502A | CDE | -25 | 350.1 | 0.1 | °C | °C |  |  |  | 0.17588 | 0 | 0.17482 | 0
5502A | CDE | 350 | 650.1 | 0.1 | °C | °C |  |  |  | 0.19683 | 0 | 0.19628 | 0
5502A | CDE | 650 | 1000 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
5502A | CDE | 1000 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
5502A | CDE | -250 | -99.9 | 1 | °C | °C |  |  |  | 0.82286 | 0 | 0.81817 | 0
5502A | CDE | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.60874 | 0 | 0.60751 | 0
5502A | CDE | -25 | 350.1 | 1 | °C | °C |  |  |  | 0.60171 | 0 | 0.60047 | 0
5502A | CDE | 350 | 650.1 | 1 | °C | °C |  |  |  | 0.60777 | 0 | 0.60706 | 0
5502A | CDE | 650 | 1000 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
5502A | CDE | 1000 | 1000.1 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
5502A | CDJ | -210 | -99.9 | 0.1 | °C | °C |  |  |  | 0.32482 | 0 | 0.31951 | 0
5502A | CDJ | -100 | -29.9 | 0.1 | °C | °C |  |  |  | 0.20266 | 0 | 0.19649 | 0
5502A | CDJ | -30 | 150.1 | 0.1 | °C | °C |  |  |  | 0.18177 | 0 | 0.17487 | 0
5502A | CDJ | 150 | 760.1 | 0.1 | °C | °C |  |  |  | 0.2089 | 0 | 0.20722 | 0
5502A | CDJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
5502A | CDJ | 1200 | 1200.1 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
5502A | CDJ | -210 | -99.9 | 1 | °C | °C |  |  |  | 0.66767 | 0 | 0.65733 | 0
5502A | CDJ | -100 | -29.9 | 1 | °C | °C |  |  |  | 0.61519 | 0 | 0.60713 | 0
5502A | CDJ | -30 | 150.1 | 1 | °C | °C |  |  |  | 0.60863 | 0 | 0.60048 | 0
5502A | CDJ | 150 | 760.1 | 1 | °C | °C |  |  |  | 0.61296 | 0 | 0.61069 | 0
5502A | CDJ | 760 | 1200 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
5502A | CDJ | 1200 | 1200.1 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
5502A | CDK | -200 | -99.9 | 0.1 | °C | °C |  |  |  | 0.39349 | 0 | 0.38767 | 0
5502A | CDK | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.22869 | 0 | 0.21853 | 0
5502A | CDK | -25 | 120.1 | 0.1 | °C | °C |  |  |  | 0.20428 | 0 | 0.19653 | 0
5502A | CDK | 120 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.31026 | 0 | 0.30758 | 0
5502A | CDK | 1000 | 1372 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
5502A | CDK | 1372 | 1372.1 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
5502A | CDK | -200 | -99.9 | 1 | °C | °C |  |  |  | 0.70603 | 0 | 0.69303 | 0
5502A | CDK | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.62924 | 0 | 0.61462 | 0
5502A | CDK | -25 | 120.1 | 1 | °C | °C |  |  |  | 0.61728 | 0 | 0.60715 | 0
5502A | CDK | 120 | 1000.1 | 1 | °C | °C |  |  |  | 0.65669 | 0 | 0.65162 | 0
5502A | CDK | 1000 | 1372 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
5502A | CDK | 1372 | 1372.1 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
5502A | CDN | -200 | -99.9 | 0.1 | °C | °C |  |  |  | 0.47201 | 0 | 0.46682 | 0
5502A | CDN | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.26968 | 0 | 0.26325 | 0
5502A | CDN | -25 | 120.1 | 0.1 | °C | °C |  |  |  | 0.23706 | 0 | 0.22973 | 0
5502A | CDN | 120 | 410.1 | 0.1 | °C | °C |  |  |  | 0.22303 | 0 | 0.21855 | 0
5502A | CDN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
5502A | CDN | 1300 | 1300.1 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
5502A | CDN | -200 | -99.9 | 1 | °C | °C |  |  |  | 0.75326 | 0 | 0.74022 | 0
5502A | CDN | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.64265 | 0 | 0.6319 | 0
5502A | CDN | -25 | 120.1 | 1 | °C | °C |  |  |  | 0.62966 | 0 | 0.61869 | 0
5502A | CDN | 120 | 410.1 | 1 | °C | °C |  |  |  | 0.62103 | 0 | 0.61463 | 0
5502A | CDN | 410 | 1300 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
5502A | CDN | 1300 | 1300.1 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
5502A | CDPT100 | -200 | 0.1 | 0.1 | °C | °C |  |  |  | 0.13102 | 0 | 0.11627 | 0
5502A | CDPT100 | 0 | 100.1 | 0.1 | °C | °C |  |  |  | 0.16571 | 0 | 0.15431 | 0
5502A | CDPT100 | 100 | 300.1 | 0.1 | °C | °C |  |  |  | 0.16571 | 0 | 0.15431 | 0
5502A | CDPT100 | 300 | 400.1 | 0.1 | °C | °C |  |  |  | 0.22203 | 0 | 0.21366 | 0
5502A | CDPT100 | 400 | 630.1 | 0.1 | °C | °C |  |  |  | 0.26 | 0 | 0.25288 | 0
5502A | CDPT100 | 630 | 800 | 0.1 | °C | °C |  |  |  | 0.43497 | 0 | 0.43075 | 0
5502A | CDPT100 | 800 | 800.1 | 0.1 | °C | °C |  |  |  | 0.43497 | 0 | 0.43075 | 0
5502A | CDPT100 | -200 | 0.1 | 1 | °C | °C |  |  |  | 0.60095 | 0 | 0.58869 | 0
5502A | CDPT100 | 0 | 100.1 | 1 | °C | °C |  |  |  | 0.60946 | 0 | 0.59736 | 0
5502A | CDPT100 | 100 | 300.1 | 1 | °C | °C |  |  |  | 0.60946 | 0 | 0.59736 | 0
5502A | CDPT100 | 300 | 400.1 | 1 | °C | °C |  |  |  | 0.62712 | 0 | 0.61537 | 0
5502A | CDPT100 | 400 | 630.1 | 1 | °C | °C |  |  |  | 0.64154 | 0 | 0.63006 | 0
5502A | CDPT100 | 630 | 800 | 1 | °C | °C |  |  |  | 0.73019 | 0 | 0.72012 | 0
5502A | CDPT100 | 800 | 800.1 | 1 | °C | °C |  |  |  | 0.73019 | 0 | 0.72012 | 0
5502A | CDPT1000 | -200 | 0.1 | 0.1 | °C | °C |  |  |  | 0.10309 | 0 | 0.083548 | 0
5502A | CDPT1000 | 0 | 100.1 | 0.1 | °C | °C |  |  |  | 0.11707 | 0 | 0.10029 | 0
5502A | CDPT1000 | 100 | 260.1 | 0.1 | °C | °C |  |  |  | 0.11707 | 0 | 0.10029 | 0
5502A | CDPT1000 | 260 | 300.1 | 0.1 | °C | °C |  |  |  | 0.15328 | 0 | 0.14088 | 0
5502A | CDPT1000 | 300 | 400.1 | 0.1 | °C | °C |  |  |  | 0.17 | 0 | 0.1589 | 0
5502A | CDPT1000 | 400 | 600.1 | 0.1 | °C | °C |  |  |  | 0.17204 | 0 | 0.16109 | 0
5502A | CDPT1000 | 600 | 630 | 0.1 | °C | °C |  |  |  | 0.47301 | 0 | 0.46914 | 0
5502A | CDPT1000 | 630 | 630.1 | 0.1 | °C | °C |  |  |  | 0.47301 | 0 | 0.46914 | 0
5502A | CDPT1000 | -200 | 0.1 | 1 | °C | °C |  |  |  | 0.59549 | 0 | 0.58311 | 0
5502A | CDPT1000 | 0 | 100.1 | 1 | °C | °C |  |  |  | 0.59806 | 0 | 0.58574 | 0
5502A | CDPT1000 | 100 | 260.1 | 1 | °C | °C |  |  |  | 0.59806 | 0 | 0.58574 | 0
5502A | CDPT1000 | 260 | 300.1 | 1 | °C | °C |  |  |  | 0.60619 | 0 | 0.59404 | 0
5502A | CDPT1000 | 300 | 400.1 | 1 | °C | °C |  |  |  | 0.61063 | 0 | 0.59857 | 0
5502A | CDPT1000 | 400 | 600.1 | 1 | °C | °C |  |  |  | 0.61121 | 0 | 0.59915 | 0
5502A | CDPT1000 | 600 | 630 | 1 | °C | °C |  |  |  | 0.75347 | 0 | 0.74372 | 0
5502A | CDPT1000 | 630 | 630.1 | 1 | °C | °C |  |  |  | 0.75347 | 0 | 0.74372 | 0
5502A | CDR | 0 | 250.1 | 0.1 | °C | °C |  |  |  | 0.66953 | 0 | 0.66592 | 0
5502A | CDR | 250 | 400.1 | 0.1 | °C | °C |  |  |  | 0.41654 | 0 | 0.41163 | 0
5502A | CDR | 400 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.39417 | 0 | 0.38897 | 0
5502A | CDR | 1000 | 1767 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
5502A | CDR | 1767 | 1767.1 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
5502A | CDR | 0 | 250.1 | 1 | °C | °C |  |  |  | 0.89034 | 0 | 0.87946 | 0
5502A | CDR | 250 | 400.1 | 1 | °C | °C |  |  |  | 0.71813 | 0 | 0.70671 | 0
5502A | CDR | 400 | 1000.1 | 1 | °C | °C |  |  |  | 0.70538 | 0 | 0.69376 | 0
5502A | CDR | 1000 | 1767 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
5502A | CDR | 1767 | 1767.1 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
5502A | CDS | 0 | 250.1 | 0.1 | °C | °C |  |  |  | 0.55632 | 0 | 0.55207 | 0
5502A | CDS | 250 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.42876 | 0 | 0.42399 | 0
5502A | CDS | 1000 | 1400.1 | 0.1 | °C | °C |  |  |  | 0.4386 | 0 | 0.43394 | 0
5502A | CDS | 1400 | 1767 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
5502A | CDS | 1767 | 1767.1 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
5502A | CDS | 0 | 250.1 | 1 | °C | °C |  |  |  | 0.80847 | 0 | 0.79673 | 0
5502A | CDS | 250 | 1000.1 | 1 | °C | °C |  |  |  | 0.72528 | 0 | 0.71398 | 0
5502A | CDS | 1000 | 1400.1 | 1 | °C | °C |  |  |  | 0.73114 | 0 | 0.71993 | 0
5502A | CDS | 1400 | 1767 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
5502A | CDS | 1767 | 1767.1 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
5502A | CDT | -250 | -149.9 | 0.1 | °C | °C |  |  |  | 0.55536 | 0 | 0.55058 | 0
5502A | CDT | -150 | 0.1 | 0.1 | °C | °C |  |  |  | 0.29308 | 0 | 0.28549 | 0
5502A | CDT | 0 | 120.1 | 0.1 | °C | °C |  |  |  | 0.20015 | 0 | 0.19651 | 0
5502A | CDT | 120 | 400 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | CDT | 400 | 400.1 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | CDT | -250 | -149.9 | 1 | °C | °C |  |  |  | 0.8089 | 0 | 0.7957 | 0
5502A | CDT | -150 | 0.1 | 1 | °C | °C |  |  |  | 0.65502 | 0 | 0.64149 | 0
5502A | CDT | 0 | 120.1 | 1 | °C | °C |  |  |  | 0.61188 | 0 | 0.60714 | 0
5502A | CDT | 120 | 400 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0
5502A | CDT | 400 | 400.1 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0
5502A | DCA | 0 | 0.00033001 | 1e-09 | A | µA |  |  |  | 0.024 | 180 | 0.023107 | 175.06
5502A | DCA | 0 | 0.00033001 | 1e-08 | A | µA |  |  |  | 0.024 | 173.53 | 0.023814 | 173.54
5502A | DCA | 0 | 0.00033001 | 1e-07 | A | µA |  |  |  | 0.062453 | 112.75 | 0.062384 | 112.54
5502A | DCA | 0 | 0.00033001 | 1e-06 | A | µA |  |  |  | 0.58048 | 15.83 | 0.58045 | 15.637
5502A | DCA | 0 | 0.00033001 | 1e-05 | A | µA |  |  |  | 5.8001 | 1.6501 | 5.8001 | 1.5718
5502A | DCA | 0 | 0.00033001 | 0.0001 | A | µA |  |  |  | 58 | 0.18852 | 58 | 0.15719
5502A | DCA | 0 | 0.00033001 | 0.001 | A | µA |  |  |  | 580 | 0.028252 | 580 | 0.015719
5502A | DCA | 0.00033 | 0.0033001 | 1e-08 | A | µA |  |  |  | 0.068 | 150 | 0.057598 | 152.89
5502A | DCA | 0.00033 | 0.0033001 | 1e-07 | A | µA |  |  |  | 0.076187 | 148.47 | 0.07332 | 149.02
5502A | DCA | 0.00033 | 0.0033001 | 1e-06 | A | µA |  |  |  | 0.56792 | 73.551 | 0.56576 | 73.312
5502A | DCA | 0.00033 | 0.0033001 | 1e-05 | A | µA |  |  |  | 5.7989 | 9.0557 | 5.7981 | 8.8116
5502A | DCA | 0.00033 | 0.0033001 | 0.0001 | A | µA |  |  |  | 58 | 0.98192 | 58 | 0.88328
5502A | DCA | 0.0033 | 0.033001 | 1e-07 | A | µA |  |  |  | 0.21 | 150 | 0.201 | 145.79
5502A | DCA | 0.0033 | 0.033001 | 1e-06 | A | µA |  |  |  | 0.44201 | 139.58 | 0.43216 | 139.79
5502A | DCA | 0.0033 | 0.033001 | 1e-05 | A | µA |  |  |  | 5.6422 | 61.544 | 5.6369 | 61.47
5502A | DCA | 0.0033 | 0.033001 | 0.0001 | A | µA |  |  |  | 57.983 | 7.2053 | 57.981 | 7.143
5502A | DCA | 0.0033 | 0.033001 | 0.001 | A | µA |  |  |  | 580 | 0.74076 | 580 | 0.71564
5502A | DCA | 0.033 | 0.33001 | 1e-06 | A | µA |  |  |  | 3 | 120 | 2.9414 | 115.33
5502A | DCA | 0.033 | 0.33001 | 1e-05 | A | µA |  |  |  | 5.2853 | 109.48 | 5.266 | 109.51
5502A | DCA | 0.033 | 0.33001 | 0.0001 | A | µA |  |  |  | 56.995 | 42.585 | 56.984 | 42.551
5502A | DCA | 0.033 | 0.33001 | 0.001 | A | µA |  |  |  | 579.89 | 4.761 | 579.88 | 4.7411
5502A | DCA | 0.033 | 0.33001 | 0.01 | A | µA |  |  |  | 5800 | 0.4827 | 5800 | 0.47471
5502A | DCA | 0 | 0.00033001 | 0.01 | A | mA |  |  |  | 5.8 | 6.59e-06 | 5.8 | 1.57e-06
5502A | DCA | 0 | 0.00033001 | 0.1 | A | mA |  |  |  | 58 | 2.16e-06 | 58 | 1.57e-07
5502A | DCA | 0.33 | 3.0001 | 1e-05 | A | mA |  |  |  | 0.052 | 0.44 | 0.051 | 0.43873
5502A | DCA | 0.33 | 3.0001 | 0.0001 | A | mA |  |  |  | 0.060249 | 0.43614 | 0.060204 | 0.43607
5502A | DCA | 0.33 | 3.0001 | 0.001 | A | mA |  |  |  | 0.50421 | 0.32729 | 0.50424 | 0.32696
5502A | DCA | 0.33 | 3.0001 | 0.01 | A | mA |  |  |  | 5.784 | 0.058643 | 5.7841 | 0.058299
5502A | DCA | 0.33 | 3.0001 | 0.1 | A | mA |  |  |  | 57.998 | 0.0060523 | 57.998 | 0.0059106
5502A | DCA | 3 | 11.001 | 0.0001 | A | mA |  |  |  | -0.024 | 0.92 | -0.024619 | 0.91523
5502A | DCA | 3 | 11.001 | 0.001 | A | mA |  |  |  | 0.055271 | 0.90973 | 0.052344 | 0.90974
5502A | DCA | 3 | 11.001 | 0.01 | A | mA |  |  |  | 4.4629 | 0.64947 | 4.4597 | 0.64887
5502A | DCA | 3 | 11.001 | 0.1 | A | mA |  |  |  | 57.764 | 0.1006 | 57.764 | 0.099909
5502A | DCA | 11 | 20.5 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
5502A | DCA | 11 | 20.5 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
5502A | DCA | 11 | 20.5 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
5502A | DCA | 11 | 20.5 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
5502A | DCA | 20.5 | 20.6 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
5502A | DCA | 20.5 | 20.6 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
5502A | DCA | 20.5 | 20.6 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
5502A | DCA | 20.5 | 20.6 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
5502A | DCA 10 turn | 1 | 11.001 | 0.0001 | A | mA |  |  |  | 0.082 | 2 | 0.080875 | 1.9947
5502A | DCA 10 turn | 1 | 11.001 | 0.001 | A | mA |  |  |  | 0.16836 | 1.9875 | 0.16627 | 1.9876
5502A | DCA 10 turn | 1 | 11.001 | 0.01 | A | mA |  |  |  | 4.4751 | 1.6631 | 4.4718 | 1.6632
5502A | DCA 10 turn | 1 | 11.001 | 0.1 | A | mA |  |  |  | 57.372 | 0.40227 | 57.371 | 0.40202
5502A | DCA 10 turn | 1 | 11.001 | 1 | A | mA |  |  |  | 577.32 | 0.041721 | 577.32 | 0.041613
5502A | DCA 10 turn | 11 | 30.001 | 0.001 | A | mA |  |  |  | 1.6 | 2 | 0.12764 | 2.0471
5502A | DCA 10 turn | 11 | 30.001 | 0.01 | A | mA |  |  |  | 1.6 | 2.0127 | 1.1038 | 2.0235
5502A | DCA 10 turn | 11 | 30.001 | 0.1 | A | mA |  |  |  | 49.094 | 1.1796 | 49.066 | 1.1772
5502A | DCA 10 turn | 11 | 30.001 | 1 | A | mA |  |  |  | 576.17 | 0.15058 | 576.16 | 0.14877
5502A | DCA 10 turn | 11 | 30.001 | 10 | A | mA |  |  |  | 5773.4 | 0.015656 | 5773.4 | 0.014925
5502A | DCA 10 turn | 11 | 30.001 | 100 | A | mA |  |  |  | 57736 | 0.0017851 | 57735 | 0.0014925
5502A | DCA 10 turn | 30 | 110.01 | 0.01 | A | mA |  |  |  | 4 | 2.1 | 2.509 | 2.1122
5502A | DCA 10 turn | 30 | 110.01 | 0.1 | A | mA |  |  |  | 29.551 | 1.93 | 29.511 | 1.9297
5502A | DCA 10 turn | 30 | 110.01 | 1 | A | mA |  |  |  | 565.25 | 0.52854 | 565.25 | 0.52743
5502A | DCA 10 turn | 30 | 110.01 | 10 | A | mA |  |  |  | 5772.3 | 0.055475 | 5772.3 | 0.054985
5502A | DCA 10 turn | 30 | 110.01 | 100 | A | mA |  |  |  | 57735 | 0.0056968 | 57735 | 0.005501
5502A | DCA 10 turn | 110 | 205 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
5502A | DCA 10 turn | 110 | 205 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
5502A | DCA 10 turn | 110 | 205 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
5502A | DCA 10 turn | 110 | 205 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
5502A | DCA 10 turn | 110 | 205 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
5502A | DCA 10 turn | 205 | 206 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
5502A | DCA 10 turn | 205 | 206 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
5502A | DCA 10 turn | 205 | 206 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
5502A | DCA 10 turn | 205 | 206 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
5502A | DCA 10 turn | 205 | 206 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
5502A | DCA 20 turn | 20 | 60.01 | 0.001 | A | mA |  |  |  | 3.9 | 2.1 | 0.23928 | 2.0473
5502A | DCA 20 turn | 20 | 60.01 | 0.01 | A | mA |  |  |  | 3.9 | 2.1 | 0.76998 | 2.0407
5502A | DCA 20 turn | 20 | 60.01 | 0.1 | A | mA |  |  |  | 38.447 | 1.626 | 38.404 | 1.6257
5502A | DCA 20 turn | 20 | 60.01 | 1 | A | mA |  |  |  | 573.08 | 0.28848 | 573.07 | 0.28764
5502A | DCA 20 turn | 20 | 60.01 | 10 | A | mA |  |  |  | 5773.1 | 0.029468 | 5773.1 | 0.02912
5502A | DCA 20 turn | 20 | 60.01 | 100 | A | mA |  |  |  | 57735 | 0.0030514 | 57735 | 0.0029124
5502A | DCA 20 turn | 60 | 220.01 | 0.001 | A | mA |  |  |  | 19 | 2.2 | 4.5221 | 2.114
5502A | DCA 20 turn | 60 | 220.01 | 0.01 | A | mA |  |  |  | 19 | 2.2 | 19.726 | 2.0608
5502A | DCA 20 turn | 60 | 220.01 | 0.1 | A | mA |  |  |  | 535.04 | 0.95226 | 535.04 | 0.95071
5502A | DCA 20 turn | 60 | 220.01 | 1 | A | mA |  |  |  | 5768.4 | 0.11068 | 5768.5 | 0.10982
5502A | DCA 20 turn | 60 | 220.01 | 10 | A | mA |  |  |  | 57735 | 0.011347 | 57735 | 0.011002
5502A | DCA 20 turn | 60 | 220.01 | 100 | A | mA |  |  |  | 577360 | 0.0012382 | 577350 | 0.0011002
5502A | DCA 20 turn | 220 | 410 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
5502A | DCA 20 turn | 220 | 410 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
5502A | DCA 20 turn | 220 | 410 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
5502A | DCA 20 turn | 220 | 410 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
5502A | DCA 20 turn | 220 | 410 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
5502A | DCA 20 turn | 410 | 411 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
5502A | DCA 20 turn | 410 | 411 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
5502A | DCA 20 turn | 410 | 411 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
5502A | DCA 20 turn | 410 | 411 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
5502A | DCA 20 turn | 410 | 411 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
5502A | DCA 50 turn | 55 | 150.01 | 0.001 | A | mA |  |  |  | 7.7 | 2 | 0.58265 | 2.0474
5502A | DCA 50 turn | 55 | 150.01 | 0.01 | A | mA |  |  |  | 7.7 | 2.0006 | 0.78384 | 2.0464
5502A | DCA 50 turn | 55 | 150.01 | 0.1 | A | mA |  |  |  | 19.451 | 1.9578 | 19.379 | 1.9579
5502A | DCA 50 turn | 55 | 150.01 | 1 | A | mA |  |  |  | 550.23 | 0.69408 | 550.2 | 0.69349
5502A | DCA 50 turn | 55 | 150.01 | 10 | A | mA |  |  |  | 5770.6 | 0.074862 | 5770.6 | 0.074565
5502A | DCA 50 turn | 55 | 150.01 | 100 | A | mA |  |  |  | 57735 | 0.0075815 | 57732 | 0.0074625
5502A | DCA 50 turn | 150 | 550.01 | 0.01 | A | mA |  |  |  | 19 | 2.2 | -17.053 | 2.1654
5502A | DCA 50 turn | 150 | 550.01 | 0.1 | A | mA |  |  |  | 19 | 2.2 | -10.261 | 2.1556
5502A | DCA 50 turn | 150 | 550.01 | 1 | A | mA |  |  |  | 409.49 | 1.6357 | 409.01 | 1.6349
5502A | DCA 50 turn | 150 | 550.01 | 10 | A | mA |  |  |  | 5740.6 | 0.27618 | 5740.5 | 0.27485
5502A | DCA 50 turn | 150 | 550.01 | 100 | A | mA |  |  |  | 57732 | 0.028329 | 57732 | 0.027783
5502A | DCA 50 turn | 550 | 1025 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
5502A | DCA 50 turn | 550 | 1025 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
5502A | DCA 50 turn | 550 | 1025 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
5502A | DCA 50 turn | 550 | 1025 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
5502A | DCA 50 turn | 550 | 1025 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
5502A | DCA 50 turn | 1025 | 1026 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
5502A | DCA 50 turn | 1025 | 1026 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
5502A | DCA 50 turn | 1025 | 1026 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
5502A | DCA 50 turn | 1025 | 1026 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
5502A | DCA 50 turn | 1025 | 1026 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
5502A | DCV | 0 | 0.33 | 1e-07 | V | µV |  |  |  | 3.6 | 69 | 3.4727 | 69.314
5502A | DCV | 0 | 0.33 | 1e-06 | V | µV |  |  |  | 3.6 | 69.151 | 3.5196 | 69.191
5502A | DCV | 0 | 0.33 | 1e-05 | V | µV |  |  |  | 6.7738 | 61.287 | 6.7344 | 61.342
5502A | DCV | 0 | 0.33 | 0.0001 | V | µV |  |  |  | 58.11 | 17.06 | 58.091 | 17.007
5502A | DCV | 0 | 0.33 | 0.001 | V | µV |  |  |  | 580.02 | 1.8105 | 580.01 | 1.785
5502A | DCV | 0 | 0.33 | 0.01 | V | µV |  |  |  | 5800 | 0.1888 | 5800 | 0.17859
5502A | DCV | 0 | 0.33 | 0.1 | V | mV |  |  |  | 58 | 2.1943e-05 | 58 | 1.7859e-05
5502A | DCV | 0 | 0.33 | 1 | V | mV |  |  |  | 580 | 3.42e-06 | 580 | 1.79e-06
5502A | DCV | 0 | 0.33 | 10 | V | mV |  |  |  | 5800 | 8.32e-07 | 5800 | 1.79e-07
5502A | DCV | 0.33 | 3.3 | 1e-06 | V | µV |  |  |  | 5.9 | 58 | 5.7708 | 57.787
5502A | DCV | 0.33 | 3.3 | 1e-05 | V | µV |  |  |  | 6.8066 | 57.529 | 6.4957 | 57.593
5502A | DCV | 0.33 | 3.3 | 0.0001 | V | µV |  |  |  | 47.82 | 47.701 | 47.342 | 47.73
5502A | DCV | 0.33 | 3.3 | 0.001 | V | µV |  |  |  | 577.17 | 10.824 | 576.99 | 10.721
5502A | DCV | 0.33 | 3.3 | 0.01 | V | µV |  |  |  | 5799.8 | 1.1473 | 5799.7 | 1.1022
5502A | DCV | 0.33 | 3.3 | 0.1 | V | µV |  |  |  | 58000 | 0.12833 | 5799.8 | 0.11025
5502A | DCV | 3 | 33.0001 | 1e-05 | V | mV |  |  |  | 0.059 | 0.058 | 0.058348 | 0.057919
5502A | DCV | 3 | 33.0001 | 0.0001 | V | mV |  |  |  | 0.067758 | 0.057661 | 0.064652 | 0.057725
5502A | DCV | 3 | 33.0001 | 0.001 | V | mV |  |  |  | 0.47784 | 0.047829 | 0.47304 | 0.047859
5502A | DCV | 3 | 33.0001 | 0.01 | V | mV |  |  |  | 5.7715 | 0.010868 | 5.7698 | 0.010764
5502A | DCV | 3 | 33.0001 | 0.1 | V | mV |  |  |  | 57.998 | 0.0011519 | 57.997 | 0.0011068
5502A | DCV | 3 | 33.0001 | 1 | V | mV |  |  |  | 580 | 0.00012879 | 580 | 0.00011071
5502A | DCV | 33 | 330.001 | 0.0001 | V | mV |  |  |  | 0.59 | 0.064 | 0.57834 | 0.063644
5502A | DCV | 33 | 330.001 | 0.001 | V | mV |  |  |  | 0.67471 | 0.063403 | 0.64578 | 0.063463
5502A | DCV | 33 | 330.001 | 0.01 | V | mV |  |  |  | 4.6629 | 0.053693 | 4.6153 | 0.053731
5502A | DCV | 33 | 330.001 | 0.1 | V | mV |  |  |  | 57.654 | 0.012974 | 57.637 | 0.012872
5502A | DCV | 33 | 330.001 | 1 | V | mV |  |  |  | 579.97 | 0.0013757 | 579.96 | 0.0013305
5502A | DCV | 33 | 330.001 | 10 | V | mV |  |  |  | 5800 | 0.00015118 | 5800 | 0.0001331
5502A | DCV | 330 | 1000.01 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
5502A | DCV | 330 | 1000.01 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
5502A | DCV | 330 | 1000.01 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
5502A | DCV | 330 | 1000.01 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
5502A | DCV | 330 | 1000.01 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
5502A | DCV | 1000 | 1001 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
5502A | DCV | 1000 | 1001 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
5502A | DCV | 1000 | 1001 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
5502A | DCV | 1000 | 1001 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
5502A | DCV | 1000 | 1001 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
5502A | DCW | 0 | 0.000109 | 1e-10 | WDC | WDC |  |  |  | 0 | 0.0036 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-09 | WDC | WDC |  |  |  | 0 | 0.0036 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-08 | WDC | WDC |  |  |  | 0 | 0.0068024 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.05811 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.58001 | 0 | 0
5502A | DCW | 0 | 0.000109 | 1e-05 | WDC | WDC |  |  |  | 0 | 5.8 | 0 | 0
5502A | DCW | 0 | 0.000109 | 0.0001 | WDC | WDC |  |  |  | 0 | 58 | 0 | 0
5502A | DCW | 0 | 0.000109 | 0.001 | WDC | WDC |  |  |  | 0 | 580 | 0 | 0
5502A | DCW | 0 | 0.000109 | 0.01 | WDC | WDC |  |  |  | 0 | 5800 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-08 | WDC | WDC |  |  |  | 0 | 0.00035 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.00063401 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.0053337 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.053213 | 0 | 0
5502A | DCW | 0.000109 | 1.09e-06 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.53211 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.00024 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.00058093 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.0053277 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.053212 | 0 | 0
5502A | DCW | 0.00109 | 0.0109 | 0.001 | WDC | WDC |  |  |  | 0 | 0.53211 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.00019 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.00056395 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.0053308 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 0.001 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 0.0109 | 0.109 | 0.01 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.00057856 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.001 | WDC | WDC |  |  |  | 0 | 0.0053324 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.01 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 0.109 | 1.0901 | 0.1 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00057825 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.0053322 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 1.09 | 10.9001 | 1 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.0053323 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
5502A | DCW | 109 | 330.0001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00015 | 0 | 0
5502A | DCW | 109 | 330.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.00015255 | 0 | 0
5502A | DCW | 109 | 330.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.00055202 | 0 | 0
5502A | DCW | 109 | 330.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.0053282 | 0 | 0
5502A | DCW | 109 | 330.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.05326 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.0006 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.00062246 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.0018565 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.017586 | 0 | 0
5502A | DCW | 330 | 3000.0001 | 100 | WDC | WDC |  |  |  | 0 | 0.17576 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 10 | WDC | WDC |  |  |  | 0 | 0.00091 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 100 | WDC | WDC |  |  |  | 0 | 0.00091 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 1000 | WDC | WDC |  |  |  | 0 | 0.00092998 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 10000 | WDC | WDC |  |  |  | 0 | 0.0021367 | 0 | 0
5502A | DCW | 3000 | 11000.1 | 100000 | WDC | WDC |  |  |  | 0 | 0.019355 | 0 | 0
5502A | DCW | 11000 | 24500 | 10 | WDC | WDC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 11000 | 24500 | 100 | WDC | WDC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 11000 | 24500 | 1000 | WDC | WDC |  |  |  | 0 | 0.0014429 | 0 | 0
5502A | DCW | 11000 | 24500 | 10000 | WDC | WDC |  |  |  | 0 | 0.0054411 | 0 | 0
5502A | DCW | 11000 | 24500 | 100000 | WDC | WDC |  |  |  | 0 | 0.052745 | 0 | 0
5502A | DCW | 24500 | 24501 | 10 | WAC | WAC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 24500 | 24501 | 100 | WAC | WAC |  |  |  | 0 | 0.0014 | 0 | 0
5502A | DCW | 24500 | 24501 | 1000 | WAC | WAC |  |  |  | 0 | 0.0014429 | 0 | 0
5502A | DCW | 24500 | 24501 | 10000 | WAC | WAC |  |  |  | 0 | 0.0054411 | 0 | 0
5502A | DCW | 24500 | 24501 | 100000 | WAC | WAC |  |  |  | 0 | 0.052745 | 0 | 0
5502A | DCW | 10.9 | 109.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.00057887 | 0 | 0
5502A | FDE | -418 | -147.9 | 0.1 | °F | °F |  |  |  | 1.0462 | 0 | 1.0433 | 0
5502A | FDE | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.3466 | 0 | 0.34484 | 0
5502A | FDE | -13 | 662.1 | 0.1 | °F | °F |  |  |  | 0.30437 | 0 | 0.30237 | 0
5502A | FDE | 662 | 1202.1 | 0.1 | °F | °F |  |  |  | 0.34352 | 0 | 0.34251 | 0
5502A | FDE | 1202 | 1832 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
5502A | FDE | 1832 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
5502A | FDE | -418 | -147.9 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
5502A | FDE | -148 | -12.9 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
5502A | FDE | -13 | 662.1 | 1 | °F | °F |  |  |  | 0.6529 | 0 | 0.64917 | 0
5502A | FDE | 662 | 1202.1 | 1 | °F | °F |  |  |  | 0.67089 | 0 | 0.66882 | 0
5502A | FDE | 1202 | 1832 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
5502A | FDE | 1832 | 1832.1 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
5502A | FDJ | -346 | -147.9 | 0.1 | °F | °F |  |  |  | 0.57727 | 0 | 0.56758 | 0
5502A | FDJ | -148 | -21.9 | 0.1 | °F | °F |  |  |  | 0.35408 | 0 | 0.34263 | 0
5502A | FDJ | -22 | 302.1 | 0.1 | °F | °F |  |  |  | 0.31531 | 0 | 0.3024 | 0
5502A | FDJ | 302 | 1400.1 | 0.1 | °F | °F |  |  |  | 0.36584 | 0 | 0.36275 | 0
5502A | FDJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
5502A | FDJ | 2192 | 2192.1 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
5502A | FDJ | -346 | -147.9 | 1 | °F | °F |  |  |  | 0.83458 | 0 | 0.80755 | 0
5502A | FDJ | -148 | -21.9 | 1 | °F | °F |  |  |  | 0.69231 | 0 | 0.66888 | 0
5502A | FDJ | -22 | 302.1 | 1 | °F | °F |  |  |  | 0.67331 | 0 | 0.64919 | 0
5502A | FDJ | 302 | 1400.1 | 1 | °F | °F |  |  |  | 0.686 | 0 | 0.6794 | 0
5502A | FDJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
5502A | FDJ | 2192 | 2192.1 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
5502A | FDK | -328 | -147.9 | 0.1 | °F | °F |  |  |  | 0.70186 | 0 | 0.69128 | 0
5502A | FDK | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.40198 | 0 | 0.38321 | 0
5502A | FDK | -13 | 248.1 | 0.1 | °F | °F |  |  |  | 0.35702 | 0 | 0.34265 | 0
5502A | FDK | 248 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.55155 | 0 | 0.54665 | 0
5502A | FDK | 1832 | 2500 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
5502A | FDK | 2500 | 2500.1 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
5502A | FDK | -328 | -147.9 | 1 | °F | °F |  |  |  | 0.93103 | 0 | 0.89881 | 0
5502A | FDK | -148 | -12.9 | 1 | °F | °F |  |  |  | 0.73198 | 0 | 0.69054 | 0
5502A | FDK | -13 | 248.1 | 1 | °F | °F |  |  |  | 0.69831 | 0 | 0.66889 | 0
5502A | FDK | 248 | 1832.1 | 1 | °F | °F |  |  |  | 0.80643 | 0 | 0.79299 | 0
5502A | FDK | 1832 | 2500 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
5502A | FDK | 2500 | 2500.1 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
5502A | FDN | -328 | -147.9 | 0.1 | °F | °F |  |  |  | 0.84492 | 0 | 0.83553 | 0
5502A | FDN | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.47674 | 0 | 0.46496 | 0
5502A | FDN | -13 | 248.1 | 0.1 | °F | °F |  |  |  | 0.41715 | 0 | 0.40364 | 0
5502A | FDN | 248 | 770.1 | 0.1 | °F | °F |  |  |  | 0.3915 | 0 | 0.38322 | 0
5502A | FDN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
5502A | FDN | 2372 | 2372.1 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
5502A | FDN | -328 | -147.9 | 1 | °F | °F |  |  |  | 1.0446 | 0 | 1.014 | 0
5502A | FDN | -148 | -12.9 | 1 | °F | °F |  |  |  | 0.76848 | 0 | 0.73904 | 0
5502A | FDN | -13 | 248.1 | 1 | °F | °F |  |  |  | 0.73301 | 0 | 0.70208 | 0
5502A | FDN | 248 | 770.1 | 1 | °F | °F |  |  |  | 0.70888 | 0 | 0.69055 | 0
5502A | FDN | 770 | 2372 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
5502A | FDN | 2372 | 2372.1 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
5502A | FDPT100 | -328 | 32.1 | 0.1 | °F | °F |  |  |  | 0.21928 | 0 | 0.19044 | 0
5502A | FDPT100 | 32 | 212.1 | 0.1 | °F | °F |  |  |  | 0.28537 | 0 | 0.26385 | 0
5502A | FDPT100 | 212 | 572.1 | 0.1 | °F | °F |  |  |  | 0.28537 | 0 | 0.26385 | 0
5502A | FDPT100 | 572 | 752.1 | 0.1 | °F | °F |  |  |  | 0.39012 | 0 | 0.37466 | 0
5502A | FDPT100 | 752 | 1166.1 | 0.1 | °F | °F |  |  |  | 0.45987 | 0 | 0.44683 | 0
5502A | FDPT100 | 1166 | 1472 | 0.1 | °F | °F |  |  |  | 0.77811 | 0 | 0.77048 | 0
5502A | FDPT100 | 1472 | 1472.1 | 0.1 | °F | °F |  |  |  | 0.77811 | 0 | 0.77048 | 0
5502A | FDPT100 | -328 | 32.1 | 1 | °F | °F |  |  |  | 0.60123 | 0 | 0.58897 | 0
5502A | FDPT100 | 32 | 212.1 | 1 | °F | °F |  |  |  | 0.61101 | 0 | 0.59895 | 0
5502A | FDPT100 | 212 | 572.1 | 1 | °F | °F |  |  |  | 0.61101 | 0 | 0.59895 | 0
5502A | FDPT100 | 572 | 752.1 | 1 | °F | °F |  |  |  | 0.63118 | 0 | 0.61951 | 0
5502A | FDPT100 | 752 | 1166.1 | 1 | °F | °F |  |  |  | 0.64673 | 0 | 0.63534 | 0
5502A | FDPT100 | 1166 | 1472 | 1 | °F | °F |  |  |  | 0.739 | 0 | 0.72906 | 0
5502A | FDPT100 | 1472 | 1472.1 | 1 | °F | °F |  |  |  | 0.739 | 0 | 0.72906 | 0
5502A | FDPT1000 | -328 | 32.1 | 0.1 | °F | °F |  |  |  | 0.16401 | 0 | 0.1228 | 0
5502A | FDPT1000 | 32 | 212.1 | 0.1 | °F | °F |  |  |  | 0.19201 | 0 | 0.15827 | 0
5502A | FDPT1000 | 212 | 500.1 | 0.1 | °F | °F |  |  |  | 0.19201 | 0 | 0.15827 | 0
5502A | FDPT1000 | 500 | 572.1 | 0.1 | °F | °F |  |  |  | 0.2619 | 0 | 0.23826 | 0
5502A | FDPT1000 | 572 | 752.1 | 0.1 | °F | °F |  |  |  | 0.29342 | 0 | 0.27254 | 0
5502A | FDPT1000 | 752 | 1112.1 | 0.1 | °F | °F |  |  |  | 0.29725 | 0 | 0.27666 | 0
5502A | FDPT1000 | 1112 | 1166 | 0.1 | °F | °F |  |  |  | 0.84699 | 0 | 0.83998 | 0
5502A | FDPT1000 | 1166 | 1166.1 | 0.1 | °F | °F |  |  |  | 0.84699 | 0 | 0.83998 | 0
5502A | FDPT1000 | -328 | 32.1 | 1 | °F | °F |  |  |  | 0.59551 | 0 | 0.58313 | 0
5502A | FDPT1000 | 212 | 500.1 | 1 | °F | °F |  |  |  | 0.59861 | 0 | 0.58629 | 0
5502A | FDPT1000 | 500 | 572.1 | 1 | °F | °F |  |  |  | 0.61003 | 0 | 0.59795 | 0
5502A | FDPT1000 | 572 | 752.1 | 1 | °F | °F |  |  |  | 0.61481 | 0 | 0.60282 | 0
5502A | FDPT1000 | 752 | 1112.1 | 1 | °F | °F |  |  |  | 0.61665 | 0 | 0.6047 | 0
5502A | FDPT1000 | 1112 | 1166 | 1 | °F | °F |  |  |  | 1.0262 | 0 | 1.0191 | 0
5502A | FDPT1000 | 1166 | 1166.1 | 1 | °F | °F |  |  |  | 1.0262 | 0 | 1.0191 | 0
5502A | FDR | 32 | 482.1 | 0.1 | °F | °F |  |  |  | 1.1966 | 0 | 1.19 | 0
5502A | FDR | 482 | 752.1 | 0.1 | °F | °F |  |  |  | 0.74217 | 0 | 0.73324 | 0
5502A | FDR | 752 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.70146 | 0 | 0.69201 | 0
5502A | FDR | 1832 | 3213 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
5502A | FDR | 3213 | 3213.1 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
5502A | FDR | 32 | 482.1 | 1 | °F | °F |  |  |  | 1.3448 | 0 | 1.3214 | 0
5502A | FDR | 482 | 752.1 | 1 | °F | °F |  |  |  | 0.95934 | 0 | 0.93147 | 0
5502A | FDR | 752 | 1832.1 | 1 | °F | °F |  |  |  | 0.92821 | 0 | 0.89938 | 0
5502A | FDR | 1832 | 3213 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
5502A | FDR | 3213 | 3213.1 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
5502A | FDS | 32 | 482.1 | 0.1 | °F | °F |  |  |  | 0.99102 | 0 | 0.98329 | 0
5502A | FDS | 482 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.76311 | 0 | 0.75443 | 0
5502A | FDS | 1832 | 2552.1 | 0.1 | °F | °F |  |  |  | 0.78274 | 0 | 0.77428 | 0
5502A | FDS | 2552 | 3213 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
5502A | FDS | 3213 | 3213.1 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
5502A | FDS | 32 | 482.1 | 1 | °F | °F |  |  |  | 1.1653 | 0 | 1.1388 | 0
5502A | FDS | 482 | 1832.1 | 1 | °F | °F |  |  |  | 0.97706 | 0 | 0.94824 | 0
5502A | FDS | 1832 | 2552.1 | 1 | °F | °F |  |  |  | 0.99106 | 0 | 0.96411 | 0
5502A | FDS | 2552 | 3213 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
5502A | FDS | 3213 | 3213.1 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
5502A | FDT | -418 | -237.9 | 0.1 | °F | °F |  |  |  | 0.55536 | 0 | 0.55058 | 0
5502A | FDT | -238 | 32.1 | 0.1 | °F | °F |  |  |  | 0.29308 | 0 | 0.28549 | 0
5502A | FDT | 32 | 248.1 | 0.1 | °F | °F |  |  |  | 0.20015 | 0 | 0.19651 | 0
5502A | FDT | 248 | 752 | 0.1 | °F | °F |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | FDT | 752 | 752.1 | 0.1 | °F | °F |  |  |  | 0.17895 | 0 | 0.17487 | 0
5502A | FDT | -418 | -237.9 | 1 | °F | °F |  |  |  | 0.8089 | 0 | 0.7957 | 0
5502A | FDT | -238 | 32.1 | 1 | °F | °F |  |  |  | 0.65502 | 0 | 0.64149 | 0
5502A | FDT | 32 | 248.1 | 1 | °F | °F |  |  |  | 0.61188 | 0 | 0.60714 | 0
5502A | FDT | 248 | 752 | 1 | °F | °F |  |  |  | 0.60528 | 0 | 0.60048 | 0
5502A | FDT | 752 | 752.1 | 1 | °F | °F |  |  |  | 0.60528 | 0 | 0.60048 | 0
5502A | Frequency | 0 | 120.001 | 0.001 | Hz | Hz |  |  |  | 0.006 | 1.3e-05 | 0.0012889 | 2.8049e-05
5502A | Frequency | 0 | 120.001 | 0.01 | Hz | Hz |  |  |  | 0.0060576 | 1.5807e-05 | 0.005881 | 1.2606e-05
5502A | Frequency | 0 | 120.001 | 0.1 | Hz | Hz |  |  |  | 0.057819 | 3.3129e-06 | 0.057746 | 1.4481e-06
5502A | Frequency | 0 | 120.001 | 1 | Hz | Hz |  |  |  | 0.57739 | 8.9636e-07 | 0.57736 | 1.4506e-07
5502A | Frequency | 0 | 120.001 | 10 | Hz | Hz |  |  |  | 5.7736 | 3.1508e-07 | 5.7736 | 1.4506e-08
5502A | Frequency | 0 | 120.001 | 100 | Hz | Hz |  |  |  | 57.736 | 1.2168e-07 | 57.736 | 1.45e-09
5502A | Frequency | 0 | 120.001 | 1000 | Hz | Hz |  |  |  | 577.36 | 4.8236e-08 | 577.36 | 1.45e-10
5502A | Frequency | 120 | 1200.01 | 0.01 | Hz | Hz |  |  |  | 0.058 | 9.5e-06 | 0.0041865 | 2.6727e-05
5502A | Frequency | 120 | 1200.01 | 0.1 | Hz | Hz |  |  |  | 0.058519 | 1.0244e-05 | 0.056808 | 9.2702e-06
5502A | Frequency | 120 | 1200.01 | 1 | Hz | Hz |  |  |  | 0.57795 | 1.5973e-06 | 0.57725 | 1.0094e-06
5502A | Frequency | 120 | 1200.01 | 10 | Hz | Hz |  |  |  | 5.7739 | 3.3757e-07 | 5.7735 | 1.0104e-07
5502A | Frequency | 1200 | 12001 | 0.1 | Hz | Hz |  |  |  | 0.58 | 9.1e-06 | 0.036333 | 2.6333e-05
5502A | Frequency | 1200 | 12001 | 0.1 | Hz | Hz |  |  |  | 0.58497 | 9.8286e-06 | 0.56786 | 8.8368e-06
5502A | Frequency | 1200 | 12001 | 1 | Hz | Hz |  |  |  | 5.7795 | 1.5456e-06 | 5.7725 | 9.5754e-07
5502A | Frequency | 1200 | 12001 | 10 | Hz | Hz |  |  |  | 57.739 | 3.3237e-07 | 57.735 | 9.584e-08
5502A | Frequency | 12000 | 120010 | 1 | Hz | Hz |  |  |  | 19 | 2.8e-05 | 17.33 | 2.8854e-05
5502A | Frequency | 12000 | 120010 | 10 | Hz | Hz |  |  |  | 19 | 2.8e-05 | 18.255 | 2.7641e-05
5502A | Frequency | 12000 | 120010 | 100 | Hz | Hz |  |  |  | 60.592 | 1.1287e-05 | 60.269 | 9.1178e-06
5502A | Frequency | 12000 | 120010 | 1000 | Hz | Hz |  |  |  | 577.75 | 1.9059e-06 | 577.61 | 9.6075e-07
5502A | Frequency | 120000 | 1200010 | 10 | Hz | Hz |  |  |  | 60 | 1.6e-05 | 18.159 | 2.8435e-05
5502A | Frequency | 120000 | 1200010 | 100 | Hz | Hz |  |  |  | 60.876 | 1.8461e-05 | 59.55 | 1.5104e-05
5502A | Frequency | 120000 | 1200010 | 1000 | Hz | Hz |  |  |  | 578.01 | 3.8671e-06 | 577.51 | 1.8144e-06
5502A | Frequency | 120000 | 1200010 | 10000 | Hz | Hz |  |  |  | 5773.8 | 1.0097e-06 | 5773.6 | 1.8186e-07
5502A | Frequency | 1200000 | 2000000 | 100 | Hz | kHz |  |  |  | 0.43 | 0.00013 | -0.35824 | 0.00036326
5502A | Frequency | 1200000 | 2000000 | 1000 | Hz | kHz |  |  |  | 0.43 | 0.00014137 | 0.42566 | 0.00012836
5502A | Frequency | 1200000 | 2000000 | 10000 | Hz | kHz |  |  |  | 5.7549 | 2.1951e-05 | 5.757 | 1.4015e-05
5502A | Frequency | 1200000 | 2000000 | 100000 | Hz | kHz |  |  |  | 57.733 | 4.5961e-06 | 57.734 | 1.4029e-06
5502A | Frequency | 2000000 | 2001000 | 100 | Hz | kHz |  |  |  | 0.43 | 0.00013 | -0.35824 | 0.00036326
5502A | Frequency | 2000000 | 2001000 | 1000 | Hz | kHz |  |  |  | 0.43 | 0.00014137 | 0.42566 | 0.00012836
5502A | Frequency | 2000000 | 2001000 | 10000 | Hz | kHz |  |  |  | 5.7549 | 2.1951e-05 | 5.757 | 1.4015e-05
5502A | Frequency | 2000000 | 2001000 | 100000 | Hz | kHz |  |  |  | 57.733 | 4.5961e-06 | 57.734 | 1.4029e-06
5502A | Resistance | 0 | 11.001 | 10000000 | Ω | mΩ |  |  |  | 5800000000 | 0 | 5800000000 | 0
5502A | Resistance | 0 | 11.001 | 1000000 | Ω | mΩ |  |  |  | 580000000 | 9.5149e-07 | 580000000 | 0
5502A | Resistance | 0 | 11.001 | 0.001 | Ω | mΩ |  |  |  | 12 | 0.14 | 11.576 | 0.1383
5502A | Resistance | 0 | 11.001 | 0.01 | Ω | mΩ |  |  |  | 12.939 | 0.12581 | 12.935 | 0.12525
5502A | Resistance | 0 | 11.001 | 0.1 | Ω | mΩ |  |  |  | 59.144 | 0.029356 | 59.141 | 0.028788
5502A | Resistance | 0 | 11.001 | 1 | Ω | mΩ |  |  |  | 580.12 | 0.003176 | 580.12 | 0.0029427
5502A | Resistance | 0 | 11.001 | 10 | Ω | mΩ |  |  |  | 5800 | 0.00038767 | 5800 | 0.00029434
5502A | Resistance | 0 | 11.001 | 100 | Ω | mΩ |  |  |  | 58000 | 6.6765e-05 | 58000 | 2.9434e-05
5502A | Resistance | 0 | 11.001 | 1000 | Ω | mΩ |  |  |  | 580000 | 1.7876e-05 | 580000 | 2.9434e-06
5502A | Resistance | 0 | 11.001 | 10000 | Ω | mΩ |  |  |  | 5800000 | 6.2674e-06 | 5800000 | 2.9435e-07
5502A | Resistance | 0 | 11.001 | 100000 | Ω | mΩ |  |  |  | 58000000 | 2.4177e-06 | 58000000 | 2.9392e-08
5502A | Resistance | 11 | 33.001 | 0.001 | Ω | mΩ |  |  |  | 18 | 0.14 | 17.339 | 0.13885
5502A | Resistance | 11 | 33.001 | 0.01 | Ω | mΩ |  |  |  | 18.263 | 0.13372 | 18.258 | 0.13357
5502A | Resistance | 11 | 33.001 | 0.1 | Ω | mΩ |  |  |  | 60.487 | 0.046304 | 60.481 | 0.04604
5502A | Resistance | 11 | 33.001 | 1 | Ω | mΩ |  |  |  | 580.26 | 0.0049938 | 580.25 | 0.0048783
5502A | Resistance | 33 | 110.01 | 0.001 | Ω | mΩ |  |  |  | 18 | 0.1 | 17.342 | 0.10406
5502A | Resistance | 33 | 110.01 | 0.01 | Ω | mΩ |  |  |  | 18.224 | 0.10151 | 18.221 | 0.10128
5502A | Resistance | 33 | 110.01 | 0.1 | Ω | mΩ |  |  |  | 60.259 | 0.041284 | 60.259 | 0.04082
5502A | Resistance | 33 | 110.01 | 1 | Ω | mΩ |  |  |  | 580.23 | 0.0046533 | 580.23 | 0.0044422
5502A | Resistance | 110 | 330.01 | 0.001 | Ω | mΩ |  |  |  | 25 | 0.3 | 24.812 | 0.29687
5502A | Resistance | 110 | 330.01 | 0.01 | Ω | mΩ |  |  |  | 25.358 | 0.29479 | 25.35 | 0.29441
5502A | Resistance | 110 | 330.01 | 0.1 | Ω | mΩ |  |  |  | 61.466 | 0.18435 | 61.46 | 0.18305
5502A | Resistance | 110 | 330.01 | 1 | Ω | mΩ |  |  |  | 580.25 | 0.024237 | 580.25 | 0.023449
5502A | Resistance | 330 | 1100.01 | 0.01 | Ω | mΩ |  |  |  | 28 | 0.1 | 23.818 | 0.10366
5502A | Resistance | 330 | 1100.01 | 0.1 | Ω | mΩ |  |  |  | 52.968 | 0.087892 | 52.867 | 0.087789
5502A | Resistance | 330 | 1100.01 | 1 | Ω | mΩ |  |  |  | 577.2 | 0.017401 | 577.18 | 0.017221
5502A | Resistance | 330 | 1100.01 | 10 | Ω | mΩ |  |  |  | 5799.7 | 0.0018248 | 5799.7 | 0.00175
5502A | Resistance | 1100 | 3300.1 | 0.01 | Ω | Ω |  |  |  | 0.25 | 0.0001 | 0.23096 | 0.00010411
5502A | Resistance | 1100 | 3300.1 | 0.1 | Ω | Ω |  |  |  | 0.25 | 0.00010329 | 0.23666 | 0.00010325
5502A | Resistance | 1100 | 3300.1 | 1 | Ω | Ω |  |  |  | 0.60487 | 6.4429e-05 | 0.60441 | 6.4221e-05
5502A | Resistance | 1100 | 3300.1 | 10 | Ω | Ω |  |  |  | 5.8014 | 8.3761e-06 | 5.8012 | 8.2294e-06
5502A | Resistance | 3300 | 11001 | 0.1 | Ω | Ω |  |  |  | 0.17 | 0.0001 | 0.12474 | 0.00010353
5502A | Resistance | 3300 | 11001 | 1 | Ω | Ω |  |  |  | 0.46536 | 8.4151e-05 | 0.46464 | 8.4044e-05
5502A | Resistance | 3300 | 11001 | 10 | Ω | Ω |  |  |  | 5.7682 | 1.5398e-05 | 5.7682 | 1.5238e-05
5502A | Resistance | 3300 | 11001 | 100 | Ω | Ω |  |  |  | 57.997 | 1.6098e-06 | 57.997 | 1.5438e-06
5502A | Resistance | 11000 | 33001 | 0.1 | Ω | Ω |  |  |  | 1.3 | 0.0001 | 1.1561 | 0.00010412
5502A | Resistance | 11000 | 33001 | 1 | Ω | Ω |  |  |  | 1.3 | 0.00010253 | 1.2442 | 0.00010252
5502A | Resistance | 11000 | 33001 | 10 | Ω | Ω |  |  |  | 5.6614 | 5.2695e-05 | 5.6605 | 5.2621e-05
5502A | Resistance | 11000 | 33001 | 100 | Ω | Ω |  |  |  | 57.978 | 6.2175e-06 | 57.978 | 6.1738e-06
5502A | Resistance | 33000 | 110010 | 1 | Ω | Ω |  |  |  | 1.3 | 0.00013 | 1.2347 | 0.00012665
5502A | Resistance | 33000 | 110010 | 10 | Ω | Ω |  |  |  | 4.353 | 0.00010805 | 4.3501 | 0.00010797
5502A | Resistance | 33000 | 110010 | 100 | Ω | Ω |  |  |  | 57.521 | 2.2173e-05 | 57.522 | 2.2051e-05
5502A | Resistance | 33000 | 110010 | 1000 | Ω | Ω |  |  |  | 579.95 | 2.2971e-06 | 579.95 | 2.2465e-06
5502A | Resistance | 110000 | 330010 | 1 | Ω | Ω |  |  |  | 12 | 0.00014 | 11.565 | 0.0001393
5502A | Resistance | 110000 | 330010 | 10 | Ω | Ω |  |  |  | 12.289 | 0.00013784 | 12.284 | 0.00013782
5502A | Resistance | 110000 | 330010 | 100 | Ω | Ω |  |  |  | 55.026 | 8.0748e-05 | 55.024 | 8.0667e-05
5502A | Resistance | 110000 | 330010 | 1000 | Ω | Ω |  |  |  | 579.51 | 1.0142e-05 | 579.51 | 1.0091e-05
5502A | Resistance | 330000 | 1100010 | 10 | Ω | Ω |  |  |  | 16 | 0.00017 | 12.341 | 0.00017274
5502A | Resistance | 330000 | 1100010 | 100 | Ω | Ω |  |  |  | 38.744 | 0.00015613 | 38.702 | 0.00015611
5502A | Resistance | 330000 | 1100010 | 1000 | Ω | Ω |  |  |  | 571.18 | 3.9241e-05 | 571.18 | 3.9165e-05
5502A | Resistance | 330000 | 1100010 | 10000 | Ω | Ω |  |  |  | 5799.1 | 4.0776e-06 | 5799.1 | 4.0446e-06
5502A | Resistance | 1100000 | 3301000 | 10 | Ω | kΩ |  |  |  | 0.19 | 1.7e-07 | 0.17339 | 1.7318e-07
5502A | Resistance | 1100000 | 3301000 | 100 | Ω | kΩ |  |  |  | 0.19 | 1.7215e-07 | 0.1791 | 1.7213e-07
5502A | Resistance | 1100000 | 3301000 | 1000 | Ω | kΩ |  |  |  | 0.55516 | 1.1804e-07 | 0.55501 | 1.1789e-07
5502A | Resistance | 1100000 | 3301000 | 10000 | Ω | kΩ |  |  |  | 5.7933 | 1.659e-08 | 5.7933 | 1.6469e-08
5502A | Resistance | 3300000 | 11001000 | 100 | Ω | kΩ |  |  |  | 0.33 | 6.9e-07 | 0.29033 | 6.9272e-07
5502A | Resistance | 3300000 | 11001000 | 1000 | Ω | kΩ |  |  |  | 0.37289 | 6.8715e-07 | 0.37252 | 6.8716e-07
5502A | Resistance | 3300000 | 11001000 | 10000 | Ω | kΩ |  |  |  | 4.8626 | 4.4972e-07 | 4.8623 | 4.4966e-07
5502A | Resistance | 3300000 | 11001000 | 100000 | Ω | kΩ |  |  |  | 57.852 | 6.2356e-08 | 57.852 | 6.2304e-08
5502A | Resistance | 11000000 | 33001000 | 100 | Ω | kΩ |  |  |  | 2.9 | 1.2e-06 | 2.8873 | 1.1548e-06
5502A | Resistance | 11000000 | 33001000 | 1000 | Ω | kΩ |  |  |  | 2.9004 | 1.1545e-06 | 2.9 | 1.1545e-06
5502A | Resistance | 11000000 | 33001000 | 10000 | Ω | kΩ |  |  |  | 4.2484 | 1.1259e-06 | 4.2478 | 1.1259e-06
5502A | Resistance | 11000000 | 33001000 | 100000 | Ω | kΩ |  |  |  | 54.575 | 4.9852e-07 | 54.575 | 4.9846e-07
5502A | Resistance | 33000000 | 110010000 | 1000 | Ω | kΩ |  |  |  | 3.5 | 5.8e-06 | 3.4663 | 5.7735e-06
5502A | Resistance | 33000000 | 110010000 | 10000 | Ω | kΩ |  |  |  | 3.5803 | 5.7727e-06 | 3.5792 | 5.7727e-06
5502A | Resistance | 33000000 | 110010000 | 100000 | Ω | kΩ |  |  |  | 14.463 | 5.6975e-06 | 14.465 | 5.6974e-06
5502A | Resistance | 33000000 | 110010000 | 1000000 | Ω | kΩ |  |  |  | 503.99 | 3.2606e-06 | 503.99 | 3.2605e-06
5502A | Resistance | 110000000 | 330010000 | 1000 | Ω | kΩ |  |  |  | 120 | 5.8e-06 | 115.47 | 5.7735e-06
5502A | Resistance | 110000000 | 330010000 | 10000 | Ω | kΩ |  |  |  | 120 | 5.7734e-06 | 115.52 | 5.7734e-06
5502A | Resistance | 110000000 | 330010000 | 100000 | Ω | kΩ |  |  |  | 120 | 5.7671e-06 | 118.42 | 5.7671e-06
5502A | Resistance | 110000000 | 330010000 | 1000000 | Ω | kΩ |  |  |  | 371.66 | 5.2445e-06 | 371.67 | 5.2444e-06
5502A | Resistance | 330000000 | 1100010000 | 10000 | Ω | kΩ |  |  |  | 930 | 1.7e-05 | 577.19 | 1.7321e-05
5502A | Resistance | 330000000 | 1100010000 | 100000 | Ω | kΩ |  |  |  | 930 | 1.732e-05 | 578 | 1.732e-05
5502A | Resistance | 330000000 | 1100010000 | 1000000 | Ω | kΩ |  |  |  | 930 | 1.7297e-05 | 611.79 | 1.7297e-05
5502A | Resistance | 330000000 | 1100010000 | 10000000 | Ω | kΩ |  |  |  | 3454.2 | 1.5468e-05 | 3454.1 | 1.5468e-05
5502A | Resistance | 1100000000 | 1101000000 | 10000 | Ω | kΩ |  |  |  | 930 | 1.7e-05 | 577.19 | 1.7321e-05
5502A | Resistance | 1100000000 | 1101000000 | 100000 | Ω | kΩ |  |  |  | 930 | 1.732e-05 | 578 | 1.732e-05
5502A | Resistance | 1100000000 | 1101000000 | 1000000 | Ω | kΩ |  |  |  | 930 | 1.7297e-05 | 611.79 | 1.7297e-05
5502A | Resistance | 1100000000 | 1101000000 | 10000000 | Ω | kΩ |  |  |  | 3454.2 | 1.5468e-05 | 3454.1 | 1.5468e-05
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 0 | kHz | 0.25 | 4.6 | 0.23096 | 4.6188
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 0 | kHz | 0.25 | 4.6 | 0.231 | 4.6188
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 0 | kHz | 0.25 | 4.6003 | 0.23368 | 4.6165
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 0 | kHz | 0.46103 | 4.4237 | 0.45975 | 4.4244
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 0 | kHz | 5.6242 | 1.9179 | 5.6235 | 1.9171
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 0 | kHz | 0.28 | 1.8 | 0.23096 | 1.8476
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 0 | kHz | 0.28 | 1.8 | 0.23102 | 1.8476
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 0 | kHz | 0.28 | 1.8 | 0.23545 | 1.844
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 0 | kHz | 0.55279 | 1.6061 | 0.55118 | 1.6064
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 0 | kHz | 5.7507 | 0.38797 | 5.75 | 0.38676
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 0 | kHz | 0.24 | 0.92 | 0.23091 | 0.92432
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 0 | kHz | 0.24 | 0.92 | 0.23101 | 0.92427
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 0 | kHz | 0.24 | 0.92 | 0.23664 | 0.92018
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 0 | kHz | 0.59376 | 0.70044 | 0.59215 | 0.69986
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 0 | kHz | 5.7717 | 0.11872 | 5.771 | 0.11743
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 0 | 5 | kHz | 0.24 | 1.2 | 0.22816 | 1.1855
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 0 | 5 | kHz | 0.24 | 1.2 | 0.2283 | 1.1854
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 0 | 5 | kHz | 0.24 | 1.2 | 0.23356 | 1.1814
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 0 | 5 | kHz | 0.58062 | 0.9487 | 0.57901 | 0.94846
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 0 | 5 | kHz | 5.767 | 0.17953 | 5.7663 | 0.17826
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 5 | 20 | kHz | 0.27 | 3.5 | 0.22313 | 3.5466
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 5 | 20 | kHz | 0.27 | 3.5331 | 0.22319 | 3.5466
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 5 | 20 | kHz | 0.27 | 3.5338 | 0.22642 | 3.5438
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 5 | 20 | kHz | 0.48562 | 3.329 | 0.48418 | 3.3297
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 5 | 20 | kHz | 5.6818 | 1.2155 | 5.6811 | 1.2146
3458A | ACA | 0.1 | 1 | 1e-06 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44152 | 11.792
3458A | ACA | 0.1 | 1 | 1e-05 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44162 | 11.792
3458A | ACA | 0.1 | 1 | 0.0001 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44312 | 11.791
3458A | ACA | 0.1 | 1 | 0.001 | A | mA | 20 | 50 | kHz | 0.53612 | 11.934 | 0.5524 | 11.697
3458A | ACA | 0.1 | 1 | 0.01 | A | mA | 20 | 50 | kHz | 5.1678 | 8.3949 | 5.1615 | 8.3731
3458A | ACA | 1 | 10 | 1e-05 | A | mA | 0 | 0 | kHz | 0.55 | 0.12 | 0.51999 | 0.093087
3458A | ACA | 1 | 10 | 0.0001 | A | mA | 0 | 0 | kHz | 0.55 | 0.12 | 0.52375 | 0.092861
3458A | ACA | 1 | 10 | 0.001 | A | mA | 0 | 0 | kHz | 0.75909 | 0.099091 | 0.76542 | 0.079775
3458A | ACA | 1 | 10 | 0.01 | A | mA | 0 | 0 | kHz | 5.7992 | 0.016313 | 5.7916 | 0.016333
3458A | ACA | 1 | 10 | 0.1 | A | mA | 0 | 0 | kHz | 57.741 | 0.0016634 | 57.738 | 0.0016635
3458A | ACA | 1 | 10 | 1e-05 | A | mA | 0 | 1 | kHz | 0.46 | 0.17 | 0.49897 | 0.12063
3458A | ACA | 1 | 10 | 0.0001 | A | mA | 0 | 1 | kHz | 0.46 | 0.16806 | 0.50269 | 0.12039
3458A | ACA | 1 | 10 | 0.001 | A | mA | 0 | 1 | kHz | 0.72094 | 0.14191 | 0.74419 | 0.10576
3458A | ACA | 1 | 10 | 0.01 | A | mA | 0 | 1 | kHz | 5.7925 | 0.023676 | 5.7849 | 0.023706
3458A | ACA | 1 | 10 | 0.1 | A | mA | 0 | 1 | kHz | 57.74 | 0.002428 | 57.737 | 0.0024282
3458A | ACA | 1 | 10 | 1e-05 | A | mA | 1 | 5 | kHz | 1.4 | 0.66 | 1.326 | 0.66218
3458A | ACA | 1 | 10 | 0.0001 | A | mA | 1 | 5 | kHz | 1.4 | 0.66168 | 1.3279 | 0.66196
3458A | ACA | 1 | 10 | 0.001 | A | mA | 1 | 5 | kHz | 1.4676 | 0.64572 | 1.4594 | 0.64665
3458A | ACA | 1 | 10 | 0.01 | A | mA | 1 | 5 | kHz | 5.82 | 0.34986 | 5.8127 | 0.35015
3458A | ACA | 1 | 10 | 0.1 | A | mA | 1 | 5 | kHz | 57.725 | 0.044886 | 57.722 | 0.044889
3458A | ACA | 10 | 20 | 1e-05 | A | mA | 0 | 0 | kHz | 0.92 | 0.49 | 4.6374 | 0.081764
3458A | ACA | 10 | 20 | 0.0001 | A | mA | 0 | 0 | kHz | 0.92 | 0.49 | 4.6379 | 0.081759
3458A | ACA | 10 | 20 | 0.001 | A | mA | 0 | 0 | kHz | 0.92 | 0.49 | 4.6724 | 0.081363
3458A | ACA | 10 | 20 | 0.01 | A | mA | 0 | 0 | kHz | 5.1799 | 0.27701 | 7.3622 | 0.058218
3458A | ACA | 10 | 20 | 0.1 | A | mA | 0 | 0 | kHz | 57.914 | 0.0082611 | 57.911 | 0.0082616
3458A | ACA | 10 | 20 | 1e-05 | A | mA | 0 | 1 | kHz | 0 | 0.5 | 2.1832 | 0.11356
3458A | ACA | 10 | 20 | 0.0001 | A | mA | 0 | 1 | kHz | 0 | 0.5 | 2.1841 | 0.11354
3458A | ACA | 10 | 20 | 0.001 | A | mA | 0 | 1 | kHz | 0 | 0.5 | 2.2467 | 0.11228
3458A | ACA | 10 | 20 | 0.01 | A | mA | 0 | 1 | kHz | 3.3354 | 0.33323 | 6.0285 | 0.063256
3458A | ACA | 10 | 20 | 0.1 | A | mA | 0 | 1 | kHz | 57.758 | 0.0076264 | 57.755 | 0.0076268
3458A | ACA | 10 | 20 | 1e-05 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.0702 | 0.65987
3458A | ACA | 10 | 20 | 0.0001 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.0706 | 0.65986
3458A | ACA | 10 | 20 | 0.001 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.1025 | 0.65885
3458A | ACA | 10 | 20 | 0.01 | A | mA | 1 | 5 | kHz | 0.20999 | 0.9395 | 3.8087 | 0.57918
3458A | ACA | 10 | 20 | 0.1 | A | mA | 1 | 5 | kHz | 57.017 | 0.12296 | 57.014 | 0.12297
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 5 | 20 | kHz | 0.25 | 710 | 0.23471 | 715.83
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 5 | 20 | kHz | 0.25 | 711.93 | 0.23479 | 715.8
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 5 | 20 | kHz | 0.25 | 712.22 | 0.24058 | 711.87
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 5 | 20 | kHz | 0.6031 | 511.67 | 0.60272 | 509.95
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 5 | 20 | kHz | 5.7741 | 79.21 | 5.774 | 77.713
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 20 | 50 | kHz | 0.47 | 4800 | 0.46627 | 4761.7
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 20 | 50 | kHz | 0.47 | 4761.6 | 0.46636 | 4761.6
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 20 | 50 | kHz | 0.47 | 4759.9 | 0.46837 | 4760
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 20 | 50 | kHz | 0.64627 | 4616 | 0.64414 | 4616.1
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 20 | 50 | kHz | 5.636 | 2159.7 | 5.6349 | 2155.3
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 50 | 100 | kHz | 1.9 | 6500 | 1.7592 | 6558
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 50 | 100 | kHz | 1.9 | 6531.9 | 1.7594 | 6558.1
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 50 | 100 | kHz | 1.9 | 6537.4 | 1.7607 | 6558
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 50 | 100 | kHz | 1.9 | 6516.3 | 1.8354 | 6506.9
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 50 | 100 | kHz | 5.8456 | 4362.3 | 5.8321 | 4309.3
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 5 | 20 | kHz | 2.5 | 710 | 2.2891 | 721.65
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 5 | 20 | kHz | 2.5 | 712.46 | 2.2904 | 721.65
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 5 | 20 | kHz | 2.5 | 716.59 | 2.3512 | 717.7
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 5 | 20 | kHz | 6.0251 | 518.68 | 6.0044 | 513.52
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 5 | 20 | kHz | 57.745 | 83.689 | 57.738 | 78.853
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 20 | 50 | kHz | 5.1 | 4700 | 4.5719 | 4747.9
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 20 | 50 | kHz | 5.1 | 4739.1 | 4.5722 | 4747.9
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 20 | 50 | kHz | 5.1 | 4739.9 | 4.5921 | 4746.3
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 20 | 50 | kHz | 6.3814 | 4600.8 | 6.366 | 4600.8
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 20 | 50 | kHz | 56.354 | 2143.3 | 56.346 | 2139.9
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 50 | 100 | kHz | 19 | 6500 | 17.151 | 6602.1
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 50 | 100 | kHz | 19 | 6531.1 | 17.153 | 6602.1
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 50 | 100 | kHz | 19 | 6534.1 | 17.164 | 6601.7
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 50 | 100 | kHz | 19 | 6550 | 17.919 | 6548
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 50 | 100 | kHz | 58.261 | 4342.2 | 58.146 | 4318.8
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 5 | 20 | kHz | 25 | 710 | 22.891 | 721.64
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 5 | 20 | kHz | 25 | 712.1 | 22.9 | 721.62
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 5 | 20 | kHz | 25 | 715.16 | 23.497 | 717.59
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 5 | 20 | kHz | 60.128 | 515.53 | 60.018 | 512.89
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 5 | 20 | kHz | 577.41 | 80.741 | 577.37 | 78.263
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 20 | 50 | kHz | 51 | 4700 | 45.602 | 4746.9
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 20 | 50 | kHz | 51 | 4736.8 | 45.606 | 4746.9
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 20 | 50 | kHz | 51 | 4738 | 45.823 | 4745.2
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 20 | 50 | kHz | 64.008 | 4598.6 | 63.621 | 4599.4
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 20 | 50 | kHz | 563.71 | 2145.2 | 563.49 | 2139.2
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 50 | 100 | kHz | 190 | 6500 | 171.51 | 6602.2
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 50 | 100 | kHz | 190 | 6538.1 | 171.63 | 6602.5
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 50 | 100 | kHz | 190 | 6565.8 | 179.15 | 6551.2
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 50 | 100 | kHz | 582.09 | 4394.2 | 581.36 | 4329.2
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 50 | 100 | kHz | 5772.5 | 669.65 | 5772.4 | 620.3
3458A | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0 | 5 | kHz | 0.037 | 700 | 0.035478 | 705.46
3458A | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0 | 5 | kHz | 0.037 | 701.57 | 0.035484 | 705.45
3458A | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0 | 5 | kHz | 0.037 | 703.27 | 0.035899 | 702.95
3458A | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0 | 5 | kHz | 0.066315 | 546.37 | 0.066275 | 544.96
3458A | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0 | 5 | kHz | 0.57804 | 91.593 | 0.57802 | 90.228
3458A | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0 | 5 | kHz | 5.7736 | 9.7715 | 5.7728 | 9.2143
3458A | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0 | 5 | kHz | 57.736 | 1.1862 | 57.735 | 0.96332
3458A | ACA | 0 | 0.0001 | 0.001 | A | µA | 0 | 5 | kHz | 577.36 | 0.20221 | 577.35 | 0.11305
3458A | ACA | 0 | 0.0001 | 0.01 | A | µA | 0 | 5 | kHz | 5773.6 | 0.053657 | 5773.6 | 0.017992
3458A | ACA | 0 | 0.0001 | 0.1 | A | µA | 0 | 5 | kHz | 57736 | 0.01874 | 57736 | 0.0044741
3458A | ACA | 0 | 0.0001 | 1 | A | µA | 0 | 5 | kHz | 577360 | 0.0072242 | 577360 | 0.0015173
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 5 | kHz | 0.25 | 350 | 0.23001 | 364.09
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 5 | kHz | 0.25 | 354.91 | 0.23009 | 364.09
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 5 | kHz | 0.25 | 359.01 | 0.23663 | 360.5
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 5 | kHz | 0.61486 | 217.08 | 0.61454 | 214.58
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 5 | kHz | 5.7771 | 28.936 | 5.777 | 27.424
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 5 | kHz | 2.5 | 350 | 2.2952 | 364.62
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 5 | kHz | 2.5 | 355.75 | 2.2965 | 364.68
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 5 | kHz | 2.5 | 362.33 | 2.3641 | 361.33
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 5 | kHz | 6.1659 | 223.94 | 6.1472 | 216.14
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 5 | kHz | 57.778 | 33.176 | 57.771 | 28.291
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 5 | kHz | 25 | 350 | 22.95 | 364.61
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 5 | kHz | 25 | 355.17 | 22.961 | 364.62
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 5 | kHz | 25 | 360.06 | 23.625 | 361.08
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 5 | kHz | 61.548 | 219.19 | 61.449 | 215.18
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 5 | kHz | 577.74 | 30.199 | 577.71 | 27.695
3458A | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.077 | 4200 | 0.076632 | 4198.9
3458A | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.077 | 4200 | 0.076633 | 4198.9
3458A | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.077 | 4200 | 0.076839 | 4197.2
3458A | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.095326 | 4045.9 | 0.095288 | 4045.8
3458A | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0 | 0 | kHz | 0.58121 | 1804 | 0.58119 | 1803.1
3458A | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0 | 0 | kHz | 5.7739 | 210.05 | 5.7723 | 209.55
3458A | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0 | 0 | kHz | 57.736 | 21.197 | 57.734 | 20.994
3458A | ACA | 0 | 0.0001 | 0.001 | A | µA | 0 | 0 | kHz | 577.36 | 2.1805 | 577.35 | 2.0995
3458A | ACA | 0 | 0.0001 | 0.01 | A | µA | 0 | 0 | kHz | 5773.6 | 0.24237 | 5773.6 | 0.20995
3458A | ACA | 0 | 0.0001 | 0.1 | A | µA | 0 | 0 | kHz | 57736 | 0.033964 | 57736 | 0.020995
3458A | ACA | 0 | 0.0001 | 1 | A | µA | 0 | 0 | kHz | 577360 | 0.0072871 | 577360 | 0.002099
3458A | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.038 | 1700 | 0.034644 | 1732.1
3458A | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.038 | 1700 | 0.034648 | 1732.1
3458A | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.038 | 1700 | 0.034998 | 1729.4
3458A | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.062399 | 1534.5 | 0.062349 | 1534
3458A | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0 | 0 | kHz | 0.57594 | 378.74 | 0.57592 | 377.44
3458A | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0 | 0 | kHz | 5.7734 | 39.656 | 5.77 | 39.099
3458A | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0 | 0 | kHz | 57.736 | 4.1759 | 57.732 | 3.953
3458A | ACA | 0 | 0.0001 | 0.001 | A | µA | 0 | 0 | kHz | 577.36 | 0.50118 | 577.35 | 0.41202
3458A | ACA | 0 | 0.0001 | 0.01 | A | µA | 0 | 0 | kHz | 5773.6 | 0.083554 | 5773.5 | 0.047889
3458A | ACA | 0 | 0.0001 | 0.1 | A | µA | 0 | 0 | kHz | 57736 | 0.02173 | 57736 | 0.0074638
3458A | ACA | 0 | 0.0001 | 1 | A | µA | 0 | 0 | kHz | 577360 | 0.007523 | 577360 | 0.0018174
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.25 | 4600 | 0.23095 | 4618.8
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.25 | 4600 | 0.23098 | 4618.8
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.25 | 4600.3 | 0.23363 | 4616.5
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.4599 | 4424.7 | 0.45952 | 4424.6
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 0 | kHz | 5.6235 | 1918.3 | 5.6233 | 1917.2
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.27 | 1700 | 0.23094 | 1732.1
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.27 | 1700 | 0.231 | 1732.1
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.27 | 1700 | 0.23546 | 1728.5
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.5563 | 1491.3 | 0.55586 | 1490.6
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 0 | kHz | 5.7533 | 346.75 | 5.7532 | 345.32
3458A | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0 | 0 | kHz | 0.24 | 690 | 0.23094 | 693.02
3458A | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0 | 0 | kHz | 0.24 | 690 | 0.23102 | 692.99
3458A | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0 | 0 | kHz | 0.24 | 690 | 0.23691 | 689
3458A | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0 | 0 | kHz | 0.60249 | 489.7 | 0.6021 | 487.94
3458A | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0 | 0 | kHz | 5.7742 | 74.832 | 5.7741 | 73.333
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 0 | kHz | 2.5 | 4600 | 2.3096 | 4618.8
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 0 | kHz | 2.5 | 4600 | 2.3101 | 4618.8
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 0 | kHz | 2.5 | 4600.7 | 2.3373 | 4616.5
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 0 | kHz | 4.6177 | 4424.3 | 4.5991 | 4424.5
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 0 | kHz | 56.243 | 1921.2 | 56.235 | 1917.8
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 0 | kHz | 2.7 | 1700 | 2.3095 | 1732.1
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 0 | kHz | 2.7 | 1700 | 2.311 | 1732
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 0 | kHz | 2.7 | 1700 | 2.3566 | 1728.5
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 0 | kHz | 5.585 | 1492.6 | 5.5626 | 1490.9
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 0 | kHz | 57.541 | 350.7 | 57.534 | 346.11
3458A | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0 | 0 | kHz | 2.4 | 690 | 2.3095 | 693.04
3458A | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0 | 0 | kHz | 2.4 | 690 | 2.3107 | 693.04
3458A | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0 | 0 | kHz | 2.4 | 690 | 2.3717 | 689.14
3458A | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0 | 0 | kHz | 6.0458 | 494.17 | 6.0252 | 488.84
3458A | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0 | 0 | kHz | 57.749 | 79.009 | 57.742 | 74.17
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 0 | kHz | 25 | 4600 | 23.095 | 4618.8
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 0 | kHz | 25 | 4600 | 23.099 | 4618.8
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 0 | kHz | 25 | 4600.4 | 23.366 | 4616.5
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 0 | kHz | 46.067 | 4424.4 | 45.964 | 4424.6
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 0 | kHz | 562.38 | 1919.1 | 562.34 | 1917.3
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 0 | kHz | 27 | 1700 | 23.095 | 1732.1
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 0 | kHz | 27 | 1700 | 23.102 | 1732.1
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 0 | kHz | 27 | 1700 | 23.554 | 1728.5
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 0 | kHz | 55.716 | 1491.6 | 55.6 | 1490.7
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 0 | kHz | 575.37 | 347.9 | 575.33 | 345.55
3458A | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0 | 0 | kHz | 24 | 690 | 23.094 | 693.05
3458A | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0 | 0 | kHz | 24 | 690 | 23.103 | 693.03
3458A | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0 | 0 | kHz | 24 | 690 | 23.701 | 689.05
3458A | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0 | 0 | kHz | 60.336 | 490.93 | 60.227 | 488.21
3458A | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0 | 0 | kHz | 577.45 | 76.063 | 577.41 | 73.582
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 1000 | 2000 | kHz | 6.1 | 17 | 1.1338 | 17.487
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1338 | 17.487
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1338 | 17.487
3458A | ACV | 1 | 10 | 0.001 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1448 | 17.486
3458A | ACV | 1 | 10 | 0.01 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.167 | 2.0915 | 17.401
3458A | ACV | 1 | 10 | 0.1 | V | mV | 1000 | 2000 | kHz | 46.847 | 13.84 | 43.262 | 13.841
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 2000 | 4000 | kHz | 12 | 46 | 8.0599 | 46.328
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.06 | 46.328
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.06 | 46.328
3458A | ACV | 1 | 10 | 0.001 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.0603 | 46.328
3458A | ACV | 1 | 10 | 0.01 | V | mV | 2000 | 4000 | kHz | 12 | 46.042 | 8.3953 | 46.298
3458A | ACV | 1 | 10 | 0.1 | V | mV | 2000 | 4000 | kHz | 35.391 | 43.95 | 33.018 | 43.95
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 4000 | 8000 | kHz | 15 | 46 | 9.1632 | 46.576
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1632 | 46.576
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1632 | 46.576
3458A | ACV | 1 | 10 | 0.001 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1658 | 46.576
3458A | ACV | 1 | 10 | 0.01 | V | mV | 4000 | 8000 | kHz | 15 | 46.329 | 9.49 | 46.547
3458A | ACV | 1 | 10 | 0.1 | V | mV | 4000 | 8000 | kHz | 36.034 | 44.241 | 33.704 | 44.241
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 8000 | 10000 | kHz | 46 | 170 | 11.528 | 173.36
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.529 | 173.36
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.529 | 173.36
3458A | ACV | 1 | 10 | 0.001 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.584 | 173.35
3458A | ACV | 1 | 10 | 0.01 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.629 | 173.35
3458A | ACV | 1 | 10 | 0.1 | V | mV | 8000 | 10000 | kHz | 46 | 170.24 | 20.344 | 172.48
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 0 | 0 | kHz | 0.46 | 0.092 | 0.45053 | 0.092199
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 0 | 0 | kHz | 0.46 | 0.091968 | 0.45057 | 0.092197
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 0 | 0 | kHz | 0.46 | 0.092019 | 0.4538 | 0.091994
3458A | ACV | 1 | 10 | 0.001 | V | mV | 0 | 0 | kHz | 0.71619 | 0.077546 | 0.71499 | 0.077402
3458A | ACV | 1 | 10 | 0.01 | V | mV | 0 | 0 | kHz | 5.7845 | 0.015263 | 5.784 | 0.01505
3458A | ACV | 1 | 10 | 0.1 | V | mV | 0 | 0 | kHz | 57.737 | 0.0016174 | 57.723 | 0.001529
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 0 | 1 | kHz | 0.25 | 0.084 | 0.22625 | 0.085545
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 0 | 1 | kHz | 0.25 | 0.084438 | 0.22632 | 0.08554
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 0 | 1 | kHz | 0.25 | 0.084683 | 0.23198 | 0.085127
3458A | ACV | 1 | 10 | 0.001 | V | mV | 0 | 1 | kHz | 0.59428 | 0.063506 | 0.59284 | 0.06333
3458A | ACV | 1 | 10 | 0.01 | V | mV | 0 | 1 | kHz | 5.7723 | 0.010443 | 5.7717 | 0.010227
3458A | ACV | 1 | 10 | 0.1 | V | mV | 0 | 1 | kHz | 57.736 | 0.0011207 | 57.726 | 0.0010323
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 1 | 20 | kHz | 0.28 | 0.16 | 0.22835 | 0.16434
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 1 | 20 | kHz | 0.28 | 0.16134 | 0.22837 | 0.16434
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 1 | 20 | kHz | 0.28 | 0.16156 | 0.23295 | 0.16397
3458A | ACV | 1 | 10 | 0.001 | V | mV | 1 | 20 | kHz | 0.56386 | 0.14001 | 0.55818 | 0.14006
3458A | ACV | 1 | 10 | 0.01 | V | mV | 1 | 20 | kHz | 5.7578 | 0.031829 | 5.7555 | 0.031388
3458A | ACV | 1 | 10 | 0.1 | V | mV | 1 | 20 | kHz | 57.735 | 0.0034135 | 57.705 | 0.003222
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.22967 | 0.34776
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.22973 | 0.34775
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.23282 | 0.34749
3458A | ACV | 1 | 10 | 0.001 | V | mV | 20 | 50 | kHz | 0.50358 | 0.32533 | 0.49042 | 0.32615
3458A | ACV | 1 | 10 | 0.01 | V | mV | 20 | 50 | kHz | 5.6913 | 0.11808 | 5.6847 | 0.11766
3458A | ACV | 1 | 10 | 0.1 | V | mV | 20 | 50 | kHz | 57.728 | 0.013146 | 57.621 | 0.01289
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 50 | 100 | kHz | 0.27 | 0.94 | 0.22603 | 0.94121
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 50 | 100 | kHz | 0.27 | 0.93759 | 0.22604 | 0.94121
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 50 | 100 | kHz | 0.27 | 0.93768 | 0.22761 | 0.94107
3458A | ACV | 1 | 10 | 0.001 | V | mV | 50 | 100 | kHz | 0.38942 | 0.92692 | 0.37412 | 0.92813
3458A | ACV | 1 | 10 | 0.01 | V | mV | 50 | 100 | kHz | 5.3092 | 0.59368 | 5.2965 | 0.59386
3458A | ACV | 1 | 10 | 0.1 | V | mV | 50 | 100 | kHz | 57.665 | 0.087823 | 57.153 | 0.087462
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1359 | 3.4831
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1359 | 3.4831
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1365 | 3.483
3458A | ACV | 1 | 10 | 0.001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1754 | 3.4796
3458A | ACV | 1 | 10 | 0.01 | V | mV | 100 | 300 | kHz | 4.1895 | 3.2244 | 4.1679 | 3.2259
3458A | ACV | 1 | 10 | 0.1 | V | mV | 100 | 300 | kHz | 56.807 | 1.1229 | 54.694 | 1.1224
3458A | ACV | 1 | 10 | 1e-06 | V | mV | 300 | 1000 | kHz | 0.83 | 12 | 1.127 | 11.688
3458A | ACV | 1 | 10 | 1e-05 | V | mV | 300 | 1000 | kHz | 0.83 | 11.988 | 1.127 | 11.688
3458A | ACV | 1 | 10 | 0.0001 | V | mV | 300 | 1000 | kHz | 0.83 | 11.988 | 1.1272 | 11.688
3458A | ACV | 1 | 10 | 0.001 | V | mV | 300 | 1000 | kHz | 0.84663 | 11.986 | 1.141 | 11.687
3458A | ACV | 1 | 10 | 0.01 | V | mV | 300 | 1000 | kHz | 2.2239 | 11.849 | 2.4896 | 11.566
3458A | ACV | 1 | 10 | 0.1 | V | mV | 300 | 1000 | kHz | 51.132 | 8.0252 | 47.575 | 8.0258
3458A | ACV | 0.1 | 1 | 1e-07 | V | mV | 2000 | 4000 | kHz | 1.2 | 46 | 0.80072 | 46.326
3458A | ACV | 0.1 | 1 | 1e-06 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.029 | 0.80072 | 46.326
3458A | ACV | 0.1 | 1 | 1e-05 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.029 | 0.80073 | 46.326
3458A | ACV | 0.1 | 1 | 0.0001 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.03 | 0.8015 | 46.325
3458A | ACV | 0.1 | 1 | 0.001 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.034 | 0.83431 | 46.296
3458A | ACV | 0.1 | 1 | 0.01 | V | mV | 2000 | 4000 | kHz | 3.5375 | 43.944 | 3.2985 | 43.945
3458A | ACV | 0.1 | 1 | 1e-07 | V | mV | 4000 | 8000 | kHz | 1.3 | 46 | 0.91606 | 46.326
3458A | ACV | 0.1 | 1 | 1e-06 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91606 | 46.326
3458A | ACV | 0.1 | 1 | 1e-05 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91607 | 46.326
3458A | ACV | 0.1 | 1 | 0.0001 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91665 | 46.325
3458A | ACV | 0.1 | 1 | 0.001 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.051 | 0.94915 | 46.296
3458A | ACV | 0.1 | 1 | 0.01 | V | mV | 4000 | 8000 | kHz | 3.6131 | 43.982 | 3.378 | 43.984
3458A | ACV | 0.1 | 1 | 1e-07 | V | mV | 8000 | 10000 | kHz | 4.6 | 170 | 1.146 | 173.36
3458A | ACV | 0.1 | 1 | 1e-06 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.146 | 173.36
3458A | ACV | 0.1 | 1 | 1e-05 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.146 | 173.36
3458A | ACV | 0.1 | 1 | 0.0001 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.1461 | 173.36
3458A | ACV | 0.1 | 1 | 0.001 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.156 | 173.35
3458A | ACV | 0.1 | 1 | 0.01 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.23 | 2.0278 | 172.49
3458A | ACV | 10 | 100 | 1e-05 | V | mV | 100 | 300 | kHz | 15 | 4.6 | 11.575 | 4.6268
3458A | ACV | 10 | 100 | 0.0001 | V | mV | 100 | 300 | kHz | 15 | 4.6028 | 11.575 | 4.6268
3458A | ACV | 10 | 100 | 0.001 | V | mV | 100 | 300 | kHz | 15 | 4.6028 | 11.576 | 4.6268
3458A | ACV | 10 | 100 | 0.01 | V | mV | 100 | 300 | kHz | 15 | 4.6033 | 11.89 | 4.624
3458A | ACV | 10 | 100 | 0.1 | V | mV | 100 | 300 | kHz | 37.775 | 4.4 | 37.726 | 4.4003
3458A | ACV | 10 | 100 | 1 | V | mV | 100 | 300 | kHz | 561.72 | 1.8549 | 536.24 | 1.8547
3458A | ACV | 10 | 100 | 1e-05 | V | mV | 300 | 1000 | kHz | 46 | 17 | 12.317 | 17.334
3458A | ACV | 10 | 100 | 0.0001 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.318 | 17.334
3458A | ACV | 10 | 100 | 0.001 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.318 | 17.334
3458A | ACV | 10 | 100 | 0.01 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.427 | 17.333
3458A | ACV | 10 | 100 | 0.1 | V | mV | 300 | 1000 | kHz | 46 | 17.03 | 21.972 | 17.247
3458A | ACV | 10 | 100 | 1 | V | mV | 300 | 1000 | kHz | 469.61 | 13.691 | 433.99 | 13.692
3458A | ACV | 100 | 1000 | 0.0001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.055 | 0.69387
3458A | ACV | 100 | 1000 | 0.001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.063 | 0.69386
3458A | ACV | 100 | 1000 | 0.01 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.256 | 0.69369
3458A | ACV | 100 | 1000 | 0.1 | V | mV | 1 | 20 | kHz | 40.883 | 0.68212 | 41.185 | 0.67806
3458A | ACV | 100 | 1000 | 1 | V | mV | 1 | 20 | kHz | 547.47 | 0.37317 | 547.4 | 0.3731
3458A | ACV | 100 | 1000 | 0.0001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 22.895 | 1.3896
3458A | ACV | 100 | 1000 | 0.001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 22.896 | 1.3896
3458A | ACV | 100 | 1000 | 0.01 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.007 | 1.3895
3458A | ACV | 100 | 1000 | 0.1 | V | mV | 20 | 50 | kHz | 33.01 | 1.39 | 33.875 | 1.3798
3458A | ACV | 100 | 1000 | 1 | V | mV | 20 | 50 | kHz | 496.92 | 1.0292 | 496.68 | 1.0293
3458A | ACV | 100 | 1000 | 0.0001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.116 | 3.4656
3458A | ACV | 100 | 1000 | 0.001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.117 | 3.4656
3458A | ACV | 100 | 1000 | 0.01 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.153 | 3.4656
3458A | ACV | 100 | 1000 | 0.1 | V | mV | 50 | 100 | kHz | 24.483 | 3.4985 | 28.03 | 3.4612
3458A | ACV | 100 | 1000 | 1 | V | mV | 50 | 100 | kHz | 369.25 | 3.1671 | 368.87 | 3.1673
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 0 | 0 | kHz | 4.7 | 0.23 | 4.6186 | 0.23291
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 0 | 0 | kHz | 4.7 | 0.23 | 4.619 | 0.2329
3458A | ACV | 10 | 20 | 0.001 | V | mV | 0 | 0 | kHz | 4.7 | 0.23 | 4.6486 | 0.23231
3458A | ACV | 10 | 20 | 0.01 | V | mV | 0 | 0 | kHz | 7.1399 | 0.18969 | 7.1403 | 0.18932
3458A | ACV | 10 | 20 | 0.1 | V | mV | 0 | 0 | kHz | 57.827 | 0.032722 | 57.828 | 0.0324
3458A | ACV | 10 | 20 | 1 | V | mV | 0 | 0 | kHz | 577.36 | 0.003403 | 577.07 | 0.0032722
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 0 | 1 | kHz | 2.4 | 0.23 | 2.298 | 0.23306
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 0 | 1 | kHz | 2.4 | 0.23 | 2.2985 | 0.23305
3458A | ACV | 10 | 20 | 0.001 | V | mV | 0 | 1 | kHz | 2.4 | 0.23 | 2.3458 | 0.23187
3458A | ACV | 10 | 20 | 0.01 | V | mV | 0 | 1 | kHz | 5.7569 | 0.1647 | 5.7574 | 0.16425
3458A | ACV | 10 | 20 | 0.1 | V | mV | 0 | 1 | kHz | 57.687 | 0.023591 | 57.688 | 0.023267
3458A | ACV | 10 | 20 | 1 | V | mV | 0 | 1 | kHz | 577.35 | 0.0024695 | 577.14 | 0.0023387
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.298 | 0.23306
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.2985 | 0.23305
3458A | ACV | 10 | 20 | 0.001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.3458 | 0.23187
3458A | ACV | 10 | 20 | 0.01 | V | mV | 1 | 20 | kHz | 5.7569 | 0.1647 | 5.7574 | 0.16425
3458A | ACV | 10 | 20 | 0.1 | V | mV | 1 | 20 | kHz | 57.687 | 0.023591 | 57.688 | 0.023267
3458A | ACV | 10 | 20 | 1 | V | mV | 1 | 20 | kHz | 577.35 | 0.0024695 | 577.14 | 0.0023387
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.2656 | 0.40924
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.266 | 0.40923
3458A | ACV | 10 | 20 | 0.001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.302 | 0.40822
3458A | ACV | 10 | 20 | 0.01 | V | mV | 20 | 50 | kHz | 5.2425 | 0.33547 | 5.2373 | 0.3351
3458A | ACV | 10 | 20 | 0.1 | V | mV | 20 | 50 | kHz | 57.496 | 0.059412 | 57.495 | 0.058915
3458A | ACV | 10 | 20 | 1 | V | mV | 20 | 50 | kHz | 577.33 | 0.0061608 | 576.8 | 0.0059565
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3767 | 1.3889
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3769 | 1.3889
3458A | ACV | 10 | 20 | 0.001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3922 | 1.3884
3458A | ACV | 10 | 20 | 0.01 | V | mV | 50 | 100 | kHz | 3.8217 | 1.3443 | 3.8184 | 1.3442
3458A | ACV | 10 | 20 | 0.1 | V | mV | 50 | 100 | kHz | 54.831 | 0.51571 | 54.83 | 0.51529
3458A | ACV | 10 | 20 | 1 | V | mV | 50 | 100 | kHz | 577.03 | 0.05599 | 572.43 | 0.055786
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 100 | 300 | kHz | 13 | 4.6 | 11.573 | 4.6268
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 100 | 300 | kHz | 13 | 4.6109 | 11.574 | 4.6268
3458A | ACV | 10 | 20 | 0.001 | V | mV | 100 | 300 | kHz | 13 | 4.6111 | 11.577 | 4.6267
3458A | ACV | 10 | 20 | 0.01 | V | mV | 100 | 300 | kHz | 13 | 4.6141 | 11.988 | 4.6141
3458A | ACV | 10 | 20 | 0.1 | V | mV | 100 | 300 | kHz | 44.406 | 3.7323 | 44.403 | 3.7322
3458A | ACV | 10 | 20 | 1 | V | mV | 100 | 300 | kHz | 573.82 | 0.64234 | 542.92 | 0.64214
3458A | ACV | 10 | 20 | 1e-05 | V | mV | 300 | 1000 | kHz | 19 | 17 | 12.388 | 17.324
3458A | ACV | 10 | 20 | 0.0001 | V | mV | 300 | 1000 | kHz | 19 | 17.087 | 12.389 | 17.324
3458A | ACV | 10 | 20 | 0.001 | V | mV | 300 | 1000 | kHz | 19 | 17.087 | 12.389 | 17.324
3458A | ACV | 10 | 20 | 0.01 | V | mV | 300 | 1000 | kHz | 19 | 17.092 | 12.517 | 17.32
3458A | ACV | 10 | 20 | 0.1 | V | mV | 300 | 1000 | kHz | 25.325 | 16.908 | 25.323 | 16.908
3458A | ACV | 10 | 20 | 1 | V | mV | 300 | 1000 | kHz | 533.13 | 7.3338 | 437.38 | 7.3336
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 0 | 0 | kHz | 5 | 0.23 | 4.6009 | 0.23351
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 0 | 0 | kHz | 5 | 0.23167 | 4.6011 | 0.23351
3458A | ACV | 20 | 100 | 0.001 | V | mV | 0 | 0 | kHz | 5 | 0.23175 | 4.6195 | 0.23336
3458A | ACV | 20 | 100 | 0.01 | V | mV | 0 | 0 | kHz | 6.5042 | 0.22022 | 6.5015 | 0.22022
3458A | ACV | 20 | 100 | 0.1 | V | mV | 0 | 0 | kHz | 57.057 | 0.070854 | 57.056 | 0.070819
3458A | ACV | 20 | 100 | 1 | V | mV | 0 | 0 | kHz | 577.28 | 0.0075312 | 576.01 | 0.0075149
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 0 | 1 | kHz | 2.7 | 0.23 | 2.2758 | 0.23386
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 0 | 1 | kHz | 2.7 | 0.2318 | 2.276 | 0.23386
3458A | ACV | 20 | 100 | 0.001 | V | mV | 0 | 1 | kHz | 2.7 | 0.23188 | 2.2998 | 0.23364
3458A | ACV | 20 | 100 | 0.01 | V | mV | 0 | 1 | kHz | 4.7115 | 0.21577 | 4.7084 | 0.21577
3458A | ACV | 20 | 100 | 0.1 | V | mV | 0 | 1 | kHz | 56.896 | 0.062832 | 56.895 | 0.062796
3458A | ACV | 20 | 100 | 1 | V | mV | 0 | 1 | kHz | 577.27 | 0.0066108 | 576.14 | 0.0065945
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8564 | 0.20787
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8565 | 0.20787
3458A | ACV | 20 | 100 | 0.001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.882 | 0.20768
3458A | ACV | 20 | 100 | 0.01 | V | mV | 1 | 20 | kHz | 7.1093 | 0.19179 | 7.1056 | 0.1918
3458A | ACV | 20 | 100 | 0.1 | V | mV | 1 | 20 | kHz | 57.594 | 0.05585 | 57.592 | 0.055819
3458A | ACV | 20 | 100 | 1 | V | mV | 1 | 20 | kHz | 577.34 | 0.0058763 | 576.84 | 0.0058618
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 20 | 50 | kHz | 6.8 | 0.37 | 6.7323 | 0.36645
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36642 | 6.7324 | 0.36645
3458A | ACV | 20 | 100 | 0.001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36631 | 6.7495 | 0.36632
3458A | ACV | 20 | 100 | 0.01 | V | mV | 20 | 50 | kHz | 8.3578 | 0.35409 | 8.3518 | 0.35408
3458A | ACV | 20 | 100 | 0.1 | V | mV | 20 | 50 | kHz | 57.163 | 0.15069 | 57.159 | 0.15056
3458A | ACV | 20 | 100 | 1 | V | mV | 20 | 50 | kHz | 577.28 | 0.017106 | 575.94 | 0.01704
3458A | ACV | 20 | 100 | 1e-05 | V | mV | 50 | 100 | kHz | 22 | 1.2 | 17.722 | 1.2358
3458A | ACV | 20 | 100 | 0.0001 | V | mV | 50 | 100 | kHz | 22 | 1.2034 | 17.723 | 1.2358
3458A | ACV | 20 | 100 | 0.001 | V | mV | 50 | 100 | kHz | 22 | 1.2034 | 17.724 | 1.2358
3458A | ACV | 20 | 100 | 0.01 | V | mV | 50 | 100 | kHz | 22 | 1.2047 | 18.315 | 1.2311
3458A | ACV | 20 | 100 | 0.1 | V | mV | 50 | 100 | kHz | 55.391 | 0.9727 | 55.373 | 0.9727
3458A | ACV | 20 | 100 | 1 | V | mV | 50 | 100 | kHz | 576.34 | 0.18075 | 568.41 | 0.18063
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 0 | 0 | kHz | 48 | 0.46 | 46.118 | 0.46325
3458A | ACV | 100 | 330 | 0.001 | V | mV | 0 | 0 | kHz | 48 | 0.46086 | 46.12 | 0.46325
3458A | ACV | 100 | 330 | 0.01 | V | mV | 0 | 0 | kHz | 48 | 0.46129 | 46.341 | 0.46283
3458A | ACV | 100 | 330 | 0.1 | V | mV | 0 | 0 | kHz | 66.332 | 0.42712 | 66.294 | 0.42698
3458A | ACV | 100 | 330 | 1 | V | mV | 0 | 0 | kHz | 573.42 | 0.11327 | 573.42 | 0.11294
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 0 | 1 | kHz | 25 | 0.46 | 23.034 | 0.46337
3458A | ACV | 100 | 330 | 0.001 | V | mV | 0 | 1 | kHz | 25 | 0.46067 | 23.037 | 0.46337
3458A | ACV | 100 | 330 | 0.01 | V | mV | 0 | 1 | kHz | 25 | 0.46116 | 23.338 | 0.46274
3458A | ACV | 100 | 330 | 0.1 | V | mV | 0 | 1 | kHz | 49.032 | 0.41285 | 48.983 | 0.41271
3458A | ACV | 100 | 330 | 1 | V | mV | 0 | 1 | kHz | 571.92 | 0.096252 | 571.92 | 0.095921
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.059 | 0.69384
3458A | ACV | 100 | 330 | 0.001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.061 | 0.69384
3458A | ACV | 100 | 330 | 0.01 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.291 | 0.69334
3458A | ACV | 100 | 330 | 0.1 | V | mV | 1 | 20 | kHz | 44.083 | 0.65012 | 43.963 | 0.65028
3458A | ACV | 100 | 330 | 1 | V | mV | 1 | 20 | kHz | 565.1 | 0.19691 | 565.03 | 0.19677
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.053 | 1.3879
3458A | ACV | 100 | 330 | 0.001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.054 | 1.3879
3458A | ACV | 100 | 330 | 0.01 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.188 | 1.3876
3458A | ACV | 100 | 330 | 0.1 | V | mV | 20 | 50 | kHz | 35.929 | 1.3608 | 35.878 | 1.3595
3458A | ACV | 100 | 330 | 1 | V | mV | 20 | 50 | kHz | 533.77 | 0.66047 | 533.57 | 0.66042
3458A | ACV | 100 | 330 | 0.0001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.11 | 3.4656
3458A | ACV | 100 | 330 | 0.001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.111 | 3.4656
3458A | ACV | 100 | 330 | 0.01 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.164 | 3.4655
3458A | ACV | 100 | 330 | 0.1 | V | mV | 50 | 100 | kHz | 24.914 | 3.4942 | 28.927 | 3.4523
3458A | ACV | 100 | 330 | 1 | V | mV | 50 | 100 | kHz | 418.09 | 2.6787 | 417.63 | 2.6793
3458A | ACV | 330 | 1000 | 0.0001 | V | mV | 0 | 0 | kHz | 50 | 0.46 | 46.107 | 0.46338
3458A | ACV | 330 | 1000 | 0.001 | V | mV | 0 | 0 | kHz | 50 | 0.46054 | 46.108 | 0.46338
3458A | ACV | 330 | 1000 | 0.01 | V | mV | 0 | 0 | kHz | 50 | 0.46059 | 46.301 | 0.46322
3458A | ACV | 330 | 1000 | 0.1 | V | mV | 0 | 0 | kHz | 64.184 | 0.4486 | 64.129 | 0.44862
3458A | ACV | 330 | 1000 | 1 | V | mV | 0 | 0 | kHz | 564.15 | 0.20595 | 564.12 | 0.20589
3458A | ACV | 330 | 1000 | 0.0001 | V | mV | 0 | 1 | kHz | 27 | 0.46 | 23.031 | 0.46343
3458A | ACV | 330 | 1000 | 0.001 | V | mV | 0 | 1 | kHz | 27 | 0.46051 | 23.032 | 0.46343
3458A | ACV | 330 | 1000 | 0.01 | V | mV | 0 | 1 | kHz | 27 | 0.46056 | 23.29 | 0.46321
3458A | ACV | 330 | 1000 | 0.1 | V | mV | 0 | 1 | kHz | 45.921 | 0.44399 | 45.855 | 0.44402
3458A | ACV | 330 | 1000 | 1 | V | mV | 0 | 1 | kHz | 562.27 | 0.1928 | 562.23 | 0.19274
3458A | ACV | 330 | 1000 | 0.0001 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.056 | 0.69387
3458A | ACV | 330 | 1000 | 0.001 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.063 | 0.69386
3458A | ACV | 330 | 1000 | 0.01 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.256 | 0.69369
3458A | ACV | 330 | 1000 | 0.1 | V | mV | 1 | 10 | kHz | 40.883 | 0.68212 | 41.185 | 0.67806
3458A | ACV | 330 | 1000 | 1 | V | mV | 1 | 10 | kHz | 547.47 | 0.37317 | 547.4 | 0.3731
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 1000 | 4000 | kHz | 23 | 80000 | 22.679 | 79411
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 1000 | 4000 | kHz | 23 | 79583 | 22.679 | 79411
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 1000 | 4000 | kHz | 23 | 79581 | 22.688 | 79410
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 1000 | 4000 | kHz | 23.293 | 79521 | 23.228 | 79358
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 1000 | 4000 | kHz | 57.909 | 76095 | 57.746 | 76108
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 1000 | 4000 | kHz | 573.98 | 42637 | 570.56 | 42634
3458A | ACV | 0 | 0.01 | 0.01 | V | µV | 1000 | 4000 | kHz | 5773.1 | 5803.4 | 5766 | 5798.9
3458A | ACV | 0 | 0.01 | 0.1 | V | µV | 1000 | 4000 | kHz | 57735 | 584.6 | 57728 | 582.75
3458A | ACV | 0 | 0.01 | 1 | V | µV | 1000 | 4000 | kHz | 577360 | 59.02 | 577350 | 58.278
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 4000 | 8000 | kHz | 31 | 230000 | 30.119 | 228890
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 4000 | 8000 | kHz | 31 | 228990 | 30.12 | 228890
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 4000 | 8000 | kHz | 31 | 228990 | 30.12 | 228890
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 4000 | 8000 | kHz | 31 | 228940 | 30.434 | 228860
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 4000 | 8000 | kHz | 56.084 | 226390 | 55.738 | 226400
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 4000 | 8000 | kHz | 561.69 | 182820 | 557.14 | 182830
3458A | ACV | 0 | 0.01 | 0.01 | V | µV | 4000 | 8000 | kHz | 5769.3 | 45269 | 5751.2 | 45261
3458A | ACV | 0 | 0.01 | 0.1 | V | µV | 4000 | 8000 | kHz | 57735 | 4703.5 | 57713 | 4700
3458A | ACV | 0 | 0.01 | 1 | V | µV | 4000 | 8000 | kHz | 577360 | 471.57 | 577330 | 470.19
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 590 | 17000 | 108.37 | 17479
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 108.37 | 17479
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 108.38 | 17479
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 109.37 | 17478
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 1000 | 2000 | kHz | 590 | 17163 | 204.79 | 17392
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 1000 | 2000 | kHz | 4683.4 | 13828 | 4325.4 | 13829
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 0 | 0 | kHz | 4.5 | 310 | 4.3849 | 312.12
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 0 | 0 | kHz | 4.5 | 312.16 | 4.3853 | 312.11
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 0 | 0 | kHz | 4.5 | 310.77 | 4.4226 | 310.57
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 0 | 0 | kHz | 7.2506 | 223.08 | 7.2466 | 222.31
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 0 | 0 | kHz | 57.903 | 32.603 | 57.901 | 32.045
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 0 | 0 | kHz | 577.37 | 3.4481 | 577.37 | 3.2224
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 0 | 1 | kHz | 3 | 160 | 2.9725 | 162.69
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 0 | 1 | kHz | 3 | 162.76 | 2.9731 | 162.67
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 0 | 1 | kHz | 3.0304 | 161.11 | 3.028 | 160.75
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 0 | 1 | kHz | 6.4968 | 89.969 | 6.4924 | 88.922
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 0 | 1 | kHz | 57.814 | 11.229 | 57.812 | 10.667
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 0 | 1 | kHz | 577.36 | 1.2948 | 577.36 | 1.0691
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 1 | 20 | kHz | 3.1 | 250 | 3.0556 | 254.12
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 1 | 20 | kHz | 3.1 | 254.16 | 3.0563 | 254.09
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 1 | 20 | kHz | 3.1119 | 251.96 | 3.1096 | 251.7
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 1 | 20 | kHz | 6.5335 | 152.12 | 6.5291 | 151.19
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 1 | 20 | kHz | 57.818 | 19.6 | 57.816 | 19.039
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 1 | 20 | kHz | 577.36 | 2.1354 | 577.36 | 1.9097
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 20 | 50 | kHz | 5.6 | 840 | 5.5358 | 839.05
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 20 | 50 | kHz | 5.6 | 839.02 | 5.5362 | 839.03
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 20 | 50 | kHz | 5.6 | 837.24 | 5.5656 | 837.27
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 20 | 50 | kHz | 7.9956 | 709.13 | 7.9861 | 708.96
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 20 | 50 | kHz | 58 | 140.31 | 57.994 | 139.7
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 20 | 50 | kHz | 577.38 | 14.459 | 577.37 | 14.2
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 50 | 100 | kHz | 8.8 | 5100 | 8.3103 | 5147.5
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 50 | 100 | kHz | 8.8 | 5147.5 | 8.3103 | 5147.5
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 50 | 100 | kHz | 8.8 | 5145.7 | 8.3291 | 5145.9
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 50 | 100 | kHz | 10.062 | 5001.3 | 10.046 | 5001.8
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 50 | 100 | kHz | 58.168 | 2497.7 | 58.157 | 2495.6
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 50 | 100 | kHz | 577.4 | 306.39 | 577.17 | 305.02
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.8932 | 46131
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.8935 | 46131
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 100 | 300 | kHz | 4.2005 | 46000 | 2.919 | 46128
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 100 | 300 | kHz | 4.9503 | 45929 | 4.8763 | 45936
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 100 | 300 | kHz | 54.13 | 41367 | 54.086 | 41369
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 100 | 300 | kHz | 575.77 | 16511 | 573.27 | 16507
3458A | ACV | 0 | 0.01 | 1e-08 | V | µV | 300 | 1000 | kHz | 18 | 45000 | 12.904 | 45496
3458A | ACV | 0 | 0.01 | 1e-07 | V | µV | 300 | 1000 | kHz | 18 | 45496 | 12.904 | 45496
3458A | ACV | 0 | 0.01 | 1e-06 | V | µV | 300 | 1000 | kHz | 18 | 45495 | 12.915 | 45495
3458A | ACV | 0 | 0.01 | 1e-05 | V | µV | 300 | 1000 | kHz | 18 | 45403 | 13.84 | 45406
3458A | ACV | 0 | 0.01 | 0.0001 | V | µV | 300 | 1000 | kHz | 56.207 | 41523 | 56.163 | 41525
3458A | ACV | 0 | 0.01 | 0.001 | V | µV | 300 | 1000 | kHz | 575.97 | 16722 | 573.47 | 16718
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 0 | 0 | kHz | 47 | 82 | 80.62 | 80.62
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 0 | 0 | kHz | 47 | 82 | 80.618 | 80.618
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 0 | 0 | kHz | 47 | 82 | 80.426 | 80.426
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 0 | 0 | kHz | 72.743 | 66.872 | 66.849 | 66.849
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 0 | 0 | kHz | 578.69 | 12.536 | 12.498 | 12.498
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 0 | 0 | kHz | 5773.7 | 1.2831 | 1.2674 | 1.2674
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 0 | 1 | kHz | 24 | 82 | 23.524 | 80.445
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 0 | 1 | kHz | 24 | 82 | 23.53 | 80.441
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 0 | 1 | kHz | 24 | 82 | 24.088 | 80.041
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 0 | 1 | kHz | 59.943 | 59.049 | 59.9 | 59.024
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 0 | 1 | kHz | 577.3 | 9.3986 | 577.28 | 9.3602
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 0 | 1 | kHz | 5773.6 | 0.9599 | 5772.7 | 0.94416
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.438 | 161.34
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.442 | 161.34
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.893 | 160.98
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 1 | 20 | kHz | 56.377 | 137.35 | 56.262 | 137.33
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 1 | 20 | kHz | 575.7 | 30.683 | 575.66 | 30.548
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.4 | 3.1917 | 5770.6 | 0.31603
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.379 | 346.14
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.383 | 346.14
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.693 | 345.87
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 20 | 50 | kHz | 49.713 | 324.48 | 49.371 | 324.63
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 20 | 50 | kHz | 568.73 | 117.17 | 568.57 | 116.91
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 20 | 50 | kHz | 5772.6 | 12.94 | 5762.2 | 12.802
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 50 | 100 | kHz | 27 | 940 | 23.141 | 938.86
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 50 | 100 | kHz | 27 | 935.56 | 23.148 | 938.85
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 50 | 100 | kHz | 27 | 935.62 | 23.298 | 938.72
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 50 | 100 | kHz | 38.506 | 925.39 | 37.913 | 925.82
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 50 | 100 | kHz | 530.37 | 592.16 | 529.89 | 592.07
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 50 | 100 | kHz | 5766.2 | 87.392 | 5715.5 | 87.123
3458A | ACV | 0.1 | 1 | 1e-07 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.46 | 3463.3
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.47 | 3463.3
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.55 | 3463.2
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 100 | 300 | kHz | 120 | 3500 | 120.41 | 3459.8
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 100 | 300 | kHz | 420.55 | 3205.9 | 419.31 | 3206.7
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 100 | 300 | kHz | 5681.4 | 1112.8 | 5471.4 | 1112.4
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
3458A | ACV | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
3458A | ACV | 0.1 | 1 | 1e-05 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
3458A | ACV | 0.1 | 1 | 0.0001 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 111.87 | 11674
3458A | ACV | 0.1 | 1 | 0.001 | V | µV | 300 | 1000 | kHz | 212.67 | 11907 | 246.94 | 11553
3458A | ACV | 0.1 | 1 | 0.01 | V | µV | 300 | 1000 | kHz | 5113.2 | 8011.6 | 4758 | 8011.9
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 1000 | 2000 | kHz | 53 | 17000 | 11.488 | 17414
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.488 | 17414
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.489 | 17414
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.588 | 17413
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 53 | 17065 | 21.129 | 17327
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 1000 | 2000 | kHz | 469.05 | 13766 | 433.19 | 13767
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.778 | 46223
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.778 | 46223
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.779 | 46223
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.809 | 46223
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 84.139 | 46193
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 2000 | 4000 | kHz | 354.54 | 43842 | 330.66 | 43842
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.256 | 46261
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.256 | 46261
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.257 | 46261
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.287 | 46261
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 95.542 | 46232
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 4000 | 8000 | kHz | 363.53 | 43909 | 338.44 | 43920
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 115.4 | 173280
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 8000 | 10000 | kHz | 450 | 170030 | 115.4 | 173280
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 115.4 | 173280
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 115.41 | 173280
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 116.4 | 173270
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 8000 | 10000 | kHz | 450 | 170150 | 203.58 | 172400
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 0 | 0 | kHz | 5.4 | 83 | 5.316 | 83.165
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 0 | 0 | kHz | 5.4 | 83.166 | 5.3163 | 83.163
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 0 | 0 | kHz | 5.4 | 83.01 | 5.3447 | 83
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 0 | 0 | kHz | 7.7286 | 70.835 | 7.7258 | 70.789
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 0 | 0 | kHz | 57.923 | 14.075 | 57.922 | 14.014
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 0 | 0 | kHz | 577.37 | 1.4498 | 577.25 | 1.4244
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 0 | 1 | kHz | 3.5 | 81 | 3.4429 | 80.684
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 0 | 1 | kHz | 3.5 | 80.684 | 3.4434 | 80.682
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 0 | 1 | kHz | 3.5 | 80.42 | 3.4848 | 80.411
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 0 | 1 | kHz | 6.538 | 63.487 | 6.5346 | 63.435
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 0 | 1 | kHz | 57.784 | 10.953 | 57.783 | 10.891
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 0 | 1 | kHz | 577.36 | 1.1266 | 577.26 | 1.1012
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 1 | 20 | kHz | 3.4 | 160 | 3.3551 | 157.72
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 1 | 20 | kHz | 3.4 | 157.72 | 3.3552 | 157.72
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 1 | 20 | kHz | 3.4 | 157.45 | 3.3911 | 157.45
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 1 | 20 | kHz | 6.2204 | 137.65 | 6.2175 | 137.62
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 1 | 20 | kHz | 57.628 | 32.01 | 57.626 | 31.951
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 1 | 20 | kHz | 577.34 | 3.3107 | 577.06 | 3.2853
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 20 | 50 | kHz | 4.7 | 330 | 4.6688 | 326.81
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 20 | 50 | kHz | 4.7 | 327.36 | 4.6691 | 326.81
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 20 | 50 | kHz | 4.7 | 327.1 | 4.6916 | 326.63
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 20 | 50 | kHz | 6.7172 | 310.87 | 6.7063 | 310.87
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 20 | 50 | kHz | 57.119 | 116.68 | 57.114 | 116.5
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 20 | 50 | kHz | 577.28 | 12.896 | 576.24 | 12.803
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 50 | 100 | kHz | 6 | 910 | 5.9729 | 906.03
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 50 | 100 | kHz | 6 | 906.02 | 5.973 | 906.03
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 50 | 100 | kHz | 6 | 905.89 | 5.9854 | 905.92
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 50 | 100 | kHz | 7.1698 | 895.95 | 7.1433 | 896.05
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 50 | 100 | kHz | 53.812 | 587.62 | 53.788 | 587.3
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 50 | 100 | kHz | 576.69 | 87.306 | 571.68 | 86.954
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.521 | 3476.1
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.521 | 3476.1
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.522 | 3476.1
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.911 | 3472.7
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 100 | 300 | kHz | 41.904 | 3219 | 41.802 | 3219.4
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 100 | 300 | kHz | 568.06 | 1120 | 547.01 | 1119.2
3458A | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
3458A | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
3458A | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
3458A | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.635 | 11630
3458A | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 300 | 1000 | kHz | 21.308 | 11907 | 25.139 | 11509
3458A | ACV | 0.01 | 0.1 | 0.001 | V | µV | 300 | 1000 | kHz | 511.75 | 7972.6 | 476.24 | 7972.4
3458A | CDE | -270 | 149 | 1 | °C | °C |  |  |  | 0.58502 | 0 | 0.57839 | 0
3458A | CDE | 149.001 | 593 | 1 | °C | °C |  |  |  | 0.57916 | 0 | 0.57842 | 0
3458A | CDE | 593.001 | 1000 | 1 | °C | °C |  |  |  | 0.57945 | 0 | 0.57871 | 0
3458A | CDE | -270 | 149 | 0.1 | °C | °C |  |  |  | 0.080416 | 0 | 0.067383 | 0
3458A | CDE | 149.001 | 593 | 0.1 | °C | °C |  |  |  | 0.069131 | 0 | 0.067566 | 0
3458A | CDE | 593.001 | 1000 | 0.1 | °C | °C |  |  |  | 0.071536 | 0 | 0.070024 | 0
3458A | CDJ | -210 | 259.999 | 1 | °C | °C |  |  |  | 0.59049 | 0 | 0.57878 | 0
3458A | CDJ | 260 | 732 | 1 | °C | °C |  |  |  | 0.58166 | 0 | 0.57849 | 0
3458A | CDJ | 732.001 | 1200 | 1 | °C | °C |  |  |  | 0.58132 | 0 | 0.57892 | 0
3458A | CDJ | -210 | 259.999 | 0.1 | °C | °C |  |  |  | 0.0917 | 0 | 0.070601 | 0
3458A | CDJ | 260 | 732 | 0.1 | °C | °C |  |  |  | 0.073715 | 0 | 0.068215 | 0
3458A | CDJ | 732.001 | 1200 | 0.1 | °C | °C |  |  |  | 0.076447 | 0 | 0.071753 | 0
3458A | CDK | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.74458 | 0 | 0.73021 | 0
3458A | CDK | -240 | 259.999 | 1 | °C | °C |  |  |  | 0.59718 | 0 | 0.58175 | 0
3458A | CDK | 260 | 816 | 1 | °C | °C |  |  |  | 0.58441 | 0 | 0.5787 | 0
3458A | CDK | 816.001 | 1371 | 1 | °C | °C |  |  |  | 0.58744 | 0 | 0.57956 | 0
3458A | CDK | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.45663 | 0 | 0.45079 | 0
3458A | CDK | -240 | 259.999 | 0.1 | °C | °C |  |  |  | 0.11392 | 0 | 0.091809 | 0
3458A | CDK | 260 | 816 | 0.1 | °C | °C |  |  |  | 0.08098 | 0 | 0.069993 | 0
3458A | CDK | 816.001 | 1371 | 0.1 | °C | °C |  |  |  | 0.090471 | 0 | 0.076709 | 0
3458A | CDN | -270 | -240.001 | 1 | °C | °C |  |  |  | 1.018 | 0 | 1.0084 | 0
3458A | CDN | -240 | -101.001 | 1 | °C | °C |  |  |  | 0.60402 | 0 | 0.58768 | 0
3458A | CDN | -101 | 371 | 1 | °C | °C |  |  |  | 0.59027 | 0 | 0.57855 | 0
3458A | CDN | 371.001 | 816 | 1 | °C | °C |  |  |  | 0.58489 | 0 | 0.57864 | 0
3458A | CDN | 816.001 | 1300 | 1 | °C | °C |  |  |  | 0.58599 | 0 | 0.57919 | 0
3458A | CDN | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.8317 | 0 | 0.82877 | 0
3458A | CDN | -240 | -101.001 | 0.1 | °C | °C |  |  |  | 0.14224 | 0 | 0.12394 | 0
3458A | CDN | -101 | 371 | 0.1 | °C | °C |  |  |  | 0.090241 | 0 | 0.068695 | 0
3458A | CDN | 371.001 | 816 | 0.1 | °C | °C |  |  |  | 0.081513 | 0 | 0.069464 | 0
3458A | CDN | 816.001 | 1300 | 0.1 | °C | °C |  |  |  | 0.086274 | 0 | 0.073915 | 0
3458A | CDPT100 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.008485 | 0 | 0.0076317 | 0
3458A | CDPT100 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.010087 | 0 | 0.0093806 | 0
3458A | CDPT100 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.012329 | 0 | 0.011526 | 0
3458A | CDPT100 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015887 | 0 | 0.015272 | 0
3458A | CDPT100 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019744 | 0 | 0.01888 | 0
3458A | CDPT100 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57756 | 0 | 0.57737 | 0
3458A | CDPT100 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57759 | 0 | 0.5774 | 0
3458A | CDPT100 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.5777 | 0 | 0.57744 | 0
3458A | CDPT100 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57752 | 0
3458A | CDPT100 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57809 | 0 | 0.57763 | 0
3458A | CDPT100 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058423 | 0 | 0.05795 | 0
3458A | CDPT100 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058677 | 0 | 0.058206 | 0
3458A | CDPT100 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.059241 | 0 | 0.058591 | 0
3458A | CDPT100 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060082 | 0 | 0.059441 | 0
3458A | CDPT100 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.061563 | 0 | 0.060469 | 0
3458A | CDPT1000 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.0069515 | 0 | 0.0065655 | 0
3458A | CDPT1000 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.0097193 | 0 | 0.0093614 | 0
3458A | CDPT1000 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.011768 | 0 | 0.011474 | 0
3458A | CDPT1000 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015792 | 0 | 0.015173 | 0
3458A | CDPT1000 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019042 | 0 | 0.018733 | 0
3458A | CDPT1000 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57743 | 0 | 0.57736 | 0
3458A | CDPT1000 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57749 | 0 | 0.5774 | 0
3458A | CDPT1000 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.57753 | 0 | 0.57744 | 0
3458A | CDPT1000 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57752 | 0
3458A | CDPT1000 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57763 | 0
3458A | CDPT1000 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058 | 0 | 0.05782 | 0
3458A | CDPT1000 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058438 | 0 | 0.058203 | 0
3458A | CDPT1000 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.058813 | 0 | 0.05858 | 0
3458A | CDPT1000 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060057 | 0 | 0.059416 | 0
3458A | CDPT1000 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.060808 | 0 | 0.060423 | 0
3458A | CDR | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60597 | 0 | 0.58986 | 0
3458A | CDR | 0 | 593 | 1 | °C | °C |  |  |  | 0.59803 | 0 | 0.5817 | 0
3458A | CDR | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59357 | 0 | 0.57971 | 0
3458A | CDR | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.595 | 0 | 0.58117 | 0
3458A | CDR | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.15085 | 0 | 0.13394 | 0
3458A | CDR | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.11485 | 0 | 0.091518 | 0
3458A | CDR | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.10063 | 0 | 0.077849 | 0
3458A | CDR | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.10872 | 0 | 0.088053 | 0
3458A | CDS | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60145 | 0 | 0.58521 | 0
3458A | CDS | 0 | 593 | 1 | °C | °C |  |  |  | 0.5968 | 0 | 0.58081 | 0
3458A | CDS | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59384 | 0 | 0.57998 | 0
3458A | CDS | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.59549 | 0 | 0.58168 | 0
3458A | CDS | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.13149 | 0 | 0.11169 | 0
3458A | CDS | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.10976 | 0 | 0.085663 | 0
3458A | CDS | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.10063 | 0 | 0.079868 | 0
3458A | CDS | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.11141 | 0 | 0.091359 | 0
3458A | CDT | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.63746 | 0 | 0.62062 | 0
3458A | CDT | -240 | 92.999 | 1 | °C | °C |  |  |  | 0.59575 | 0 | 0.58083 | 0
3458A | CDT | 93 | 400 | 1 | °C | °C |  |  |  | 0.58328 | 0 | 0.57831 | 0
3458A | CDT | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.2459 | 0 | 0.23488 | 0
3458A | CDT | -240 | 92.999 | 0.1 | °C | °C |  |  |  | 0.10841 | 0 | 0.08582 | 0
3458A | CDT | 93 | 400 | 0.1 | °C | °C |  |  |  | 0.076722 | 0 | 0.066645 | 0
3458A | DCA | 0 | 1e-07 | 1e-12 | A | nA |  |  |  | 0.048 | 34000 | 0.046246 | 34649
3458A | DCA | 0 | 1e-07 | 1e-11 | A | nA |  |  |  | 0.048 | 34000 | 0.046601 | 34403
3458A | DCA | 0 | 1e-07 | 1e-10 | A | nA |  |  |  | 0.074257 | 24027 | 0.07397 | 22146
3458A | DCA | 0 | 1e-07 | 1e-09 | A | nA |  |  |  | 0.57935 | 3897.9 | 0.5792 | 2870.4
3458A | DCA | 0 | 1e-07 | 1e-08 | A | nA |  |  |  | 5.7738 | 700.9 | 5.7737 | 288.02
3458A | DCA | 0 | 1e-07 | 1e-07 | A | nA |  |  |  | 57.736 | 193.97 | 57.736 | 28.803
3458A | DCA | 0 | 1e-07 | 1e-06 | A | nA |  |  |  | 577.36 | 68.945 | 577.36 | 2.8803
3458A | DCA | 0 | 1e-07 | 1e-05 | A | nA |  |  |  | 5773.6 | 26.714 | 5773.6 | 0.28803
3458A | DCA | 0 | 1e-07 | 0.0001 | A | nA |  |  |  | 57736 | 10.599 | 57736 | 0.028736
3458A | DCA | 1e-07 | 1e-06 | 1e-12 | A | nA |  |  |  | 0.048 | 23000 | 0.046073 | 24826
3458A | DCA | 1e-07 | 1e-06 | 1e-11 | A | nA |  |  |  | 0.048 | 23271 | 0.046424 | 24708
3458A | DCA | 1e-07 | 1e-06 | 1e-10 | A | nA |  |  |  | 0.074091 | 17690 | 0.07366 | 17772
3458A | DCA | 1e-07 | 1e-06 | 1e-09 | A | nA |  |  |  | 0.57936 | 2553.2 | 0.57914 | 2554.1
3458A | DCA | 1e-07 | 1e-06 | 1e-08 | A | nA |  |  |  | 5.7738 | 256.81 | 5.7737 | 256.81
3458A | DCA | 1e-07 | 1e-06 | 1e-07 | A | nA |  |  |  | 57.736 | 25.682 | 57.736 | 25.682
3458A | DCA | 1e-06 | 1e-05 | 1e-12 | A | nA |  |  |  | 0.13 | 25000 | 0.11375 | 25625
3458A | DCA | 1e-06 | 1e-05 | 1e-11 | A | nA |  |  |  | 0.13 | 25119 | 0.11389 | 25617
3458A | DCA | 1e-06 | 1e-05 | 1e-10 | A | nA |  |  |  | 0.13 | 24832 | 0.12602 | 24846
3458A | DCA | 1e-06 | 1e-05 | 1e-09 | A | nA |  |  |  | 0.58396 | 10197 | 0.58374 | 10200
3458A | DCA | 1e-06 | 1e-05 | 1e-08 | A | nA |  |  |  | 5.7742 | 1129.1 | 5.7741 | 1129.1
3458A | DCA | 1e-06 | 1e-05 | 1e-07 | A | nA |  |  |  | 57.736 | 113.04 | 57.735 | 113.04
3458A | DCA | 1e-05 | 0.0001 | 1e-11 | A | nA |  |  |  | 0.16 | 23000 | 0.11411 | 23000
3458A | DCA | 1e-05 | 0.0001 | 1e-10 | A | nA |  |  |  | 0.16 | 23403 | 0.11928 | 23403
3458A | DCA | 1e-05 | 0.0001 | 1e-09 | A | nA |  |  |  | 0.47124 | 20574 | 0.46866 | 20574
3458A | DCA | 1e-05 | 0.0001 | 1e-08 | A | nA |  |  |  | 5.7304 | 5486.9 | 5.7294 | 5486.9
3458A | DCA | 1e-05 | 0.0001 | 1e-07 | A | nA |  |  |  | 57.731 | 581.51 | 57.731 | 581.51
3458A | DCA | 1e-05 | 0.0001 | 1e-06 | A | nA |  |  |  | 577.35 | 61.585 | 577.35 | 61.585
3458A | DCA | 0.0001 | 0.001 | 1e-10 | A | nA |  |  |  | 6.3 | 23000 | 5.7456 | 23552
3458A | DCA | 0.0001 | 0.001 | 1e-09 | A | nA |  |  |  | 6.3 | 23340 | 5.7675 | 23536
3458A | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 7.7388 | 22133 | 7.7351 | 22126
3458A | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 57.586 | 7177.2 | 57.585 | 7158.5
3458A | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 577.34 | 770.78 | 577.34 | 762.27
3458A | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.6 | 79.686 | 5773.5 | 76.28
3458A | DCA | 0.001 | 0.01 | 1e-09 | A | µA |  |  |  | 0.064 | 23 | 0.057395 | 23.593
3458A | DCA | 0.001 | 0.01 | 1e-08 | A | µA |  |  |  | 0.064 | 23.49 | 0.057621 | 23.576
3458A | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 0.07733 | 22.173 | 0.077294 | 22.166
3458A | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 0.57584 | 7.1949 | 0.57583 | 7.1763
3458A | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.7734 | 0.77277 | 5.7734 | 0.76426
3458A | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.736 | 0.079885 | 57.735 | 0.076479
3458A | DCA | 0.01 | 0.1 | 1e-08 | A | µA |  |  |  | 0.67 | 40 | 0.57451 | 40.858
3458A | DCA | 0.01 | 0.1 | 1e-07 | A | µA |  |  |  | 0.67 | 40.354 | 0.57629 | 40.844
3458A | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 0.75326 | 39.474 | 0.74494 | 39.51
3458A | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 5.6886 | 17.431 | 5.683 | 17.368
3458A | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 57.726 | 2.0334 | 57.724 | 1.9935
3458A | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.35 | 0.21573 | 577.34 | 0.19968
3458A | DCA | 0.1 | 1 | 1e-07 | A | µA |  |  |  | 12 | 130 | 11.536 | 127.22
3458A | DCA | 0.1 | 1 | 1e-06 | A | µA |  |  |  | 12 | 130 | 11.548 | 127.21
3458A | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 12 | 130 | 12.277 | 126.6
3458A | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 52.949 | 97.342 | 52.884 | 97.405
3458A | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 576.12 | 17.674 | 576.1 | 17.7
3458A | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5773.4 | 1.7853 | 5771.8 | 1.7958
3458A | DCA | 1 | 20 | 1e-06 | A | mA |  |  |  | 0.068 | 0.0036 | 0.043 | 0.0047621
3458A | DCA | 1 | 20 | 1e-05 | A | mA |  |  |  | 0.068 | 0.0036 | 0.043318 | 0.0047475
3458A | DCA | 1 | 20 | 0.0001 | A | mA |  |  |  | 0.070893 | 0.0038965 | 0.068928 | 0.0038258
3458A | DCA | 1 | 20 | 0.001 | A | mA |  |  |  | 0.57834 | 0.00081073 | 0.57809 | 0.0006992
3458A | DCA | 1 | 20 | 0.01 | A | mA |  |  |  | 5.7739 | 0.00011677 | 5.7736 | 7.0888e-05
3458A | DCA | 1 | 20 | 0.1 | A | mA |  |  |  | 57.736 | 2.5449e-05 | 57.735 | 7.09e-06
3458A | DCA | 0 | 1e-07 | 0.001 | A | nA |  |  |  | 577360 | 4.2309 | 577360 | 0
3458A | DCA | 0 | 1e-07 | 0.01 | A | nA |  |  |  | 5773600 | 1.6933 | 5773600 | 0
3458A | DCA | 0 | 1e-07 | 0.1 | A | nA |  |  |  | 57736000 | 0.75258 | 57736000 | 0
3458A | DCV | 0 | 0.1 | 1e-08 | V | µV |  |  |  | 0.62 | 8 | 0.60546 | 8.1092
3458A | DCV | 0 | 0.1 | 1e-07 | V | µV |  |  |  | 0.62 | 8.0238 | 0.60819 | 8.0937
3458A | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 0.83938 | 6.9596 | 0.83737 | 6.9316
3458A | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 5.8321 | 1.4327 | 5.831 | 1.3947
3458A | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.004 | 0.15763 | 58.003 | 0.14188
3458A | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.020493 | 580 | 0.014191
3458A | DCV | 0 | 0.1 | 0.01 | V | mV |  |  |  | 5.8 | 3.94e-06 | 5.7999 | 1.42e-06
3458A | DCV | 0 | 0.1 | 0.1 | V | mV |  |  |  | 58 | 1.15e-06 | 58 | 1.42e-07
3458A | DCV | 0 | 0.1 | 1 | V | mV |  |  |  | 580 | 4.18e-07 | 580 | 1.42e-08
3458A | DCV | 0.1 | 1 | 1e-08 | V | µV |  |  |  | 0.44 | 6 | 0.43733 | 5.9926
3458A | DCV | 0.1 | 1 | 1e-07 | V | µV |  |  |  | 0.44094 | 5.99 | 0.43909 | 5.9911
3458A | DCV | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 0.60901 | 5.85 | 0.60244 | 5.8536
3458A | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 5.589 | 3.0792 | 5.5845 | 3.0749
3458A | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 57.973 | 0.3881 | 57.971 | 0.38452
3458A | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 580 | 0.040018 | 579.96 | 0.038571
3458A | DCV | 1 | 10 | 1e-07 | V | µV |  |  |  | 1.2 | 5.3 | 0.75905 | 5.3437
3458A | DCV | 1 | 10 | 1e-06 | V | µV |  |  |  | 1.2 | 5.3409 | 0.78898 | 5.341
3458A | DCV | 1 | 10 | 1e-05 | V | µV |  |  |  | 3.301 | 5.1205 | 3.2984 | 5.1207
3458A | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 55.981 | 2.3401 | 55.98 | 2.34
3458A | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 579.76 | 0.27724 | 579.76 | 0.27716
3458A | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5800 | 0.027808 | 5799.7 | 0.027777
3458A | DCV | 10 | 100 | 1e-06 | V | µV |  |  |  | 34 | 7.8 | 33.132 | 7.7866
3458A | DCV | 10 | 100 | 1e-05 | V | µV |  |  |  | 34 | 7.7848 | 33.3 | 7.7851
3458A | DCV | 10 | 100 | 0.0001 | V | µV |  |  |  | 48.99 | 7.6505 | 48.723 | 7.6514
3458A | DCV | 10 | 100 | 0.001 | V | µV |  |  |  | 545.46 | 4.5283 | 545.28 | 4.5242
3458A | DCV | 10 | 100 | 0.01 | V | µV |  |  |  | 5795 | 0.61987 | 5794.9 | 0.61637
3458A | DCV | 100 | 1000 | 1e-05 | V | mV |  |  |  | 0.11 | 0.0087 | 0.10064 | 0.0086547
3458A | DCV | 100 | 1000 | 0.0001 | V | mV |  |  |  | 0.11 | 0.0086529 | 0.10253 | 0.008653
3458A | DCV | 100 | 1000 | 0.001 | V | mV |  |  |  | 0.27989 | 0.0084976 | 0.27707 | 0.0084975
3458A | DCV | 100 | 1000 | 0.01 | V | mV |  |  |  | 5.3677 | 0.0051444 | 5.3663 | 0.0051359
3458A | DCV | 100 | 1000 | 0.1 | V | mV |  |  |  | 57.936 | 0.00072803 | 57.936 | 0.00072118
3458A | DCV | 100 | 1000 | 1 | V | mV |  |  |  | 579.99 | 7.53e-05 | 579.93 | 7.2528e-05
3458A | FDE | -454 | 300 | 1 | °F | °F |  |  |  | 0.60177 | 0 | 0.58066 | 0
3458A | FDE | 300.001 | 1100 | 1 | °F | °F |  |  |  | 0.58299 | 0 | 0.5806 | 0
3458A | FDE | 1100.001 | 1832 | 1 | °F | °F |  |  |  | 0.58352 | 0 | 0.58114 | 0
3458A | FDE | -454 | 300 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.08462 | 0
3458A | FDE | 300.001 | 1100 | 0.1 | °F | °F |  |  |  | 0.09 | 0 | 0.084258 | 0
3458A | FDE | 1100.001 | 1832 | 0.1 | °F | °F |  |  |  | 0.091751 | 0 | 0.08789 | 0
3458A | FDJ | -346 | 499.999 | 1 | °F | °F |  |  |  | 0.61872 | 0 | 0.58175 | 0
3458A | FDJ | 500 | 1350 | 1 | °F | °F |  |  |  | 0.59113 | 0 | 0.58077 | 0
3458A | FDJ | 1350.001 | 2192 | 1 | °F | °F |  |  |  | 0.58927 | 0 | 0.58157 | 0
3458A | FDJ | -346 | 499.999 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.091814 | 0
3458A | FDJ | 500 | 1350 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.085419 | 0
3458A | FDJ | 1350.001 | 2192 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.090692 | 0
3458A | FDK | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.99057 | 0 | 0.95529 | 0
3458A | FDK | -400 | 499.999 | 1 | °F | °F |  |  |  | 0.63828 | 0 | 0.59031 | 0
3458A | FDK | 500 | 1500 | 1 | °F | °F |  |  |  | 0.59947 | 0 | 0.58127 | 0
3458A | FDK | 1500.001 | 2500 | 1 | °F | °F |  |  |  | 0.60796 | 0 | 0.58292 | 0
3458A | FDK | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.77443 | 0 | 0.76327 | 0
3458A | FDK | -400 | 499.999 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13588 | 0
3458A | FDK | 500 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.088716 | 0
3458A | FDK | 1500.001 | 2500 | 0.1 | °F | °F |  |  |  | 0.13135 | 0 | 0.098982 | 0
3458A | FDN | -454 | -400.001 | 1 | °F | °F |  |  |  | 1.5603 | 0 | 1.54 | 0
3458A | FDN | -400 | -150.001 | 1 | °F | °F |  |  |  | 0.65772 | 0 | 0.60786 | 0
3458A | FDN | -150 | 700 | 1 | °F | °F |  |  |  | 0.61817 | 0 | 0.58117 | 0
3458A | FDN | 700.001 | 1500 | 1 | °F | °F |  |  |  | 0.60109 | 0 | 0.58115 | 0
3458A | FDN | 1500.001 | 2372 | 1 | °F | °F |  |  |  | 0.60386 | 0 | 0.58223 | 0
3458A | FDN | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 1.5 | 0 | 1.4288 | 0
3458A | FDN | -400 | -150.001 | 0.1 | °F | °F |  |  |  | 0.24 | 0 | 0.19873 | 0
3458A | FDN | -150 | 700 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.088083 | 0
3458A | FDN | 700.001 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.087954 | 0
3458A | FDN | 1500.001 | 2372 | 0.1 | °F | °F |  |  |  | 0.13 | 0 | 0.094806 | 0
3458A | FDPT100 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.013 | 0 | 0.010679 | 0
3458A | FDPT100 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014506 | 0
3458A | FDPT100 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.021 | 0 | 0.018862 | 0
3458A | FDPT100 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.027261 | 0 | 0.026097 | 0
3458A | FDPT100 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.034473 | 0 | 0.032867 | 0
3458A | FDPT100 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57804 | 0 | 0.57742 | 0
3458A | FDPT100 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57812 | 0 | 0.5775 | 0
3458A | FDPT100 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57849 | 0 | 0.57763 | 0
3458A | FDPT100 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57877 | 0 | 0.57791 | 0
3458A | FDPT100 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57975 | 0 | 0.57826 | 0
3458A | FDPT100 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.059936 | 0 | 0.05843 | 0
3458A | FDPT100 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.060734 | 0 | 0.059249 | 0
3458A | FDPT100 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.062482 | 0 | 0.060463 | 0
3458A | FDPT100 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.065034 | 0 | 0.063096 | 0
3458A | FDPT100 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.069375 | 0 | 0.066183 | 0
3458A | FDPT1000 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.0090498 | 0 | 0.0080621 | 0
3458A | FDPT1000 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014466 | 0
3458A | FDPT1000 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.02 | 0 | 0.018759 | 0
3458A | FDPT1000 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.02708 | 0 | 0.025909 | 0
3458A | FDPT1000 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.033169 | 0 | 0.032594 | 0
3458A | FDPT1000 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57761 | 0 | 0.57738 | 0
3458A | FDPT1000 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57781 | 0 | 0.5775 | 0
3458A | FDPT1000 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57793 | 0 | 0.57763 | 0
3458A | FDPT1000 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57876 | 0 | 0.5779 | 0
3458A | FDPT1000 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57876 | 0 | 0.57824 | 0
3458A | FDPT1000 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.058588 | 0 | 0.058009 | 0
3458A | FDPT1000 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.059981 | 0 | 0.059239 | 0
3458A | FDPT1000 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.061159 | 0 | 0.060431 | 0
3458A | FDPT1000 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.064958 | 0 | 0.063018 | 0
3458A | FDPT1000 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.067183 | 0 | 0.066048 | 0
3458A | FDR | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.66465 | 0 | 0.61591 | 0
3458A | FDR | 32 | 1100 | 1 | °F | °F |  |  |  | 0.64166 | 0 | 0.59103 | 0
3458A | FDR | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.6276 | 0 | 0.5841 | 0
3458A | FDR | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.63063 | 0 | 0.58736 | 0
3458A | FDR | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.26 | 0 | 0.22215 | 0
3458A | FDR | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13898 | 0
3458A | FDR | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.16 | 0 | 0.10571 | 0
3458A | FDR | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.12242 | 0
3458A | FDS | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.6516 | 0 | 0.6018 | 0
3458A | FDS | 32 | 1100 | 1 | °F | °F |  |  |  | 0.63802 | 0 | 0.58826 | 0
3458A | FDS | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.62833 | 0 | 0.5849 | 0
3458A | FDS | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.63063 | 0 | 0.58878 | 0
3458A | FDS | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.22 | 0 | 0.17936 | 0
3458A | FDS | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12667 | 0
3458A | FDS | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.10571 | 0
3458A | FDS | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.12242 | 0
3458A | FDT | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.74355 | 0 | 0.69586 | 0
3458A | FDT | -400 | 199.999 | 1 | °F | °F |  |  |  | 0.63425 | 0 | 0.58771 | 0
3458A | FDT | 200 | 752 | 1 | °F | °F |  |  |  | 0.59635 | 0 | 0.58043 | 0
3458A | FDT | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.41398 | 0 | 0.39271 | 0
3458A | FDT | -400 | 199.999 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12412 | 0
3458A | FDT | 200 | 752 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.083094 | 0
3458A | Frequency | 0 | 40 | 1e-06 | Hz | Hz |  |  |  | 0 | 6e-10 | 1.1e-11 | 5.77e-10
3458A | Frequency | 0 | 40 | 1e-05 | Hz | Hz |  |  |  | 0 | 6e-10 | 2.97e-11 | 5.77e-10
3458A | Frequency | 0 | 40 | 0.0001 | Hz | Hz |  |  |  | 0 | 6e-10 | 2.95e-09 | 5.77e-10
3458A | Frequency | 0 | 40 | 0.001 | Hz | Hz |  |  |  | 2.23e-07 | 5.94e-10 | 2.45e-07 | 5.71e-10
3458A | Frequency | 0 | 40 | 0.01 | Hz | Hz |  |  |  | 5.34e-06 | 4.67e-10 | 5.34e-06 | 4.62e-10
3458A | Frequency | 0 | 40 | 0.1 | Hz | Hz |  |  |  | 5.7625e-05 | 1.14e-10 | 5.7277e-05 | 1.14e-10
3458A | Frequency | 0 | 40 | 1 | Hz | Hz |  |  |  | 0.00057734 | 1.18e-11 | 0.00057689 | 1.18e-11
3458A | Frequency | 0 | 40 | 10 | Hz | Hz |  |  |  | 0.0057736 | 1.18e-12 | 0.0057731 | 1.18e-12
3458A | Frequency | 0 | 40 | 100 | Hz | Hz |  |  |  | 0.057736 | 1.18e-13 | 0.057735 | 1.18e-13
3458A | Frequency | 0 | 40 | 1000 | Hz | Hz |  |  |  | 0.57736 | 1.18e-14 | 0.57735 | 1.18e-14
3458A | Frequency | 0 | 40 | 10000 | Hz | Hz |  |  |  | 5.7736 | 1.18e-15 | 5.7736 | 1.18e-15
3458A | Frequency | 40 | 1000 | 0.0001 | Hz | Hz |  |  |  | 2.9e-07 | 1.3e-07 | 2.53e-07 | 1.2897e-07
3458A | Frequency | 40 | 1000 | 0.001 | Hz | Hz |  |  |  | 6.57e-07 | 1.2894e-07 | 5.98e-07 | 1.2907e-07
3458A | Frequency | 40 | 1000 | 0.01 | Hz | Hz |  |  |  | 2.609e-05 | 1.3885e-07 | 2.5942e-05 | 1.3918e-07
3458A | Frequency | 40 | 1000 | 0.1 | Hz | Hz |  |  |  | 0.00039689 | 5.557e-07 | 0.00039681 | 5.5588e-07
3458A | Frequency | 40 | 1000 | 1 | Hz | Hz |  |  |  | 0.0039872 | 5.4141e-06 | 0.0039872 | 5.4141e-06
3458A | Frequency | 40 | 1000 | 10 | Hz | Hz |  |  |  | 0.055949 | 5.4127e-05 | 0.055949 | 5.4127e-05
3458A | Frequency | 1000 | 10000 | 0.001 | Hz | Hz |  |  |  | 4.4e-09 | 1.3e-07 | 7.41e-10 | 1.2922e-07
3458A | Frequency | 1000 | 10000 | 0.01 | Hz | Hz |  |  |  | 1.95e-07 | 1.292e-07 | 1.38e-07 | 1.2921e-07
3458A | Frequency | 1000 | 10000 | 0.1 | Hz | Hz |  |  |  | 1.3574e-05 | 1.2799e-07 | 1.3571e-05 | 1.2799e-07
3458A | Frequency | 1000 | 10000 | 1 | Hz | Hz |  |  |  | 0.00068566 | 9.1517e-08 | 0.00068565 | 9.1518e-08
3458A | Frequency | 1000 | 10000 | 10 | Hz | Hz |  |  |  | 0.0057844 | 1.571e-08 | 0.0057844 | 1.571e-08
3458A | Frequency | 1000 | 10000 | 100 | Hz | Hz |  |  |  | 0.057734 | 1.59e-09 | 0.05772 | 1.59e-09
3458A | Frequency | 10000 | 100000 | 0.01 | Hz | Hz |  |  |  | 4.4e-08 | 7.41e-12 | 7.41e-09 | 1.2922e-07
3458A | Frequency | 10000 | 100000 | 0.1 | Hz | Hz |  |  |  | 1.95e-06 | 1.38e-09 | 1.38e-06 | 1.2921e-07
3458A | Frequency | 10000 | 100000 | 1 | Hz | Hz |  |  |  | 0.00013574 | 1.3571e-07 | 0.00013571 | 1.2799e-07
3458A | Frequency | 10000 | 100000 | 10 | Hz | Hz |  |  |  | 0.0068566 | 6.8565e-06 | 0.0068565 | 9.1518e-08
3458A | Frequency | 10000 | 100000 | 100 | Hz | Hz |  |  |  | 0.057844 | 5.7844e-05 | 0.057844 | 1.571e-08
3458A | Frequency | 10000 | 100000 | 1000 | Hz | Hz |  |  |  | 0.57734 | 0.0005772 | 0.5772 | 1.59e-09
3458A | Frequency | 10000 | 100000 | 10000 | Hz | Hz |  |  |  | 5.7736 | 0.0057734 | 5.7734 | 1.59e-10
3458A | Frequency | 100000 | 1000000 | 0.1 | Hz | Hz |  |  |  | 4.4e-07 | 1.3e-07 | 7.41e-08 | 1.2922e-07
3458A | Frequency | 100000 | 1000000 | 1 | Hz | Hz |  |  |  | 1.9518e-05 | 1.292e-07 | 1.3837e-05 | 1.2921e-07
3458A | Frequency | 100000 | 1000000 | 10 | Hz | Hz |  |  |  | 0.0013574 | 1.2799e-07 | 0.0013571 | 1.2799e-07
3458A | Frequency | 100000 | 1000000 | 100 | Hz | Hz |  |  |  | 0.068566 | 9.1517e-08 | 0.068565 | 9.1518e-08
3458A | Frequency | 100000 | 1000000 | 1000 | Hz | Hz |  |  |  | 0.57844 | 1.571e-08 | 0.57844 | 1.571e-08
3458A | Frequency | 100000 | 1000000 | 10000 | Hz | Hz |  |  |  | 5.7734 | 1.59e-09 | 5.772 | 1.59e-09
3458A | Frequency | 1000000 | 10000000 | 1 | Hz | Hz |  |  |  | 4.4e-06 | 1.3e-07 | 7.41e-07 | 1.2922e-07
3458A | Frequency | 1000000 | 10000000 | 10 | Hz | Hz |  |  |  | 0.00019518 | 1.292e-07 | 0.00013837 | 1.2921e-07
3458A | Frequency | 1000000 | 10000000 | 100 | Hz | Hz |  |  |  | 0.013574 | 1.2799e-07 | 0.013571 | 1.2799e-07
3458A | Frequency | 1000000 | 10000000 | 1000 | Hz | Hz |  |  |  | 0.68566 | 9.1517e-08 | 0.68565 | 9.1518e-08
3458A | Frequency | 1000000 | 10000000 | 10000 | Hz | Hz |  |  |  | 5.7844 | 1.571e-08 | 5.7844 | 1.571e-08
3458A | Frequency | 1000000 | 10000000 | 100000 | Hz | Hz |  |  |  | 57.734 | 1.59e-09 | 57.72 | 1.59e-09
3458A | Frequency | 40 | 10000000 | 0 | Hz | Hz |  |  |  | 0 | 1.2e-07 | 0 | 1.1561e-07
3458A | Resistance | 0 | 10 | 1e-05 | Ω | µΩ |  |  |  | 65 | 17 | 61.668 | 17.321
3458A | Resistance | 0 | 10 | 0.0001 | Ω | µΩ |  |  |  | 81.734 | 16.018 | 81.655 | 16.015
3458A | Resistance | 0 | 10 | 0.001 | Ω | µΩ |  |  |  | 578.24 | 4.5207 | 578.2 | 4.5079
3458A | Resistance | 0 | 10 | 0.01 | Ω | µΩ |  |  |  | 5773.6 | 0.47627 | 5773.6 | 0.47059
3458A | Resistance | 0 | 10 | 0.1 | Ω | µΩ |  |  |  | 57736 | 0.049352 | 57736 | 0.047081
3458A | Resistance | 0 | 10 | 1 | Ω | µΩ |  |  |  | 577360 | 0.0056165 | 577360 | 0.0047081
3458A | Resistance | 0 | 10 | 10 | Ω | µΩ |  |  |  | 5773600 | 0.00083417 | 5773600 | 0.00047081
3458A | Resistance | 0 | 10 | 100 | Ω | µΩ |  |  |  | 57736000 | 0.00019242 | 57736000 | 4.708e-05
3458A | Resistance | 0 | 10 | 1000 | Ω | µΩ |  |  |  | 577360000 | 6.285e-05 | 577360000 | 4.7021e-06
3458A | Resistance | 10 | 100 | 1e-05 | Ω | µΩ |  |  |  | 580 | 14 | 577.27 | 0
3458A | Resistance | 10 | 100 | 0.0001 | Ω | µΩ |  |  |  | 580 | 13.975 | 579.78 | 13.975
3458A | Resistance | 10 | 100 | 0.001 | Ω | µΩ |  |  |  | 794.35 | 12.648 | 794.24 | 12.648
3458A | Resistance | 10 | 100 | 0.01 | Ω | µΩ |  |  |  | 5786.4 | 3.1626 | 5786.3 | 3.1617
3458A | Resistance | 10 | 100 | 0.1 | Ω | µΩ |  |  |  | 57737 | 0.32669 | 57737 | 0.32628
3458A | Resistance | 10 | 100 | 1 | Ω | µΩ |  |  |  | 577360 | 0.032802 | 577350 | 0.032639
3458A | Resistance | 100 | 1000 | 0.0001 | Ω | µΩ |  |  |  | 590 | 12 | 568.01 | 11.812
3458A | Resistance | 100 | 1000 | 0.001 | Ω | µΩ |  |  |  | 647.66 | 11.942 | 669.04 | 11.724
3458A | Resistance | 100 | 1000 | 0.01 | Ω | µΩ |  |  |  | 5185.8 | 8.4755 | 5185 | 8.4748
3458A | Resistance | 100 | 1000 | 0.1 | Ω | µΩ |  |  |  | 57619 | 1.4299 | 57619 | 1.4287
3458A | Resistance | 100 | 1000 | 1 | Ω | µΩ |  |  |  | 577340 | 0.14501 | 577340 | 0.14451
3458A | Resistance | 100 | 1000 | 10 | Ω | µΩ |  |  |  | 5773600 | 0.014653 | 5773500 | 0.014453
3458A | Resistance | 1000 | 10000 | 0.001 | Ω | mΩ |  |  |  | 5.9 | 0.012 | 5.8778 | 0.011642
3458A | Resistance | 1000 | 10000 | 0.01 | Ω | mΩ |  |  |  | 6.5111 | 0.011939 | 6.8827 | 0.011555
3458A | Resistance | 1000 | 10000 | 0.1 | Ω | mΩ |  |  |  | 52.024 | 0.0083239 | 52.009 | 0.0083231
3458A | Resistance | 1000 | 10000 | 1 | Ω | mΩ |  |  |  | 576.23 | 0.0013955 | 576.23 | 0.0013939
3458A | Resistance | 1000 | 10000 | 10 | Ω | mΩ |  |  |  | 5773.4 | 0.0001416 | 5773.4 | 0.00014095
3458A | Resistance | 1000 | 10000 | 100 | Ω | mΩ |  |  |  | 57736 | 1.4357e-05 | 57735 | 1.4097e-05
3458A | Resistance | 10000 | 100000 | 0.01 | Ω | mΩ |  |  |  | 58 | 0.012 | 53.873 | 0.012027
3458A | Resistance | 10000 | 100000 | 0.1 | Ω | mΩ |  |  |  | 64.108 | 0.011939 | 63.986 | 0.011939
3458A | Resistance | 10000 | 100000 | 1 | Ω | mΩ |  |  |  | 516.51 | 0.0086656 | 516.37 | 0.0086649
3458A | Resistance | 10000 | 100000 | 10 | Ω | mΩ |  |  |  | 5761.5 | 0.0014742 | 5761.5 | 0.0014726
3458A | Resistance | 10000 | 100000 | 100 | Ω | mΩ |  |  |  | 57734 | 0.00014965 | 57734 | 0.000149
3458A | Resistance | 10000 | 100000 | 1000 | Ω | mΩ |  |  |  | 577360 | 1.5162e-05 | 577350 | 1.4902e-05
3458A | Resistance | 100000 | 1000000 | 0.1 | Ω | Ω |  |  |  | 2.3 | 1.8e-05 | 2.2687 | 1.7871e-05
3458A | Resistance | 100000 | 1000000 | 1 | Ω | Ω |  |  |  | 2.3 | 1.8e-05 | 2.3129 | 1.7835e-05
3458A | Resistance | 100000 | 1000000 | 10 | Ω | Ω |  |  |  | 5.5163 | 1.5438e-05 | 5.5116 | 1.5439e-05
3458A | Resistance | 100000 | 1000000 | 100 | Ω | Ω |  |  |  | 57.516 | 3.6361e-06 | 57.515 | 3.6327e-06
3458A | Resistance | 100000 | 1000000 | 1000 | Ω | Ω |  |  |  | 577.33 | 3.7584e-07 | 577.33 | 3.7433e-07
3458A | Resistance | 100000 | 1000000 | 10000 | Ω | Ω |  |  |  | 5773.6 | 3.8048e-08 | 5773.5 | 3.7445e-08
3458A | Resistance | 1000000 | 10000000 | 1 | Ω | Ω |  |  |  | 120 | 5.9e-05 | 115.5 | 5.7739e-05
3458A | Resistance | 1000000 | 10000000 | 10 | Ω | Ω |  |  |  | 120 | 5.9e-05 | 115.61 | 5.7731e-05
3458A | Resistance | 1000000 | 10000000 | 100 | Ω | Ω |  |  |  | 124.34 | 5.8567e-05 | 125.65 | 5.6965e-05
3458A | Resistance | 1000000 | 10000000 | 1000 | Ω | Ω |  |  |  | 569.84 | 3.3313e-05 | 569.55 | 3.3236e-05
3458A | Resistance | 1000000 | 10000000 | 10000 | Ω | Ω |  |  |  | 5771.9 | 4.3708e-06 | 5771.8 | 4.3145e-06
3458A | Resistance | 1000000 | 10000000 | 100000 | Ω | Ω |  |  |  | 57735 | 4.5579e-07 | 57731 | 4.3308e-07
3458A | Resistance | 10000000 | 100000000 | 10 | Ω | kΩ |  |  |  | 1.2 | 5.8e-07 | 1.1508 | 5.778e-07
3458A | Resistance | 10000000 | 100000000 | 100 | Ω | kΩ |  |  |  | 1.2 | 5.8e-07 | 1.1511 | 5.778e-07
3458A | Resistance | 10000000 | 100000000 | 1000 | Ω | kΩ |  |  |  | 1.2 | 5.8e-07 | 1.1777 | 5.7756e-07
3458A | Resistance | 10000000 | 100000000 | 10000 | Ω | kΩ |  |  |  | 3.4529 | 5.5768e-07 | 3.442 | 5.5771e-07
3458A | Resistance | 10000000 | 100000000 | 100000 | Ω | kΩ |  |  |  | 55.449 | 2.7074e-07 | 55.444 | 2.7056e-07
3458A | Resistance | 10000000 | 100000000 | 10000000 | Ω | kΩ |  |  |  | 5773.5 | 3.31e-09 | 5770.9 | 3.3e-09
3458A | Resistance | 100000000 | 1000000000 | 100 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 11.592 | 5.7735e-06
3458A | Resistance | 100000000 | 1000000000 | 1000 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 11.592 | 5.7735e-06
3458A | Resistance | 100000000 | 1000000000 | 10000 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 11.595 | 5.7707e-06
3458A | Resistance | 100000000 | 1000000000 | 100000 | Ω | kΩ |  |  |  | 12 | 5.8e-06 | 14.68 | 5.7707e-06
3458A | Resistance | 100000000 | 1000000000 | 10000000 | Ω | kΩ |  |  |  | 5540.2 | 2.6331e-06 | 5540.2 | 2.633e-06
3458A | Resistance | 100000000 | 1000000000 | 100000000 | Ω | kΩ |  |  |  | 57707 | 3.1791e-07 | 57475 | 3.179e-07
3458A | Resistance | 0 | 10 | 10000 | Ω | µΩ |  |  |  | 5773600000 | 2.3736e-05 | 5773600000 | 0
3458A | Resistance | 0 | 10 | 100000 | Ω | µΩ |  |  |  | 57736000000 | 8.4771e-06 | 57736000000 | 0
3458A | Resistance | 0 | 10 | 1000000 | Ω | µΩ |  |  |  | 577000000000 | 0 | 577000000000 | 0
34420A | DCVmeas | 0 | 0.001 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.027845 | 71.95
34420A | DCVmeas | 0 | 0.001 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.062207 | 53.073
34420A | DCVmeas | 0 | 0.001 | 1e-06 | V | µV |  |  |  | 0.62 | 4.5 | 0.5802 | 8.2952
34420A | DCVmeas | 0 | 0.001 | 1e-05 | V | µV |  |  |  | 5.8 | 0.89276 | 5.8 | 0.83626
34420A | DCVmeas | 0 | 0.001 | 0.0001 | V | µV |  |  |  | 58 | 0.10624 | 58 | 0.083633
34420A | DCVmeas | 0 | 0.001 | 0.001 | V | µV |  |  |  | 580 | 0.017406 | 580 | 0.0083633
34420A | DCVmeas | 0.001 | 0.01 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.060672 | 48.506
34420A | DCVmeas | 0.001 | 0.01 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.076236 | 47.254
34420A | DCVmeas | 0.001 | 0.01 | 1e-06 | V | µV |  |  |  | 0.62 | 19.155 | 0.56725 | 22.912
34420A | DCVmeas | 0.001 | 0.01 | 1e-05 | V | µV |  |  |  | 5.7983 | 2.757 | 5.7983 | 2.7323
34420A | DCVmeas | 0.001 | 0.01 | 0.0001 | V | µV |  |  |  | 58 | 0.2838 | 58 | 0.27385
34420A | DCVmeas | 0.001 | 0.01 | 0.001 | V | µV |  |  |  | 580 | 0.031364 | 580 | 0.027386
34420A | DCVmeas | 0.01 | 0.1 | 1e-07 | V | µV |  |  |  | 0.48 | 37 | 0.47363 | 36.624
34420A | DCVmeas | 0.01 | 0.1 | 1e-06 | V | µV |  |  |  | 0.6696 | 35.075 | 0.66823 | 35.079
34420A | DCVmeas | 0.01 | 0.1 | 1e-05 | V | µV |  |  |  | 5.7207 | 14.051 | 5.7199 | 14.036
34420A | DCVmeas | 0.01 | 0.1 | 0.0001 | V | µV |  |  |  | 57.991 | 1.5774 | 57.99 | 1.5689
34420A | DCVmeas | 0.01 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.16049 | 580 | 0.1571
34420A | DCVmeas | 0.01 | 0.1 | 0.01 | V | µV |  |  |  | 5800 | 0.017067 | 5800 | 0.01571
34420A | DCVmeas | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 4.7 | 41 | 4.6518 | 40.782
34420A | DCVmeas | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 6.5523 | 39.255 | 6.5389 | 39.26
34420A | DCVmeas | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 56.989 | 16.705 | 56.981 | 16.693
34420A | DCVmeas | 0.1 | 1 | 0.001 | V | µV |  |  |  | 579.88 | 1.9084 | 579.88 | 1.9012
34420A | DCVmeas | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5800 | 0.19329 | 5800 | 0.19042
34420A | DCVmeas | 0.1 | 1 | 0.1 | V | µV |  |  |  | 58000 | 0.020188 | 58000 | 0.019042
34420A | DCVmeas | 1 | 10 | 1e-05 | V | µV |  |  |  | 47 | 35 | 46.376 | 34.813
34420A | DCVmeas | 1 | 10 | 0.0001 | V | µV |  |  |  | 66.458 | 33.232 | 66.382 | 33.232
34420A | DCVmeas | 1 | 10 | 0.001 | V | µV |  |  |  | 572.81 | 12.879 | 572.76 | 12.867
34420A | DCVmeas | 1 | 10 | 0.01 | V | µV |  |  |  | 5799.2 | 1.4325 | 5799.1 | 1.4259
34420A | DCVmeas | 1 | 10 | 0.1 | V | µV |  |  |  | 58000 | 0.1454 | 58000 | 0.14276
34420A | DCVmeas | 1 | 10 | 1 | V | µV |  |  |  | 580000 | 0.015331 | 580000 | 0.014276
34420A | DCVmeas | 10 | 100 | 0.0001 | V | mV |  |  |  | 0.58 | 0.041 | 0.57895 | 0.040741
34420A | DCVmeas | 10 | 100 | 0.001 | V | mV |  |  |  | 0.7495 | 0.039399 | 0.74879 | 0.039399
34420A | DCVmeas | 10 | 100 | 0.01 | V | mV |  |  |  | 5.7109 | 0.017264 | 5.7105 | 0.017251
34420A | DCVmeas | 10 | 100 | 0.1 | V | mV |  |  |  | 57.989 | 0.0019846 | 57.989 | 0.0019773
34420A | DCVmeas | 10 | 100 | 1 | V | mV |  |  |  | 580 | 0.000201 | 580 | 0.00019806
34420A | Resistance | 0 | 1000 | 1e-06 | Ohms | µΩ |  |  |  | 3 | 0.12 | 2.3273 | 0.080813
34420A | Resistance | 0 | 1000 | 1e-05 | Ohms | µΩ |  |  |  | 3 | 0.12 | 3.9637 | 0.079377
34420A | Resistance | 0 | 1000 | 0.0001 | Ohms | µΩ |  |  |  | 63.6928 | 0.062659 | 54.2074 | 0.047163
34420A | Resistance | 0 | 1000 | 0.001 | Ohms | µΩ |  |  |  | 600.322 | 0.096481 | 579.4447 | 0.0064836
34420A | Resistance | 0 | 1000 | 0.01 | Ohms | mΩ |  |  |  | 5.8513 | 0.00024805 | 5.7999 | 6.5169e-07
34420A | Resistance | 1000 | 10000 | 1e-05 | Ohms | µΩ |  |  |  | 24.0351 | 0.069214 | 23.2955 | 0.069264
34420A | Resistance | 1000 | 10000 | 0.0001 | Ohms | µΩ |  |  |  | 66.5586 | 0.066056 | 41.3894 | 0.067687
34420A | Resistance | 1000 | 10000 | 0.001 | Ohms | µΩ |  |  |  | 643.2147 | 0.054751 | 550.1922 | 0.037118
34420A | Resistance | 1000 | 10000 | 0.01 | Ohms | mΩ |  |  |  | 6.0044 | 9.5076e-05 | 5.7959 | 4.809e-06
34420A | Resistance | 1000 | 10000 | 0.1 | Ohms | mΩ |  |  |  | 58.5135 | 0.00024789 | 57.9996 | 4.8274e-07
34420A | Resistance | 10000 | 100000 | 0.0001 | Ohms | mΩ |  |  |  | 0.24035 | 6.9214e-05 | 0.23295 | 6.9264e-05
34420A | Resistance | 10000 | 100000 | 0.001 | Ohms | mΩ |  |  |  | 0.66559 | 6.6056e-05 | 0.41389 | 6.7687e-05
34420A | Resistance | 10000 | 100000 | 0.01 | Ohms | mΩ |  |  |  | 6.4321 | 5.4751e-05 | 5.5019 | 3.7118e-05
34420A | Resistance | 10000 | 100000 | 0.1 | Ohms | mΩ |  |  |  | 60.0444 | 9.5076e-05 | 57.9593 | 4.809e-06
34420A | Resistance | 10000 | 100000 | 1 | Ohms | mΩ |  |  |  | 585.1347 | 0.00024789 | 579.9959 | 4.83e-07
34420A | Resistance | 100000 | 1000 | 1 | Ohm | mΩ |  |  |  | 2.4035 | 6.9214e-05 | 2.3295 | 6.9264e-05
34420A | Resistance | 100000 | 1000 | 10 | Ohm | mΩ |  |  |  | 6.6559 | 6.6056e-05 | 4.1389 | 6.7687e-05
34420A | Resistance | 100000 | 1000 | 100 | Ohm | mΩ |  |  |  | 64.3215 | 5.4751e-05 | 55.0192 | 3.7118e-05
34420A | Resistance | 100000 | 1000 | 1000 | Ohm | mΩ |  |  |  | 600.444 | 9.5076e-05 | 579.5927 | 4.809e-06
34420A | Resistance | 100000 | 1000 | 10000 | Ohm | Ω |  |  |  | 5.8513 | 2.4789e-07 | 5.8 | 4.83e-10
34420A | Resistance | 1000 | 10000 | 10 | Ohm | mΩ |  |  |  | 24.0351 | 6.9214e-05 | 23.2955 | 6.9264e-05
34420A | Resistance | 1000 | 10000 | 100 | Ohm | mΩ |  |  |  | 66.5586 | 6.6056e-05 | 41.3894 | 6.7688e-05
34420A | Resistance | 1000 | 10000 | 1000 | Ohm | mΩ |  |  |  | 643.2147 | 5.4751e-05 | 550.192 | 3.7118e-05
34420A | Resistance | 1000 | 10000 | 10000 | Ohm | Ω |  |  |  | 6.0044 | 9.5076e-08 | 5.7959 | 4.809e-09
34420A | Resistance | 1000 | 10000 | 100000 | Ohm | Ω |  |  |  | 58.5135 | 2.4789e-07 | 57.9996 | 4.83e-10
34420A | Resistance | 10000 | 100000 | 100 | Ohm | Ω |  |  |  | 0.46937 | 6.92e-08 | 0.46349 | 6.93e-08
34420A | Resistance | 10000 | 100000 | 1000 | Ohm | Ω |  |  |  | 0.82928 | 6.67e-08 | 0.61213 | 6.81e-08
34420A | Resistance | 10000 | 100000 | 10000 | Ohm | Ω |  |  |  | 6.4547 | 5.59e-08 | 5.5271 | 3.87e-08
34420A | Resistance | 10000 | 100000 | 100000 | Ohm | Ω |  |  |  | 60.0461 | 9.53e-08 | 57.9607 | 5.09e-09
34420A | Resistance | 10000 | 100000 | 1000000 | Ohm | Ω |  |  |  | 585.1349 | 2.48e-07 | 579.996 | 5.1e-10
34420A | Resistance | 100000 | 1000 | 1000000 | Ohm | Ω |  |  |  | 5 | 8.05e-08 | 4.6334 | 8.08e-08
34420A | Resistance | 100000 | 1000 | 10000000 | Ohm | Ω |  |  |  | 8.0349 | 7.84e-08 | 5.9988 | 7.96e-08
34420A | Resistance | 100000 | 1000 | 100000000 | Ohm | Ω |  |  |  | 63.945 | 6.39e-08 | 54.4969 | 4.88e-08
34420A | Resistance | 100000 | 1000 | 1000000000 | Ohm | Ω |  |  |  | 600.3396 | 9.67e-08 | 579.4589 | 6.8e-09
34420A | Resistance | 100000 | 1000 | 10000000000 | Ohm | kΩ |  |  |  | 5.8513 | 2.481e-10 | 5.7999 | 6.9e-13
Transmille | ACA | 1e-05 | 0.000104 | 1e-09 | A | µA | 0.01 | 2 | KHz | 0.47 | 1200 | 0.46304 | 1230
Transmille | ACA | 1e-05 | 0.000104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.47 | 1230 | 0.46308 | 1230
Transmille | ACA | 1e-05 | 0.000104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.47 | 1230 | 0.4666 | 1230
Transmille | ACA | 1e-05 | 0.000104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 0.73981 | 857 | 0.73979 | 856
Transmille | ACA | 1e-05 | 0.000104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8088 | 394 | 5.8088 | 394
Transmille | ACA | 0.0001 | 0.00104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.97 | 1200 | 0.9653 | 1190
Transmille | ACA | 0.0001 | 0.00104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.97 | 1190 | 0.96691 | 1190
Transmille | ACA | 0.0001 | 0.00104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 1.1172 | 1120 | 1.1168 | 1120
Transmille | ACA | 0.0001 | 0.00104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8401 | 354 | 5.8398 | 352
Transmille | ACA | 0.0001 | 0.00104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 57.912 | 316 | 57.912 | 315
Transmille | ACA | 0.001 | 0.0104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 9.7 | 1200 | 9.653 | 1190
Transmille | ACA | 0.001 | 0.0104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 9.7 | 1190 | 9.6691 | 1190
Transmille | ACA | 0.001 | 0.0104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 11.172 | 1120 | 11.168 | 1120
Transmille | ACA | 0.001 | 0.0104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 58.401 | 354 | 58.398 | 352
Transmille | ACA | 0.001 | 0.0104 | 0.001 | A | µA | 0.01 | 2 | KHz | 579.12 | 316 | 579.12 | 315
Transmille | ACA | 0.01 | 0.104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 97 | 1200 | 96.53 | 1190
Transmille | ACA | 0.01 | 0.104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 97 | 1190 | 96.691 | 1190
Transmille | ACA | 0.01 | 0.104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 111.72 | 1120 | 111.68 | 1120
Transmille | ACA | 0.01 | 0.104 | 0.001 | A | µA | 0.01 | 2 | KHz | 583.96 | 354 | 583.98 | 352
Transmille | ACA | 0.01 | 0.104 | 0.01 | A | µA | 0.01 | 2 | KHz | 5791.2 | 316 | 5791.2 | 315
Transmille | ACA | 0.1 | 1.04 | 1e-05 | A | mA | 0.01 | 2 | KHz | 1 | 1.5 | 0.99589 | 1.49
Transmille | ACA | 0.1 | 1.04 | 0.0001 | A | mA | 0.01 | 2 | KHz | 1 | 1.49 | 0.99738 | 1.49
Transmille | ACA | 0.1 | 1.04 | 0.001 | A | mA | 0.01 | 2 | KHz | 1.1829 | 1.41 | 1.1409 | 1.41
Transmille | ACA | 0.1 | 1.04 | 0.01 | A | mA | 0.01 | 2 | KHz | 5.9116 | 0.478 | 5.8383 | 0.476
Transmille | ACA | 0.1 | 1.04 | 0.1 | A | mA | 0.01 | 2 | KHz | 57.947 | 0.33 | 57.91 | 0.329
Transmille | ACA | 1 | 10.4 | 0.0001 | A | mA | 0.01 | 2 | KHz | 18 | 1.4 | 17.128 | 1.4
Transmille | ACA | 1 | 10.4 | 0.001 | A | mA | 0.01 | 2 | KHz | 18 | 1.4 | 17.138 | 1.4
Transmille | ACA | 1 | 10.4 | 0.01 | A | mA | 0.01 | 2 | KHz | 18.047 | 1.36 | 18.045 | 1.36
Transmille | ACA | 1 | 10.4 | 0.1 | A | mA | 0.01 | 2 | KHz | 60.057 | 0.582 | 60.055 | 0.581
Transmille | ACA | 1 | 10.4 | 1 | A | mA | 0.01 | 2 | KHz | 579.27 | 0.343 | 579.27 | 0.342
Transmille | ACV | 0 | 0.1 | 1e-06 | V | µV | 0.01 | 2 | KHz | 35 | 940 | 34.74 | 937
Transmille | ACV | 0 | 0.1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 35.162 | 934 | 35.157 | 934
Transmille | ACV | 0 | 0.1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 67.116 | 737 | 67.105 | 737
Transmille | ACV | 0 | 0.1 | 0.001 | V | µV | 0.01 | 2 | KHz | 578.33 | 132 | 578.32 | 131
Transmille | ACV | 0 | 0.1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5773.7 | 13.4 | 5773.6 | 13.3
Transmille | ACV | 0 | 0.1 | 1e-06 | V | µV | 2 | 20 | KHz | 36 | 930 | 35.471 | 926
Transmille | ACV | 0 | 0.1 | 1e-05 | V | µV | 2 | 20 | KHz | 36 | 923 | 35.865 | 923
Transmille | ACV | 0 | 0.1 | 0.0001 | V | µV | 2 | 20 | KHz | 67.126 | 727 | 67.116 | 727
Transmille | ACV | 0 | 0.1 | 0.001 | V | µV | 2 | 20 | KHz | 578.33 | 129 | 578.32 | 129
Transmille | ACV | 0 | 0.1 | 0.01 | V | µV | 2 | 20 | KHz | 5773.7 | 13.2 | 5773.6 | 13.1
Transmille | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 350 | 940 | 347.15 | 937
Transmille | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 351.19 | 935 | 351.16 | 934
Transmille | ACV | 0.1 | 1 | 0.001 | V | µV | 0.01 | 2 | KHz | 650.74 | 758 | 650.66 | 757
Transmille | ACV | 0.1 | 1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5776.6 | 138 | 5776.5 | 138
Transmille | ACV | 0.1 | 1 | 0.1 | V | µV | 0.01 | 2 | KHz | 57736 | 14.1 | 57736 | 14
Transmille | ACV | 0.1 | 1 | 1e-05 | V | µV | 2 | 20 | KHz | 850 | 1700 | 808.55 | 1740
Transmille | ACV | 0.1 | 1 | 0.0001 | V | µV | 2 | 20 | KHz | 850 | 1740 | 810.33 | 1740
Transmille | ACV | 0.1 | 1 | 0.001 | V | µV | 2 | 20 | KHz | 976.25 | 1640 | 975.89 | 1640
Transmille | ACV | 0.1 | 1 | 0.01 | V | µV | 2 | 20 | KHz | 5806.3 | 505 | 5806.1 | 504
Transmille | ACV | 0.1 | 1 | 0.1 | V | µV | 2 | 20 | KHz | 57739 | 53.4 | 57739 | 53.1
Transmille | ACV | 1 | 10 | 0.0001 | V | mV | 0.01 | 2 | KHz | 3.5 | 0.94 | 3.4755 | 0.937
Transmille | ACV | 1 | 10 | 0.001 | V | mV | 0.01 | 2 | KHz | 3.5127 | 0.935 | 3.5116 | 0.934
Transmille | ACV | 1 | 10 | 0.01 | V | mV | 0.01 | 2 | KHz | 6.5091 | 0.758 | 6.5066 | 0.757
Transmille | ACV | 1 | 10 | 0.1 | V | mV | 0.01 | 2 | KHz | 57.766 | 0.139 | 57.765 | 0.138
Transmille | ACV | 1 | 10 | 1 | V | mV | 0.01 | 2 | KHz | 577.36 | 0.0144 | 577.36 | 0.014
Transmille | ACV | 1 | 10 | 0.0001 | V | mV | 2 | 20 | KHz | 8.2 | 1.8 | 8.1347 | 1.75
Transmille | ACV | 1 | 10 | 0.001 | V | mV | 2 | 20 | KHz | 8.2 | 1.75 | 8.1525 | 1.75
Transmille | ACV | 1 | 10 | 0.01 | V | mV | 2 | 20 | KHz | 9.8126 | 1.65 | 9.799 | 1.65
Transmille | ACV | 1 | 10 | 0.1 | V | mV | 2 | 20 | KHz | 58.073 | 0.513 | 58.064 | 0.512
Transmille | ACV | 1 | 10 | 1 | V | mV | 2 | 20 | KHz | 577.39 | 0.0545 | 577.39 | 0.054
Transmille | ACV | 10 | 100 | 0.001 | V | mV | 0.04 | 1 | KHz | 35 | 0.95 | 34.943 | 0.948
Transmille | ACV | 10 | 100 | 0.01 | V | mV | 0.04 | 1 | KHz | 35.349 | 0.946 | 35.282 | 0.945
Transmille | ACV | 10 | 100 | 0.1 | V | mV | 0.04 | 1 | KHz | 65.076 | 0.769 | 64.997 | 0.768
Transmille | ACV | 10 | 100 | 1 | V | mV | 0.04 | 1 | KHz | 577.66 | 0.143 | 577.63 | 0.141
Transmille | ACV | 10 | 100 | 10 | V | mV | 0.04 | 1 | KHz | 5773.6 | 0.0151 | 5773.6 | 0.0143
Transmille | ACV | 100 | 1020 | 0.01 | V | mV | 0.04 | 1 | KHz | 360 | 1.1 | 337.15 | 1.12
Transmille | ACV | 100 | 1020 | 0.1 | V | mV | 0.04 | 1 | KHz | 360 | 1.11 | 341.11 | 1.11
Transmille | ACV | 100 | 1020 | 1 | V | mV | 0.04 | 1 | KHz | 639.27 | 0.929 | 638.46 | 0.928
Transmille | ACV | 100 | 1020 | 10 | V | mV | 0.04 | 1 | KHz | 5773 | 0.185 | 5772.7 | 0.183
Transmille | ACV | 100 | 1020 | 100 | V | mV | 0.04 | 1 | KHz | 57736 | 0.0194 | 57735 | 0.0186
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-12 | F | nF |  |  |  | 0.036 | 6600000 | 0.035476 | 6560000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-11 | F | nF |  |  |  | 0.036 | 6540000 | 0.035893 | 6530000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-10 | F | nF |  |  |  | 0.066425 | 5020000 | 0.066402 | 5010000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-09 | F | nF |  |  |  | 0.57863 | 1100000 | 0.57863 | 1100000
Transmille | Capacitance | 1e-09 | 1e-08 | 1e-08 | F | nF |  |  |  | 5.7903 | 3020000 | 5.7903 | 3020000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-11 | F | nF |  |  |  | 0.5 | 5800000 | 0.48989 | 5800000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-10 | F | nF |  |  |  | 0.5 | 5790000 | 0.49306 | 5790000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-09 | F | nF |  |  |  | 0.74933 | 4690000 | 0.74914 | 4680000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-08 | F | nF |  |  |  | 5.797 | 1100000 | 5.797 | 1090000
Transmille | Capacitance | 1e-08 | 1e-07 | 1e-07 | F | nF |  |  |  | 57.904 | 3020000 | 57.904 | 3020000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-10 | F | nF |  |  |  | 6.2 | 5200000 | 6.1439 | 5210000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-09 | F | nF |  |  |  | 6.2 | 5200000 | 6.1698 | 5200000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-08 | F | nF |  |  |  | 8.3825 | 4380000 | 8.3808 | 4370000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-07 | F | nF |  |  |  | 58.094 | 1100000 | 58.094 | 1090000
Transmille | Capacitance | 1e-07 | 1e-06 | 1e-06 | F | nF |  |  |  | 579.05 | 3020000 | 579.05 | 3020000
Transmille | DCA | 0 | 0.0001 | 1e-09 | A | nA |  |  |  | 35 | 350000 | 34.933 | 349000
Transmille | DCA | 0 | 0.0001 | 1e-08 | A | nA |  |  |  | 35.368 | 347000 | 35.366 | 347000
Transmille | DCA | 0 | 0.0001 | 1e-07 | A | nA |  |  |  | 67.34 | 231000 | 67.336 | 231000
Transmille | DCA | 0 | 0.0001 | 1e-06 | A | nA |  |  |  | 578.4 | 31400 | 578.39 | 31400
Transmille | DCA | 0 | 0.0001 | 1e-05 | A | nA |  |  |  | 5773.7 | 3160 | 5773.7 | 3150
Transmille | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 120 | 350000 | 115.86 | 347000
Transmille | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 127.29 | 339000 | 127.24 | 339000
Transmille | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 580.76 | 159000 | 580.72 | 159000
Transmille | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.7 | 18400 | 5773.7 | 18400
Transmille | DCA | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 57736 | 1850 | 57736 | 1840
Transmille | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 1.2 | 350 | 1.1597 | 347
Transmille | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 1.2739 | 339 | 1.2734 | 339
Transmille | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.8079 | 159 | 5.8074 | 159
Transmille | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.737 | 18.4 | 57.737 | 18.4
Transmille | DCA | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 1.86 | 577.36 | 1.84
Transmille | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 12 | 350 | 11.589 | 348
Transmille | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 12.73 | 341 | 12.725 | 341
Transmille | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 58.071 | 160 | 58.066 | 160
Transmille | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.37 | 18.6 | 577.37 | 18.5
Transmille | DCA | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 1.87 | 5773.6 | 1.86
Transmille | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 180 | 370 | 172.98 | 365
Transmille | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 181.28 | 360 | 181.24 | 360
Transmille | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 594.75 | 195 | 594.71 | 195
Transmille | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5775 | 23.6 | 5775 | 23.6
Transmille | DCA | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 2.38 | 57736 | 2.36
Transmille | DCA | 1 | 10.2 | 0.0001 | A | µA |  |  |  | 2.4 | 0.59 | 2.3095 | 0.58
Transmille | DCA | 1 | 10.2 | 0.001 | A | µA |  |  |  | 2.4 | 0.631 | 2.3701 | 0.576
Transmille | DCA | 1 | 10.2 | 0.01 | A | µA |  |  |  | 5.8935 | 0.563 | 6.0655 | 0.39
Transmille | DCA | 1 | 10.2 | 0.1 | A | µA |  |  |  | 57.625 | 0.183 | 57.752 | 0.0555
Transmille | DCA | 1 | 10.2 | 1 | A | µA |  |  |  | 577.31 | 0.0574 | 577.36 | 0.00558
Transmille | DCA 10 turn | 10 | 102 | 0.0001 | A | mA |  |  |  | 16 | 2 | 15.052 | 2.01
Transmille | DCA 10 turn | 10 | 102 | 0.001 | A | mA |  |  |  | 16 | 2 | 15.566 | 2
Transmille | DCA 10 turn | 10 | 102 | 0.01 | A | mA |  |  |  | 50.243 | 1.74 | 50.216 | 1.74
Transmille | DCA 10 turn | 10 | 102 | 0.1 | A | mA |  |  |  | 574.15 | 0.429 | 574.14 | 0.428
Transmille | DCA 10 turn | 10 | 102 | 1 | A | mA |  |  |  | 5773.2 | 0.0445 | 5773.2 | 0.0444
Transmille | DCA 10 turn | 10 | 102 | 10 | A | mA |  |  |  | 57735 | 0.00449 | 57735 | 0.00444
Transmille | DCA 20 turn | 20 | 204 | 0.001 | A | mA |  |  |  | 32 | 2 | 30.087 | 2.01
Transmille | DCA 20 turn | 20 | 204 | 0.01 | A | mA |  |  |  | 32 | 2.01 | 30.344 | 2.01
Transmille | DCA 20 turn | 20 | 204 | 0.1 | A | mA |  |  |  | 52.695 | 1.92 | 52.601 | 1.92
Transmille | DCA 20 turn | 20 | 204 | 1 | A | mA |  |  |  | 565.98 | 0.784 | 565.94 | 0.784
Transmille | DCA 20 turn | 20 | 204 | 10 | A | mA |  |  |  | 5772.2 | 0.0889 | 5772.2 | 0.0886
Transmille | DCA 50 turn | 5 | 51 | 0.001 | A | mA |  |  |  | 4.7 | 2 | 4.5993 | 1.98
Transmille | DCA 50 turn | 5 | 51 | 0.01 | A | mA |  |  |  | 5.9279 | 1.96 | 5.8001 | 1.96
Transmille | DCA 50 turn | 5 | 51 | 0.1 | A | mA |  |  |  | 53.039 | 1.32 | 52.916 | 1.32
Transmille | DCA 50 turn | 5 | 51 | 1 | A | mA |  |  |  | 576.56 | 0.206 | 576.52 | 0.204
Transmille | DCA 50 turn | 5 | 51 | 10 | A | mA |  |  |  | 5773.5 | 0.0215 | 5773.5 | 0.0206
Transmille | DCA 50 turn | 50 | 510 | 0.001 | A | mA |  |  |  | 80 | 2 | 75.208 | 2.01
Transmille | DCA 50 turn | 50 | 510 | 0.01 | A | mA |  |  |  | 80 | 2.01 | 75.318 | 2.01
Transmille | DCA 50 turn | 50 | 510 | 0.1 | A | mA |  |  |  | 85.569 | 1.99 | 85.292 | 1.99
Transmille | DCA 50 turn | 50 | 510 | 1 | A | mA |  |  |  | 534.37 | 1.39 | 534.07 | 1.39
Transmille | DCA 50 turn | 50 | 510 | 10 | A | mA |  |  |  | 5765.3 | 0.22 | 5765.2 | 0.22
Transmille | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 12 | 92 | 11.624 | 92.1
Transmille | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 12.963 | 86.4 | 12.956 | 86.5
Transmille | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.889 | 24.9 | 58.883 | 24.9
Transmille | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 577.47 | 2.59 | 577.47 | 2.59
Transmille | DCV | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5773.6 | 0.259 | 5773.6 | 0.259
Transmille | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 36 | 92 | 35.885 | 91.6
Transmille | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 65.685 | 74.3 | 65.574 | 74.3
Transmille | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 577.78 | 13.6 | 577.73 | 13.5
Transmille | DCV | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5773.6 | 1.39 | 5773.6 | 1.37
Transmille | DCV | 0.1 | 1 | 0.1 | V | µV |  |  |  | 57736 | 0.144 | 57736 | 0.137
Transmille | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 360 | 92 | 354.54 | 91.8
Transmille | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 653.72 | 74.4 | 653.17 | 74.3
Transmille | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5777.2 | 13.6 | 5777 | 13.5
Transmille | DCV | 1 | 10 | 0.1 | V | µV |  |  |  | 57736 | 1.4 | 57736 | 13.5
Transmille | DCV | 1 | 10 | 1 | V | µV |  |  |  | 577360 | 0.148 | 577360 | 13.5
Transmille | DCV | 10 | 100 | 0.001 | V | mV |  |  |  | 6.8 | 0.082 | 6.7338 | 0.0825
Transmille | DCV | 10 | 100 | 0.01 | V | mV |  |  |  | 8.774 | 0.0728 | 8.766 | 0.0728
Transmille | DCV | 10 | 100 | 0.1 | V | mV |  |  |  | 58.072 | 0.0158 | 58.068 | 0.0158
Transmille | DCV | 10 | 100 | 1 | V | mV |  |  |  | 577.39 | 0.00164 | 577.39 | 0.00161
Transmille | DCV | 10 | 100 | 10 | V | mV |  |  |  | 5773.6 | 0.000174 | 5773.6 | 0.000161
Transmille | DCV | 100 | 1020 | 0.01 | V | mV |  |  |  | 36 | 0.092 | 35.843 | 0.0918
Transmille | DCV | 100 | 1020 | 0.1 | V | mV |  |  |  | 65.632 | 0.0746 | 65.525 | 0.0746
Transmille | DCV | 100 | 1020 | 1 | V | mV |  |  |  | 577.76 | 0.0138 | 577.71 | 0.0137
Transmille | DCV | 100 | 1020 | 10 | V | mV |  |  |  | 5773.6 | 0.00142 | 5773.6 | 0.00139
Transmille | DCV | 100 | 1020 | 100 | V | mV |  |  |  | 57736 | 0.000151 | 57736 | 0.000139
Transmille | Frequency | 1 | 100000 | 1 | Hz | Hz |  |  |  | 0.87 | 1.6e-05 | 0.81638 | 1.64e-05
Transmille | Frequency | 1 | 100000 | 10 | Hz | Hz |  |  |  | 5.8379 | 4.82e-06 | 5.8243 | 4.68e-06
Transmille | Frequency | 1 | 100000 | 100 | Hz | Hz |  |  |  | 58.005 | 3.17e-06 | 58 | 3.11e-06
Transmille | Frequency | 1 | 100000 | 1000 | Hz | Hz |  |  |  | 579.98 | 2.65e-05 | 579.97 | 2.65e-05
Transmille | Frequency | 1 | 100000 | 10000 | Hz | Hz |  |  |  | 5799.7 | 0.000265 | 5799.7 | 0.000265
Transmille | Insulation DCV | 1 | 100 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.4054 | 0.0105
Transmille | Insulation DCV | 1 | 100 | 10 | V | V |  |  |  | 5.9149 | 0.00383 | 5.9069 | 0.00367
Transmille | Insulation DCV | 1 | 100 | 100 | V | V |  |  |  | 57.918 | 0.00337 | 57.915 | 0.0033
Transmille | Insulation DCV | 100 | 250 | 1 | V | V |  |  |  | 1.4 | 0.011 | 1.3333 | 0.0112
Transmille | Insulation DCV | 100 | 250 | 10 | V | V |  |  |  | 5.6899 | 0.00577 | 5.6798 | 0.00565
Transmille | Insulation DCV | 100 | 250 | 100 | V | V |  |  |  | 57.719 | 0.00246 | 57.716 | 0.00239
Transmille | Insulation DCV | 250 | 500 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.2696 | 0.0114
Transmille | Insulation DCV | 250 | 500 | 10 | V | V |  |  |  | 5.1032 | 0.00803 | 5.0881 | 0.00793
Transmille | Insulation DCV | 250 | 500 | 100 | V | V |  |  |  | 57.497 | 0.00222 | 57.502 | 0.00214
Transmille | Insulation DCV | 500 | 1000 | 1 | V | V |  |  |  | 1.8 | 0.011 | 1.2242 | 0.0115
Transmille | Insulation DCV | 500 | 1000 | 10 | V | V |  |  |  | 4.1693 | 0.00987 | 4.0987 | 0.00988
Transmille | Insulation DCV | 500 | 1000 | 100 | V | V |  |  |  | 57.066 | 0.00249 | 57.041 | 0.00245
Transmille | Insulation Res | 250000 | 100000000 | 10000 | Ω | kΩ |  |  |  | 18 | 9.2e-06 | 13.773 | 9.24e-06
Transmille | Insulation Res | 250000 | 100000000 | 100000 | Ω | kΩ |  |  |  | 57.523 | 8.82e-06 | 57.453 | 8.82e-06
Transmille | Insulation Res | 250000 | 100000000 | 1000000 | Ω | kΩ |  |  |  | 576.26 | 5.27e-06 | 576.24 | 5.27e-06
Transmille | Insulation Res | 250000 | 100000000 | 10000000 | Ω | kΩ |  |  |  | 5775 | 1.02e-06 | 5775 | 1.02e-06
Transmille | Insulation Res | 100000000 | 250000000 | 10000 | Ω | kΩ |  |  |  | 14 | 1.1e-05 | 13.511 | 1.08e-05
Transmille | Insulation Res | 100000000 | 250000000 | 100000 | Ω | kΩ |  |  |  | 57.138 | 1.06e-05 | 57.125 | 1.06e-05
Transmille | Insulation Res | 100000000 | 250000000 | 1000000 | Ω | kΩ |  |  |  | 575.4 | 8.79e-06 | 575.35 | 8.79e-06
Transmille | Insulation Res | 100000000 | 250000000 | 10000000 | Ω | kΩ |  |  |  | 5773 | 2.52e-06 | 5772.9 | 2.52e-06
Transmille | Insulation Res | 100000000 | 250000000 | 100000000 | Ω | kΩ |  |  |  | 57896 | 1.31e-06 | 57896 | 1.31e-06
Transmille | Insulation Res | 250000000 | 500000000 | 10000 | Ω | kΩ |  |  |  | 18 | 1.1e-05 | 12.764 | 1.1e-05
Transmille | Insulation Res | 250000000 | 500000000 | 100000 | Ω | kΩ |  |  |  | 55.079 | 1.09e-05 | 54.77 | 1.09e-05
Transmille | Insulation Res | 250000000 | 500000000 | 1000000 | Ω | kΩ |  |  |  | 572.67 | 9.95e-06 | 572.64 | 9.95e-06
Transmille | Insulation Res | 250000000 | 500000000 | 10000000 | Ω | kΩ |  |  |  | 5771.4 | 4.47e-06 | 5771.3 | 4.47e-06
Transmille | Insulation Res | 500000000 | 1000000000 | 10000 | Ω | kΩ |  |  |  | 25 | 1.1e-05 | -76.51 | 1.12e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 100000 | Ω | kΩ |  |  |  | 25 | 1.12e-05 | -75.391 | 1.12e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 1000000 | Ω | kΩ |  |  |  | 25 | 1.1e-05 | 24.396 | 1.1e-05
Transmille | Insulation Res | 500000000 | 1000000000 | 10000000 | Ω | kΩ |  |  |  | 5114.5 | 7.34e-06 | 5114.5 | 7.34e-06
Transmille | Insulation Res | 500000000 | 1000000000 | 100000000 | Ω | kΩ |  |  |  | 57638 | 1.44e-06 | 57638 | 1.44e-06
Transmille | TcindicatorJ | -346 | -148 | 0.1 | °F | °F |  |  |  | 2.4168 | 0 | 2.4145 | 0
Transmille | TcindicatorJ | -148 | 302 | 0.1 | °F | °F |  |  |  | 1.9165 | 0 | 1.9145 | 0
Transmille | TcindicatorJ | 302 | 1400 | 0.1 | °F | °F |  |  |  | 2.0393 | 0 | 2.0387 | 0
Transmille | TcindicatorJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 2.2469 | 0 | 2.2464 | 0
Transmille | TcindicatorJ | -346 | -148 | 1 | °F | °F |  |  |  | 2.4908 | 0 | 2.4819 | 0
Transmille | TcindicatorJ | -148 | 302 | 1 | °F | °F |  |  |  | 2.0067 | 0 | 1.9988 | 0
Transmille | TcindicatorJ | 302 | 1400 | 1 | °F | °F |  |  |  | 2.1202 | 0 | 2.1181 | 0
Transmille | TcindicatorJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 2.3207 | 0 | 2.3187 | 0
Transmille | TcindicatorJ | -210 | -100 | 0.1 | °C | °C |  |  |  | 1.3461 | 0 | 1.3448 | 0
Transmille | TcindicatorJ | -100 | 150 | 0.1 | °C | °C |  |  |  | 1.0665 | 0 | 1.0653 | 0
Transmille | TcindicatorJ | 150 | 760 | 0.1 | °C | °C |  |  |  | 1.1342 | 0 | 1.1339 | 0
Transmille | TcindicatorJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 1.2495 | 0 | 1.2492 | 0
Transmille | TcindicatorJ | -210 | -100 | 1 | °C | °C |  |  |  | 1.467 | 0 | 1.4624 | 0
Transmille | TcindicatorJ | -100 | 150 | 1 | °C | °C |  |  |  | 1.2144 | 0 | 1.2103 | 0
Transmille | TcindicatorJ | 150 | 760 | 1 | °C | °C |  |  |  | 1.2722 | 0 | 1.2711 | 0
Transmille | TcindicatorJ | 760 | 1200 | 1 | °C | °C |  |  |  | 1.376 | 0 | 1.375 | 0
Transmille | TcindicatorK | -328 | -148 | 0.1 | °F | °F |  |  |  | 2.5841 | 0 | 2.5813 | 0
Transmille | TcindicatorK | -148 | 248 | 0.1 | °F | °F |  |  |  | 2.0844 | 0 | 2.0809 | 0
Transmille | TcindicatorK | 248 | 2500 | 0.1 | °F | °F |  |  |  | 2.5391 | 0 | 2.5376 | 0
Transmille | TcindicatorK | -328 | -148 | 1 | °F | °F |  |  |  | 2.6555 | 0 | 2.6444 | 0
Transmille | TcindicatorK | -148 | 248 | 1 | °F | °F |  |  |  | 2.1723 | 0 | 2.1587 | 0
Transmille | TcindicatorK | 248 | 2500 | 1 | °F | °F |  |  |  | 2.6075 | 0 | 2.6018 | 0
Transmille | TcindicatorK | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.4398 | 0 | 1.4382 | 0
Transmille | TcindicatorK | -100 | 120 | 0.1 | °C | °C |  |  |  | 1.16 | 0 | 1.1581 | 0
Transmille | TcindicatorK | 120 | 1370 | 0.1 | °C | °C |  |  |  | 1.4121 | 0 | 1.4113 | 0
Transmille | TcindicatorK | -200 | -100 | 1 | °C | °C |  |  |  | 1.5545 | 0 | 1.5487 | 0
Transmille | TcindicatorK | -100 | 120 | 1 | °C | °C |  |  |  | 1.2998 | 0 | 1.2927 | 0
Transmille | TcindicatorK | 120 | 1370 | 1 | °C | °C |  |  |  | 1.5267 | 0 | 1.5237 | 0
Transmille | TcindicatorN | -328 | -148 | 0.1 | °F | °F |  |  |  | 3.205 | 0 | 3.2025 | 0
Transmille | TcindicatorN | -148 | 770 | 0.1 | °F | °F |  |  |  | 2.2918 | 0 | 2.2894 | 0
Transmille | TcindicatorN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 2.4558 | 0 | 2.4545 | 0
Transmille | TcindicatorN | -328 | -148 | 1 | °F | °F |  |  |  | 3.2633 | 0 | 3.2536 | 0
Transmille | TcindicatorN | -148 | 770 | 1 | °F | °F |  |  |  | 2.3698 | 0 | 2.3604 | 0
Transmille | TcindicatorN | 770 | 2372 | 1 | °F | °F |  |  |  | 2.5259 | 0 | 2.5208 | 0
Transmille | TcindicatorN | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.782 | 0 | 1.7806 | 0
Transmille | TcindicatorN | -100 | 410 | 0.1 | °C | °C |  |  |  | 1.2762 | 0 | 1.2749 | 0
Transmille | TcindicatorN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 1.3658 | 0 | 1.3651 | 0
Transmille | TcindicatorN | -200 | -100 | 1 | °C | °C |  |  |  | 1.8762 | 0 | 1.871 | 0
Transmille | TcindicatorN | -100 | 410 | 1 | °C | °C |  |  |  | 1.4032 | 0 | 1.3983 | 0
Transmille | TcindicatorN | 410 | 1300 | 1 | °C | °C |  |  |  | 1.4837 | 0 | 1.481 | 0
Transmille | TcindicatorR | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
Transmille | TcindicatorR | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5823 | 0 | 3.5805 | 0
Transmille | TcindicatorR | 32 | 482 | 1 | °F | °F |  |  |  | 4.8335 | 0 | 4.8271 | 0
Transmille | TcindicatorR | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6336 | 0 | 3.6263 | 0
Transmille | TcindicatorR | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6784 | 0 | 2.6775 | 0
Transmille | TcindicatorR | 250 | 1760 | 0.1 | °C | °C |  |  |  | 1.9965 | 0 | 1.9955 | 0
Transmille | TcindicatorR | 0 | 250 | 1 | °C | °C |  |  |  | 2.7419 | 0 | 2.7384 | 0
Transmille | TcindicatorR | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0804 | 0 | 2.0765 | 0
Transmille | TcindicatorS | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
Transmille | TcindicatorS | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5846 | 0 | 3.5828 | 0
Transmille | TcindicatorS | 32 | 482 | 1 | °F | °F |  |  |  | 4.8334 | 0 | 4.8271 | 0
Transmille | TcindicatorS | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6358 | 0 | 3.6285 | 0
Transmille | TcindicatorS | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6783 | 0 | 2.6775 | 0
Transmille | TcindicatorS | 250 | 1760 | 0.1 | °C | °C |  |  |  | 2.0006 | 0 | 1.9996 | 0
Transmille | TcindicatorS | 0 | 250 | 1 | °C | °C |  |  |  | 2.7418 | 0 | 2.7384 | 0
Transmille | TcindicatorS | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0844 | 0 | 2.0805 | 0
Transmille | TcindicatorT | -418 | -238 | 0.1 | °F | °F |  |  |  | 3.9624 | 0 | 3.9602 | 0
Transmille | TcindicatorT | -238 | 752 | 0.1 | °F | °F |  |  |  | 1.9193 | 0 | 1.9156 | 0
Transmille | TcindicatorT | -418 | -238 | 1 | °F | °F |  |  |  | 4.0103 | 0 | 4.0017 | 0
Transmille | TcindicatorT | -238 | 752 | 1 | °F | °F |  |  |  | 2.0141 | 0 | 1.9999 | 0
Transmille | TcindicatorT | -250 | -150 | 0.1 | °C | °C |  |  |  | 2.2146 | 0 | 2.2134 | 0
Transmille | TcindicatorT | -150 | 400 | 0.1 | °C | °C |  |  |  | 1.0694 | 0 | 1.0674 | 0
Transmille | TcindicatorT | -250 | -150 | 1 | °C | °C |  |  |  | 2.2914 | 0 | 2.2868 | 0
Transmille | TcindicatorT | -150 | 400 | 1 | °C | °C |  |  |  | 1.2194 | 0 | 1.2121 | 0

## Unc. Calculator
|  |  | RED=Key Calculator Component |  | 5502A | Standard1
|  |  | GREEN=Formula Uses a Table Lookup |  | 3458A | Standard2 |  | Troubleshooter Lookup | [11 empty] | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [44 empty] | ACV Freq. Rng Low | ACI Freq. Rng Low | ACA10T. Rng low | ACA20T. Rng low | ACA50T. Rng Low | Deg. C, TC K | Deg. C, TC J | Deg. F, TC K | Deg. F, TC J | DCW | ACW
|  |  | ORANGE= END USER RAW INPUTS |  | 34420A | Standard3 |  | Rdg | L.up ID | Standard | Base | Mult | 2nd B | 2nd Mult | Rep. Mult | Unc. Units | Result |  | Standard6 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [19 empty] | 1000 | [23 empty] | Standard1 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1163:$I$1174 | $I$1151:$I$1162 | $I$1246:$I$1309 | $I$1310:$I$1378
|  |  | BLUE= PART OF MAIN LOOKUP ID |  | Transmille | Standard4 |  | 0.003 | DCA0.000010 | Standard1 | 5.8001 | 1.6501 | 5.8001 | 1.5718 | 0 | µA | 5.8050503µA |  | Standard5 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [43 empty] | Standard2 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$888:$I$895 | $I$888:$I$895 | #N/A | #N/A
[6 empty] | Standard5 | [12 empty] | Standard4 | $D$4:$D$28 | $D$80:$D$119 | $D$29:$D$58 | $D$120:$D$149 | #N/A | $D$150:$D$164 | $D$165:$D$169 | $D$59:$D$64 | $D$65:$D$69 | $D$70:$D$79 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [43 empty] | Standard3 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[6 empty] | Standard6 | [12 empty] | Standard3 | #N/A | #N/A | #N/A | #N/A | $D$39:$D$73 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [21 empty] | Removed exception for Hz over 50 | [21 empty] | Standard4 | $I$80:$I$119 | $I$120:$I$149 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Col. Index # for frequency Lookup | Input Unit | Unc units
[6 empty] | Standard7 | [12 empty] | Standard2 | $D$4:$D$35 | $D$156:$D$552 | $D$36:$D$95 | $D$553:$D$750 | $D$96:$D$155 | #N/A | $D$751:$D$793 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | $D$806:$D$813 | $D$800:$D$805 | $D$888:$D$895 | $D$882:$D$887 | #N/A | #N/A | Unit Multiplier | [19 empty] | 2ndary Unit Multiplier | [22 empty] | Standard5 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | V11 | CN11
[19 empty] | Standard1 | $D$4:$D$40 | $D$230:$D$439 | $D$41:$D$81 | $D$471:$D$735 | $D$151:$D$229 | $D$961:$D$1030 | $D$440:$D$470 | $D$82:$D$107 | $D$108:$D$129 | $D$130:$D$150 | $D$736:$D$826 | $D$827:$D$891 | $D$892:$D$960 | $D$1055:$D$1066 | $D$1043:$D$1054 | $D$1163:$D$1174 | $D$1151:$D$1162 | $D$1246:$D$1309 | $D$1310:$D$1378 | Base unit | [19 empty] | Base Unit | [18 empty] | RESOLUTION CONVERTER |  |  |  | Standard6 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Tolerance | Converted |  |  |  |  | USER INPUTS |  |  |  |  | Function Selector | [8 empty] | 2: MAIN INPUT FUNCTION LOOKUP | V | V | Amps | Amps | kOhms | mF | kHz | Amps | Amps | Amps | Amps | Amps | Amps | °C | °C | °F | °F | Watts | Watts |  |  |  |  | LOOKED UP UNC. COMPONENTS |  |  |  |  | REPEAT USER INPUTS |  |  |  |  | CONVERTED REPEAT ENTRIES | [5 empty] | "0.000 to 0.001" | [16 empty] | Repeat Multiplier Ratio
| Uncert | Uncert | Main Lookup ID | Standard Used | Unc. | Res. | In. Value | Units | Range 2 | Units | L.upbldr | L.upbldr | L.upbldr | L.upbldr | L.upbldr | L.upbldr | Func.Sel. | Main L.up | Converted Rdg | DCV | ACV | DCA | ACA | Resistance | Capacitance | Frequency | DCA 10 TURN | DCA 20 TURN | DCA 50 TURN | ACA 10 TURN | ACA 20 TURN | ACA 50 TURN | CDK | CDJ | FDK | FDJ | DCW | ACW | DCV | ACV | DCA | ACA | Resistance | Capacitance | Frequency | DCA 10 TURN | DCA 20 TURN | DCA 50 TURN | ACA 10 TURN | ACA 20 TURN | ACA 50 TURN | CDK | CDJ | FDK | FDJ | DCW | ACW | Primary Conv. Factor | kHz | Converted Rdg | 2ndary L. Up | Base | Mult | 2nd B. | 2nd Mult. | Rept Mult | R1 | R2 | R3 | R4 | R5 | R1(Mult) | R2(Mult) | R3(Mult) | R4(Mult) | R5(Mult) | Unc. Units | Input | Conv. Txt | Result |  | Unc. Table Address for L. Ups | ACV | ACA | ACA 10 TURN | ACA 20 TURN | ACA 50 TURN | CDK | CDJ | FDK | FDJ | DCW | ACW |  | Input | Uncert. | Multiplier
3.3 | 0.14 | 140 | ACV0.00010.0330.045 | 5502A | 140 µV | 0.0001 | 300 | mV | 45 | Hz | m | V | AC | ACV |  |  | ACV | 0.033 | 0.3 | 0 | 0.033 | 0.033 | 0.033 | 0.11 | 0.11 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.001 | 0.045 | 0.045 | 58.586 | 276.03 | 58.483 | 276.21 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 2 | 0.001 | 1e-06
0.053 | 0.0025 | 2.5 | ACV0.0013.30.045 | 5502A | 2.5 mV | 0.001 | 5 | V | 500 | Hz |  | V | AC | ACV |  |  | ACV | 3.3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.5 | 0.045 | 0.78957 | 0.34543 | 0.78891 | 0.34543 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.001 | 0.001 | 0.001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 3 | 1 | 0.001
0.103 | 0.0025 | 2.5 | ACV0.0013.30.045 | 5502A | 2.5 mV | 0.001 | 5 | V | 1 | kHz |  | V | AC | ACV |  |  | ACV | 3.3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.045 | 0.78957 | 0.34543 | 0.78891 | 0.34543 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.001 | 0.001 | 0.001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 4 | 1 | 0.001
0.53 | 0.058 | 58 | ACV0.01330.045 | 5502A | 58 mV | 0.01 | 50 | V | 45 | Hz |  | V | AC | ACV |  |  | ACV | 33 | 50 | 33 | 33 | 20.5 | 11 | 33 | 33 | 12 | 30 | 20 | #N/A | 30 | 22 | #N/A | -25 | -30 | -13 | -22 | 10.9 | 10.9 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 12 | 0.92331 | 10.834 | 0.9234 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.01 | 0.01 | 0.01 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 5 | 1 | 0.001
1.03 | 0.058 | 58 | ACV0.01330.045 | 5502A | 58 mV | 0.01 | 50 | V | 1 | kHz |  | V | AC | ACV |  |  | ACV | 33 | 50 | 33 | 33 | 20.5 | 11 | 33 | 33 | 12 | 30 | 20 | #N/A | 30 | 22 | #N/A | -25 | -30 | -13 | -22 | 10.9 | 10.9 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.045 | 12 | 0.92331 | 10.834 | 0.9234 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.01 | 0.01 | 0.01 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 6 | 1 | 0.001
3.3 | 0.3 | 300 | ACV0.1330.045 | 5502A | 300 mV | 0.1 | 300 | V | 45 | Hz |  | V | AC | ACV |  |  | ACV | 33 | 300 | 33 | 33 | 20.5 | 11 | 110 | 110 | 120 | 205 | 220 | 150 | 110 | 220 | 150 | 120 | 150 | 248 | -22 | 109 | 109 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 43.309 | 0.841 | 43.21 | 0.84115 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 7 | 1 | 0.001
5.3 | 0.67 | 670 | ACV0.13300.045 | 5502A | 670 mV | 0.1 | 500 | V | 500 | Hz |  | V | AC | ACV |  |  | ACV | 330 | 500 | 330 | 330 | 20.5 | 11 | 330 | 110 | 120 | 205 | 410 | 150 | 110 | 220 | 150 | 120 | 150 | 248 | 302 | 330 | 363 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.5 | 0.045 | 510.18 | 0.32785 | 510.26 | 0.32724 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 8 | 1 | 0.001
10.3 | 0.67 | 670 | ACV0.13300.045 | 5502A | 670 mV | 0.1 | 500 | V | 1 | kHz |  | V | AC | ACV |  |  | ACV | 330 | 500 | 330 | 330 | 20.5 | 11 | 330 | 110 | 120 | 205 | 410 | 150 | 110 | 220 | 150 | 120 | 150 | 248 | 302 | 330 | 363 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.045 | 510.18 | 0.32785 | 510.26 | 0.32724 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 9 | 1 | 0.001
13 | 5.8 | 5800 | ACV13300.045 | 5502A | 5800 mV | 1 | 1000 | V | 45 | Hz |  | V | AC | ACV |  |  | ACV | 330 | 1000 | 1000 | 330 | 20.5 | 11 | 330 | 110 | 120 | 205 | 410 | 550 | 110 | 220 | 550 | 1000 | 760 | 248 | 302 | 330 | 990 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 5790.3 | 0.041506 | 5790.4 | 0.041123 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 1 | 1 | 1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 10 | 1 | 0.001
0.06 | 0.0068 | 0.0068 | Frequency0.000010 | 5502A | 0.0068 Hz | 1e-05 | 45 | Hz | 1 | V | H | z | AC | ACz |  | Frequency | Frequency | 0 | 0.045 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.0109 | 0.0109 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 |  | 0.0060576 | 0.015807 | 0.005881 | 0.012606 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Hz | 0.01 | 0.01 | 1e-05 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 11 | 1 | 1
0.06 | 0.02 | 20 | Frequency0.0112 | 5502A | 20 Hz | 0.01 | 50 | kHz | 5 | V | H | z | AC | ACz |  | Frequency | Frequency | 12 | 50 | 33 | 33 | 20.5 | 11 | 33 | 33 | 12 | 30 | 20 | #N/A | 30 | 22 | #N/A | -25 | -30 | -13 | -22 | 10.9 | 10.9 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 5 |  | 19 | 0.028 | 18.255 | 0.027641 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Hz | 0.01 | 0.01 | 0.01 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 12 | 1000 | 1
0.007 | 0.00072 | 0.72 | DCV0.0013 | 5502A | 0.72 mV | 0.001 | 5 | V |  |  |  | V | DC | DCV |  |  | DCV | 3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.47784 | 0.047829 | 0.47304 | 0.047859 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.001 | 0.001 | 0.001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 13 | 1 | 0.001
0.5 | 0.062 | 62 | DCV0.133 | 5502A | 62 mV | 0.1 | 300 | V |  |  |  | V | DC | DCV |  |  | DCV | 33 | 300 | 33 | 33 | 20.5 | 11 | 110 | 110 | 120 | 205 | 220 | 150 | 110 | 220 | 150 | 120 | 150 | 248 | -22 | 109 | 109 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 57.654 | 0.012974 | 57.637 | 0.012872 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 14 | 1 | 0.001
3 | 0.58 | 580 | DCV11000 | 5502A | 580 mV | 1 | 1000 | V |  |  |  | V | DC | DCV |  |  | DCV | 1000 | 1000 | 1000 | 330 | 20.5 | 11 | 330 | 110 | 120 | 205 | 410 | 550 | 110 | 220 | 550 | 1000 | 760 | 248 | 302 | 330 | 990 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 1 | 1 | 1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 15 | 1 | 0.001
3 | 0.58 | 580 | DCV11000 | 5502A | 580 mV | 1 | 1000 | V |  |  |  | V | DC | DCV |  |  | DCV | 1000 | 1000 | 1000 | 330 | 20.5 | 11 | 330 | 110 | 120 | 205 | 410 | 550 | 110 | 220 | 550 | 1000 | 760 | 248 | 302 | 330 | 990 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 1 | 1 | 1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 16 | 1 | 0.001
0.06 | 0.0068 | 0.0068 | Frequency0.000010 | 5502A | 0.0068 Hz | 1e-05 | 45 | Hz | 3 | V | H | z | AC | ACz |  | Frequency | Frequency | 0 | 0.045 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.0109 | 0.0109 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 3 |  | 0.0060576 | 0.015807 | 0.005881 | 0.012606 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Hz | 0.01 | 0.01 | 1e-05 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 17 | 1 | 1
0.06 | 0.02 | 20 | Frequency0.0112 | 5502A | 20 Hz | 0.01 | 50 | kHz | 30 | V | H | z | AC | ACz |  | Frequency | Frequency | 12 | 50 | 33 | 33 | 20.5 | 11 | 33 | 33 | 12 | 30 | 20 | #N/A | 30 | 22 | #N/A | -25 | -30 | -13 | -22 | 10.9 | 10.9 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 30 |  | 19 | 0.028 | 18.255 | 0.027641 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Hz | 0.01 | 0.01 | 0.01 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 18 | 1000 | 1
0.2 | 0.059 | 59 | DCV0.00010 | 5502A | 59 µV | 0.0001 | 30 | mV |  |  | m | V | DC | DCV |  |  | DCV | 0 | 0.03 | 0 | 0 | 0.0033 | 0.0033 | 0.011 | 0.011 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.0109 | 0.0109 | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 58.11 | 17.06 | 58.091 | 17.007 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 19 | 0.001 | 1e-06
0.5 | 0.063 | 63 | DCV0.00010 | 5502A | 63 µV | 0.0001 | 300 | mV |  |  | m | V | DC | DCV |  |  | DCV | 0 | 0.3 | 0 | 0.033 | 0.033 | 0.033 | 0.11 | 0.11 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 58.11 | 17.06 | 58.091 | 17.007 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 20 | 0.001 | 1e-06
0.7 | 0.076 | 76 | DCV0.00010.33 | 5502A | 76 µV | 0.0001 | 600 | mV |  |  | m | V | DC | DCV |  |  | DCV | 0.33 | 0.6 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 47.82 | 47.701 | 47.342 | 47.73 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 21 | 0.001 | 1e-06
4.7 | 0.097 | 97 | Resistance0.00010.33 | 5502A | 97 mΩ | 0.0001 | 500 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0.33 | 0.5 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 52.968 | 87.892 | 52.867 | 87.789 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 22 | 1 | 0.001
0.32 | 0.026 | 26 | Resistance1011000 | 5502A | 26 kΩ | 10 | 19 | MΩ |  |  | M | O | DC | O |  | Resistance | Resistance | 11000 | 19000 | 1000 | 330 | 20.5 | 11 | 11000 | 110 | 2000 | 205 | 410 | 1025 | 110 | 220 | 550 | 1372 | 1200 | 2500 | 2192 | 11000 | 11200 | 1 | 1 | 1 | 1 | 1000 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1000 | 1 |  |  | 4.2484 | 0.0011259 | 4.2478 | 0.0011259 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | kΩ | 0.01 | 0.01 | 10 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 23 | 1000000 | 1000
13 | 3.8 | 3.8 | Capacitance0.0000010.00033 | 5502A | 3.8 nF | 1e-06 | 900 | nF |  |  | n | F | DC | F |  | Capacitance | Capacitance | 0.00033 | 0.0009 | 0 | 0 | 0.00033 | 0.00033 | 0 | 0.00033 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.000109 | 0.000109 | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 |  |  | 1.2438 | 2860.2 | 1.2501 | 2835.6 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nF | 1 | 1 | 1e-06 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 24 | 1e-09 | 1e-09
0 | 0.062 | 62 | Resistance0.00010.011 | 5502A | 62 mΩ | 0.0001 | 25 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0.011 | 0.025 | 0 | 0 | 0.0033 | 0.0033 | 0.011 | 0.011 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.0109 | 0.0109 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 60.487 | 46.304 | 60.481 | 46.04 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 25 | 1 | 0.001
0 | 0.11 | 110 | Resistance0.00010.11 | 5502A | 110 mΩ | 0.0001 | 250 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0.11 | 0.25 | 0 | 0.033 | 0.033 | 0.033 | 0.11 | 0.11 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 61.466 | 184.35 | 61.46 | 183.05 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 26 | 1 | 0.001
0.022 | 0.0006 | 600 | DCV0.0010.33 | 5502A | 600 µV | 0.001 | 2 | V |  |  |  | V | DC | DCV |  |  | DCV | 0.33 | 2 | 0.33 | 0.33 | 0.33 | 1.1 | 1.1 | 1.1 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 577.17 | 10.824 | 576.99 | 10.721 | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.001 | 0.001 | 0.001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 27 | 1 | 1e-06
0.08 | 0.0069 | 6.9 | ACA0.000010.000330.045 | 5502A | 6.9 µA | 1e-05 | 3 | mA | 45 | Hz | m | A | AC | ACA |  |  | ACA | 0.00033 | 0.003 | 0 | 0 | 0.00033 | 0.00033 | 0 | 0.0011 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.00109 | 0.00109 | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 0.001 | 0.001 | 0.045 | 0.045 | 5.6654 | 409.29 | 5.6651 | 409.25 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µA | 0.01 | 0.01 | 1e-05 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 28 | 0.001 | 1e-06
0.78 | 0.12 | 0.12 | ACA0.000010.0331 | 5502A | 0.12 mA | 1e-05 | 50 | mA | 1 | kHz | m | A | AC | ACA |  |  | ACA | 0.033 | 0.05 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.0109 | 0.0109 | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 0.001 | 1 | 1 | 1 | 0.058 | 1.2 | 0.057825 | 1.1583 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.01 | 0.01 | 1e-05 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 29 | 0.001 | 0.001
6.3 | 4 | 4 | ACA0.00010.331 | 5502A | 4 mA | 0.0001 | 400 | mA | 1 | kHz | m | A | AC | ACA |  |  | ACA | 0.33 | 0.4 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1.2 | 6.9 | 1.1555 | 6.9278 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 30 | 0.001 | 0.001
0.063 | 0.0052 | 5.2 | ACA0.00130.045 | 5502A | 5.2 mA | 0.001 | 4 | A | 45 | Hz |  | A | AC | ACA |  |  | ACA | 3 | 4 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 2.4 | 0.69 | 2.3709 | 0.69112 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.001 | 0.001 | 0.001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 31 | 1 | 0.001
0.17 | 0.014 | 14 | ACA0.0130.1 | 5502A | 14 mA | 0.01 | 9 | A | 1 | kHz |  | A | AC | ACA |  |  | ACA | 3 | 9 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 5.1972 | 0.99041 | 5.1951 | 0.99036 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.01 | 0.01 | 0.01 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 32 | 1 | 0.001
0.06 | 0.0058 | 5.8 | DCA0.000010.00033 | 5502A | 5.8 µA | 1e-05 | 3 | mA |  |  | m | A | DC | DCA |  |  | DCA | 0.00033 | 0.003 | 0 | 0 | 0.00033 | 0.00033 | 0 | 0.0011 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.00109 | 0.00109 | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 0.001 | 1 |  |  | 5.7989 | 9.0557 | 5.7981 | 8.8116 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µA | 0.01 | 0.01 | 1e-05 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 33 | 0.001 | 1e-06
0.53 | 0.011 | 11 | DCA0.000010.033 | 5502A | 11 µA | 1e-05 | 50 | mA |  |  | m | A | DC | DCA |  |  | DCA | 0.033 | 0.05 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.0109 | 0.0109 | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 0.001 | 1 |  |  | 5.2853 | 109.48 | 5.266 | 109.51 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µA | 0.01 | 0.01 | 1e-05 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 34 | 0.001 | 1e-06
4.3 | 0.23 | 0.23 | DCA0.00010.33 | 5502A | 0.23 mA | 0.0001 | 400 | mA |  |  | m | A | DC | DCA |  |  | DCA | 0.33 | 0.4 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0.109 | 0.109 | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 0.001 | 1 |  |  | 0.060249 | 0.43614 | 0.060204 | 0.43607 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.1 | 0.1 | 0.0001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 35 | 0.001 | 0.001
0.043 | 0.0037 | 3.7 | DCA0.0013 | 5502A | 3.7 mA | 0.001 | 4 | A |  |  |  | A | DC | DCA |  |  | DCA | 3 | 4 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.055271 | 0.90973 | 0.052344 | 0.90974 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.001 | 0.001 | 0.001 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 36 | 1 | 0.001
0.12 | 0.01 | 10 | DCA0.013 | 5502A | 10 mA | 0.01 | 9 | A |  |  |  | A | DC | DCA |  |  | DCA | 3 | 9 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1.2 | 1 | #N/A | #N/A | 1 | #N/A | #N/A | -25 | -30 | -13 | -22 | 1.09 | 1.09 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 4.4629 | 0.64947 | 4.4597 | 0.64887 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.01 | 0.01 | 0.01 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 37 | 1 | 0.001
1 | 0.2 | 0.2 | CDK0.1-25 | 5502A | 0.2 °C | 0.1 | 0 | °C |  |  |  | CD | K | CDK |  |  | CDK | -25 | 0 | 0 | 0 | 0 | 0 | 0 | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -25 | -30 | -13 | -22 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.20428 | 0 | 0.19653 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | °C | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 38 | 1 | 1
1.4 | 0.23 | 0.23 | CDK0.1-100 | 5502A | 0.23 °C | 0.1 | 40 | °C |  |  |  | CD | K | CDK |  |  | CDK | -25 | 40 | 33 | 33 | 20.5 | 11 | 33 | 33 | 12 | 30 | 20 | #N/A | 30 | 22 | #N/A | -25 | -30 | -13 | -22 | 10.9 | 10.9 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.22869 | 0 | 0.21853 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | °C | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 39 | 1 | 1
5 | 0.31 | 0.31 | CDK0.1120 | 5502A | 0.31 °C | 0.1 | 400 | °C |  |  |  | CD | K | CDK |  |  | CDK | 120 | 400 | 330 | 330 | 20.5 | 11 | 330 | 110 | 120 | 205 | 220 | 150 | 110 | 220 | 150 | 120 | 150 | 248 | 302 | 330 | 363 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.31026 | 0 | 0.30758 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | °C | 0.1 | 0.1 | 0.1 | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 40 | 1 | 1
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 41 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 42 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 43 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 44 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 45 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 46 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 47 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 48 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 49 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 50 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 51 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 52 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 53 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 54 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 55 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 56 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 57 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 58 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 59 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 60 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 61 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 62 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 63 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 64 |  | #N/A
0 | 0 | * | #N/A | 3458A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard2 | $B$4:$P$956 | $I$156:$I$552 | $I$553:$I$750 | #N/A | #N/A | #N/A | $I$806:$I$813 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | $I$800:$I$805 | 65 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 66 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 67 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 68 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 69 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 70 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 71 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 72 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 73 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 74 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 75 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 76 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 77 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 78 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 79 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 80 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 81 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 82 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 83 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 84 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 85 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 86 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 87 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 88 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 89 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 90 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 91 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 92 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 93 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 94 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 95 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 96 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 97 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 98 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 99 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 100 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 101 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 102 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 103 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 104 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 105 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 106 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 107 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 108 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 109 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 110 |  | #N/A
0 | 0 | * | #N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$1376 | $I$230:$I$439 | $I$471:$I$735 | $I$736:$I$826 | $I$827:$I$891 | $I$892:$I$960 | $I$1055:$I$1066 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | $I$1043:$I$1054 | 111 |  | #N/A
Note on S1: Lookup table thatTranslates Standard Name to Generic name. i.e. "5502A" to "Standard1". See Colum "CQ"
Note on BB7: Columns AU-BA convert the reading entered into the standard unit of the table for each function
Note on BW7: Converts the 2ndary units (usually frequency) to the standard table format (kHz)
Note on T9: User inputs. Resolution is converted from "0.0" format to "0.1" format in Column CA-CE, and Column T references that
Note on Y9: Columns Y-AD select which function (column C) is used based on the inputs
Note on AH9: Columns AG- AM determine in what range the primary unit falls (Ex: 5VDC falls in the 3.3-30V Range, so it returns 3.3, as it is set to show the "Range Low" part as a marker
Note on DI9: This is where the ratio between the Input units and the Uncertainty Units is created to generate the repeat multiplier in CC11
Note on Q10: Constructed input ID in the form of Function,resolution,InputRangeLOW,InputRange2LOW
Note on T10: From Column CE
Note on U10: Returns the Absolute value of the reading (filters out Negative Numbers)
Note on Z10: This is where special characters get filtered out and converted to whatever we designate. (i.e. Ω changes to O)
Note on AE10: Determines which category in Column "C" the reading falls under
Note on AF10: Column AE Retrieves the appropriate "Range Low" result from Columns AG-AM, based on the Function Selector in Column AD. It is an Hlookup formula, desgined to return the Nth row down, where N is the number of rows
Note on AG10: Column AF Converts Column AE to whatever the standard unit is in the table, based on a multiplier in Column BI
Note on BV10: Looks up the correct converted Unit based on the function selector
Note on BW10: 2ndary unit conv. factor
Note on BY10: Returns the appropriate "Range 2 Low" value (usually frequency) from the table. Part of the INPUT ID that looks up the Uncertainty components
Note on CR10: Final Resolution Input: Part of the INPUT ID used for uncertainty Lookup

## Standard1
Updated: | 2021-11-18
|  | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.00000010 | DCV | 0 | 0.33 | 1e-07 | V | µV |  |  |  | 3.6 | 69 | 3.4727 | 69.314
| DCV0.0000010 | DCV | 0 | 0.33 | 1e-06 | V | µV |  |  |  | 3.6 | 69.151 | 3.5196 | 69.191
| DCV0.000010 | DCV | 0 | 0.33 | 1e-05 | V | µV |  |  |  | 6.7738 | 61.287 | 6.7344 | 61.342
| DCV0.00010 | DCV | 0 | 0.33 | 0.0001 | V | µV |  |  |  | 58.11 | 17.06 | 58.091 | 17.007
| DCV0.0010 | DCV | 0 | 0.33 | 0.001 | V | µV |  |  |  | 580.02 | 1.8105 | 580.01 | 1.785
| DCV0.010 | DCV | 0 | 0.33 | 0.01 | V | µV |  |  |  | 5800 | 0.1888 | 5800 | 0.17859
| DCV0.10 | DCV | 0 | 0.33 | 0.1 | V | mV |  |  |  | 58 | 2.1943e-05 | 58 | 1.7859e-05
| DCV10 | DCV | 0 | 0.33 | 1 | V | mV |  |  |  | 580 | 3.4192e-06 | 580 | 1.7859e-06
| DCV100 | DCV | 0 | 0.33 | 10 | V | mV |  |  |  | 5800 | 8.319e-07 | 5800 | 1.7859e-07
| DCV0.0000010.33 | DCV | 0.33 | 3.3 | 1e-06 | V | µV |  |  |  | 5.9 | 58 | 5.7708 | 57.787
| DCV0.000010.33 | DCV | 0.33 | 3.3 | 1e-05 | V | µV |  |  |  | 6.8066 | 57.529 | 6.4957 | 57.593
| DCV0.00010.33 | DCV | 0.33 | 3.3 | 0.0001 | V | µV |  |  |  | 47.82 | 47.701 | 47.342 | 47.73
| DCV0.0010.33 | DCV | 0.33 | 3.3 | 0.001 | V | µV |  |  |  | 577.17 | 10.824 | 576.99 | 10.721
| DCV0.010.33 | DCV | 0.33 | 3.3 | 0.01 | V | µV |  |  |  | 5799.8 | 1.1473 | 5799.7 | 1.1022
| DCV0.10.33 | DCV | 0.33 | 3.3 | 0.1 | V | µV |  |  |  | 58000 | 0.12833 | 5799.8 | 0.11025
| DCV0.000013 | DCV | 3 | 33.0001 | 1e-05 | V | mV |  |  |  | 0.059 | 0.058 | 0.058348 | 0.057919
| DCV0.00013 | DCV | 3 | 33.0001 | 0.0001 | V | mV |  |  |  | 0.067758 | 0.057661 | 0.064652 | 0.057725
| DCV0.0013 | DCV | 3 | 33.0001 | 0.001 | V | mV |  |  |  | 0.47784 | 0.047829 | 0.47304 | 0.047859
| DCV0.013 | DCV | 3 | 33.0001 | 0.01 | V | mV |  |  |  | 5.7715 | 0.010868 | 5.7698 | 0.010764
| DCV0.13 | DCV | 3 | 33.0001 | 0.1 | V | mV |  |  |  | 57.998 | 0.0011519 | 57.997 | 0.0011068
| DCV13 | DCV | 3 | 33.0001 | 1 | V | mV |  |  |  | 580 | 0.00012879 | 580 | 0.00011071
| DCV0.000133 | DCV | 33 | 330.001 | 0.0001 | V | mV |  |  |  | 0.59 | 0.064 | 0.57834 | 0.063644
| DCV0.00133 | DCV | 33 | 330.001 | 0.001 | V | mV |  |  |  | 0.67471 | 0.063403 | 0.64578 | 0.063463
| DCV0.0133 | DCV | 33 | 330.001 | 0.01 | V | mV |  |  |  | 4.6629 | 0.053693 | 4.6153 | 0.053731
| DCV0.133 | DCV | 33 | 330.001 | 0.1 | V | mV |  |  |  | 57.654 | 0.012974 | 57.637 | 0.012872
| DCV133 | DCV | 33 | 330.001 | 1 | V | mV |  |  |  | 579.97 | 0.0013757 | 579.96 | 0.0013305
| DCV1033 | DCV | 33 | 330.001 | 10 | V | mV |  |  |  | 5800 | 0.00015118 | 5800 | 0.0001331
| DCV0.001330 | DCV | 330 | 1000.01 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
| DCV0.01330 | DCV | 330 | 1000.01 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
| DCV0.1330 | DCV | 330 | 1000.01 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
| DCV1330 | DCV | 330 | 1000.01 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
| DCV10330 | DCV | 330 | 1000.01 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
| DCV0.0011000 | DCV | 1000 | 1001 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
| DCV0.011000 | DCV | 1000 | 1001 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
| DCV0.11000 | DCV | 1000 | 1001 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
| DCV11000 | DCV | 1000 | 1001 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
| DCV101000 | DCV | 1000 | 1001 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
| DCA0.0000000010 | DCA | 0 | 0.00033001 | 1e-09 | A | µA |  |  |  | 0.024 | 180 | 0.023107 | 175.06
| DCA0.000000010 | DCA | 0 | 0.00033001 | 1e-08 | A | µA |  |  |  | 0.024 | 173.53 | 0.023814 | 173.54
| DCA0.00000010 | DCA | 0 | 0.00033001 | 1e-07 | A | µA |  |  |  | 0.062453 | 112.75 | 0.062384 | 112.54
| DCA0.0000010 | DCA | 0 | 0.00033001 | 1e-06 | A | µA |  |  |  | 0.58048 | 15.83 | 0.58045 | 15.637
| DCA0.000010 | DCA | 0 | 0.00033001 | 1e-05 | A | µA |  |  |  | 5.8001 | 1.6501 | 5.8001 | 1.5718
| DCA0.00010 | DCA | 0 | 0.00033001 | 0.0001 | A | µA |  |  |  | 58 | 0.18852 | 58 | 0.15719
| DCA0.0010 | DCA | 0 | 0.00033001 | 0.001 | A | µA |  |  |  | 580 | 0.028252 | 580 | 0.015719
| DCA0.010 | DCA | 0 | 0.00033001 | 0.01 | A | mA |  |  |  | 5.8 | 6.5851e-06 | 5.8 | 1.5719e-06
| DCA0.10 | DCA | 0 | 0.00033001 | 0.1 | A | mA |  |  |  | 58 | 2.1625e-06 | 58 | 1.572e-07
| DCA0.000000010.00033 | DCA | 0.00033 | 0.0033001 | 1e-08 | A | µA |  |  |  | 0.068 | 150 | 0.057598 | 152.89
| DCA0.00000010.00033 | DCA | 0.00033 | 0.0033001 | 1e-07 | A | µA |  |  |  | 0.076187 | 148.47 | 0.07332 | 149.02
| DCA0.0000010.00033 | DCA | 0.00033 | 0.0033001 | 1e-06 | A | µA |  |  |  | 0.56792 | 73.551 | 0.56576 | 73.312
| DCA0.000010.00033 | DCA | 0.00033 | 0.0033001 | 1e-05 | A | µA |  |  |  | 5.7989 | 9.0557 | 5.7981 | 8.8116
| DCA0.00010.00033 | DCA | 0.00033 | 0.0033001 | 0.0001 | A | µA |  |  |  | 58 | 0.98192 | 58 | 0.88328
| DCA0.00000010.0033 | DCA | 0.0033 | 0.033001 | 1e-07 | A | µA |  |  |  | 0.21 | 150 | 0.201 | 145.79
| DCA0.0000010.0033 | DCA | 0.0033 | 0.033001 | 1e-06 | A | µA |  |  |  | 0.44201 | 139.58 | 0.43216 | 139.79
| DCA0.000010.0033 | DCA | 0.0033 | 0.033001 | 1e-05 | A | µA |  |  |  | 5.6422 | 61.544 | 5.6369 | 61.47
| DCA0.00010.0033 | DCA | 0.0033 | 0.033001 | 0.0001 | A | µA |  |  |  | 57.983 | 7.2053 | 57.981 | 7.143
| DCA0.0010.0033 | DCA | 0.0033 | 0.033001 | 0.001 | A | µA |  |  |  | 580 | 0.74076 | 580 | 0.71564
| DCA0.0000010.033 | DCA | 0.033 | 0.33001 | 1e-06 | A | µA |  |  |  | 3 | 120 | 2.9414 | 115.33
| DCA0.000010.033 | DCA | 0.033 | 0.33001 | 1e-05 | A | µA |  |  |  | 5.2853 | 109.48 | 5.266 | 109.51
| DCA0.00010.033 | DCA | 0.033 | 0.33001 | 0.0001 | A | µA |  |  |  | 56.995 | 42.585 | 56.984 | 42.551
| DCA0.0010.033 | DCA | 0.033 | 0.33001 | 0.001 | A | µA |  |  |  | 579.89 | 4.761 | 579.88 | 4.7411
| DCA0.010.033 | DCA | 0.033 | 0.33001 | 0.01 | A | µA |  |  |  | 5800 | 0.4827 | 5800 | 0.47471
| DCA0.000010.33 | DCA | 0.33 | 3.0001 | 1e-05 | A | mA |  |  |  | 0.052 | 0.44 | 0.051 | 0.43873
| DCA0.00010.33 | DCA | 0.33 | 3.0001 | 0.0001 | A | mA |  |  |  | 0.060249 | 0.43614 | 0.060204 | 0.43607
| DCA0.0010.33 | DCA | 0.33 | 3.0001 | 0.001 | A | mA |  |  |  | 0.50421 | 0.32729 | 0.50424 | 0.32696
| DCA0.010.33 | DCA | 0.33 | 3.0001 | 0.01 | A | mA |  |  |  | 5.784 | 0.058643 | 5.7841 | 0.058299
| DCA0.10.33 | DCA | 0.33 | 3.0001 | 0.1 | A | mA |  |  |  | 57.998 | 0.0060523 | 57.998 | 0.0059106
| DCA0.00013 | DCA | 3 | 11.001 | 0.0001 | A | mA |  |  |  | -0.024 | 0.92 | -0.024619 | 0.91523
| DCA0.0013 | DCA | 3 | 11.001 | 0.001 | A | mA |  |  |  | 0.055271 | 0.90973 | 0.052344 | 0.90974
| DCA0.013 | DCA | 3 | 11.001 | 0.01 | A | mA |  |  |  | 4.4629 | 0.64947 | 4.4597 | 0.64887
| DCA0.13 | DCA | 3 | 11.001 | 0.1 | A | mA |  |  |  | 57.764 | 0.1006 | 57.764 | 0.099909
| DCA0.000111 | DCA | 11 | 20.5 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
| DCA0.00111 | DCA | 11 | 20.5 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
| DCA0.0111 | DCA | 11 | 20.5 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
| DCA0.111 | DCA | 11 | 20.5 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
| DCA0.000120.5 | DCA | 20.5 | 20.6 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
| DCA0.00120.5 | DCA | 20.5 | 20.6 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
| DCA0.0120.5 | DCA | 20.5 | 20.6 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
| DCA0.120.5 | DCA | 20.5 | 20.6 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
| DCA 10 turn0.00011 | DCA 10 turn | 1 | 11.001 | 0.0001 | A | mA |  |  |  | 0.082 | 2 | 0.080875 | 1.9947
| DCA 10 turn0.0011 | DCA 10 turn | 1 | 11.001 | 0.001 | A | mA |  |  |  | 0.16836 | 1.9875 | 0.16627 | 1.9876
| DCA 10 turn0.011 | DCA 10 turn | 1 | 11.001 | 0.01 | A | mA |  |  |  | 4.4751 | 1.6631 | 4.4718 | 1.6632
| DCA 10 turn0.11 | DCA 10 turn | 1 | 11.001 | 0.1 | A | mA |  |  |  | 57.372 | 0.40227 | 57.371 | 0.40202
| DCA 10 turn11 | DCA 10 turn | 1 | 11.001 | 1 | A | mA |  |  |  | 577.32 | 0.041721 | 577.32 | 0.041613
| DCA 10 turn0.00111 | DCA 10 turn | 11 | 30.001 | 0.001 | A | mA |  |  |  | 1.6 | 2 | 0.12764 | 2.0471
| DCA 10 turn0.0111 | DCA 10 turn | 11 | 30.001 | 0.01 | A | mA |  |  |  | 1.6 | 2.0127 | 1.1038 | 2.0235
| DCA 10 turn0.111 | DCA 10 turn | 11 | 30.001 | 0.1 | A | mA |  |  |  | 49.094 | 1.1796 | 49.066 | 1.1772
| DCA 10 turn111 | DCA 10 turn | 11 | 30.001 | 1 | A | mA |  |  |  | 576.17 | 0.15058 | 576.16 | 0.14877
| DCA 10 turn1011 | DCA 10 turn | 11 | 30.001 | 10 | A | mA |  |  |  | 5773.4 | 0.015656 | 5773.4 | 0.014925
| DCA 10 turn10011 | DCA 10 turn | 11 | 30.001 | 100 | A | mA |  |  |  | 57736 | 0.0017851 | 57735 | 0.0014925
| DCA 10 turn0.0130 | DCA 10 turn | 30 | 110.01 | 0.01 | A | mA |  |  |  | 4 | 2.1 | 2.509 | 2.1122
| DCA 10 turn0.130 | DCA 10 turn | 30 | 110.01 | 0.1 | A | mA |  |  |  | 29.551 | 1.93 | 29.511 | 1.9297
| DCA 10 turn130 | DCA 10 turn | 30 | 110.01 | 1 | A | mA |  |  |  | 565.25 | 0.52854 | 565.25 | 0.52743
| DCA 10 turn1030 | DCA 10 turn | 30 | 110.01 | 10 | A | mA |  |  |  | 5772.3 | 0.055475 | 5772.3 | 0.054985
| DCA 10 turn10030 | DCA 10 turn | 30 | 110.01 | 100 | A | mA |  |  |  | 57735 | 0.0056968 | 57735 | 0.005501
| DCA 10 turn0.01110 | DCA 10 turn | 110 | 205 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
| DCA 10 turn0.1110 | DCA 10 turn | 110 | 205 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
| DCA 10 turn1110 | DCA 10 turn | 110 | 205 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
| DCA 10 turn10110 | DCA 10 turn | 110 | 205 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
| DCA 10 turn100110 | DCA 10 turn | 110 | 205 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
| DCA 10 turn0.01205 | DCA 10 turn | 205 | 206 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
| DCA 10 turn0.1205 | DCA 10 turn | 205 | 206 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
| DCA 10 turn1205 | DCA 10 turn | 205 | 206 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
| DCA 10 turn10205 | DCA 10 turn | 205 | 206 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
| DCA 10 turn100205 | DCA 10 turn | 205 | 206 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
| DCA 20 turn0.00120 | DCA 20 turn | 20 | 60.01 | 0.001 | A | mA |  |  |  | 3.9 | 2.1 | 0.23928 | 2.0473
| DCA 20 turn0.0120 | DCA 20 turn | 20 | 60.01 | 0.01 | A | mA |  |  |  | 3.9 | 2.1 | 0.76998 | 2.0407
| DCA 20 turn0.120 | DCA 20 turn | 20 | 60.01 | 0.1 | A | mA |  |  |  | 38.447 | 1.626 | 38.404 | 1.6257
| DCA 20 turn120 | DCA 20 turn | 20 | 60.01 | 1 | A | mA |  |  |  | 573.08 | 0.28848 | 573.07 | 0.28764
| DCA 20 turn1020 | DCA 20 turn | 20 | 60.01 | 10 | A | mA |  |  |  | 5773.1 | 0.029468 | 5773.1 | 0.02912
| DCA 20 turn10020 | DCA 20 turn | 20 | 60.01 | 100 | A | mA |  |  |  | 57735 | 0.0030514 | 57735 | 0.0029124
| DCA 20 turn0.00160 | DCA 20 turn | 60 | 220.01 | 0.001 | A | mA |  |  |  | 19 | 2.2 | 4.5221 | 2.114
| DCA 20 turn0.0160 | DCA 20 turn | 60 | 220.01 | 0.01 | A | mA |  |  |  | 19 | 2.2 | 19.726 | 2.0608
| DCA 20 turn0.160 | DCA 20 turn | 60 | 220.01 | 0.1 | A | mA |  |  |  | 535.04 | 0.95226 | 535.04 | 0.95071
| DCA 20 turn160 | DCA 20 turn | 60 | 220.01 | 1 | A | mA |  |  |  | 5768.4 | 0.11068 | 5768.5 | 0.10982
| DCA 20 turn1060 | DCA 20 turn | 60 | 220.01 | 10 | A | mA |  |  |  | 57735 | 0.011347 | 57735 | 0.011002
| DCA 20 turn10060 | DCA 20 turn | 60 | 220.01 | 100 | A | mA |  |  |  | 577360 | 0.0012382 | 577350 | 0.0011002
| DCA 20 turn0.01220 | DCA 20 turn | 220 | 410 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
| DCA 20 turn0.1220 | DCA 20 turn | 220 | 410 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
| DCA 20 turn1220 | DCA 20 turn | 220 | 410 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
| DCA 20 turn10220 | DCA 20 turn | 220 | 410 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
| DCA 20 turn100220 | DCA 20 turn | 220 | 410 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
| DCA 20 turn0.01410 | DCA 20 turn | 410 | 411 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
| DCA 20 turn0.1410 | DCA 20 turn | 410 | 411 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
| DCA 20 turn1410 | DCA 20 turn | 410 | 411 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
| DCA 20 turn10410 | DCA 20 turn | 410 | 411 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
| DCA 20 turn100410 | DCA 20 turn | 410 | 411 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
| DCA 50 turn0.00155 | DCA 50 turn | 55 | 150.01 | 0.001 | A | mA |  |  |  | 7.7 | 2 | 0.58265 | 2.0474
| DCA 50 turn0.0155 | DCA 50 turn | 55 | 150.01 | 0.01 | A | mA |  |  |  | 7.7 | 2.0006 | 0.78384 | 2.0464
| DCA 50 turn0.155 | DCA 50 turn | 55 | 150.01 | 0.1 | A | mA |  |  |  | 19.451 | 1.9578 | 19.379 | 1.9579
| DCA 50 turn155 | DCA 50 turn | 55 | 150.01 | 1 | A | mA |  |  |  | 550.23 | 0.69408 | 550.2 | 0.69349
| DCA 50 turn1055 | DCA 50 turn | 55 | 150.01 | 10 | A | mA |  |  |  | 5770.6 | 0.074862 | 5770.6 | 0.074565
| DCA 50 turn10055 | DCA 50 turn | 55 | 150.01 | 100 | A | mA |  |  |  | 57735 | 0.0075815 | 57732 | 0.0074625
| DCA 50 turn0.01150 | DCA 50 turn | 150 | 550.01 | 0.01 | A | mA |  |  |  | 19 | 2.2 | -17.053 | 2.1654
| DCA 50 turn0.1150 | DCA 50 turn | 150 | 550.01 | 0.1 | A | mA |  |  |  | 19 | 2.2 | -10.261 | 2.1556
| DCA 50 turn1150 | DCA 50 turn | 150 | 550.01 | 1 | A | mA |  |  |  | 409.49 | 1.6357 | 409.01 | 1.6349
| DCA 50 turn10150 | DCA 50 turn | 150 | 550.01 | 10 | A | mA |  |  |  | 5740.6 | 0.27618 | 5740.5 | 0.27485
| DCA 50 turn100150 | DCA 50 turn | 150 | 550.01 | 100 | A | mA |  |  |  | 57732 | 0.028329 | 57732 | 0.027783
| DCA 50 turn0.01550 | DCA 50 turn | 550 | 1025 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
| DCA 50 turn0.1550 | DCA 50 turn | 550 | 1025 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
| DCA 50 turn1550 | DCA 50 turn | 550 | 1025 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
| DCA 50 turn10550 | DCA 50 turn | 550 | 1025 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
| DCA 50 turn100550 | DCA 50 turn | 550 | 1025 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
| DCA 50 turn0.011025 | DCA 50 turn | 1025 | 1026 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
| DCA 50 turn0.11025 | DCA 50 turn | 1025 | 1026 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
| DCA 50 turn11025 | DCA 50 turn | 1025 | 1026 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
| DCA 50 turn101025 | DCA 50 turn | 1025 | 1026 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
| DCA 50 turn1001025 | DCA 50 turn | 1025 | 1026 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
| Resistance0.0000010 | Resistance | 0 | 0.011001 | 1e-06 | kOhms | mΩ |  |  |  | 12 | 140 | 11.576 | 138.3
| Resistance0.000010 | Resistance | 0 | 0.011001 | 1e-05 | kOhms | mΩ |  |  |  | 12.939 | 125.81 | 12.935 | 125.25
| Resistance0.00010 | Resistance | 0 | 0.011001 | 0.0001 | kOhms | mΩ |  |  |  | 59.144 | 29.356 | 59.141 | 28.788
| Resistance0.0010 | Resistance | 0 | 0.011001 | 0.001 | kOhms | mΩ |  |  |  | 580.12 | 3.176 | 580.12 | 2.9427
| Resistance0.010 | Resistance | 0 | 0.011001 | 0.01 | kOhms | mΩ |  |  |  | 5800 | 0.38767 | 5800 | 0.29434
| Resistance0.10 | Resistance | 0 | 0.011001 | 0.1 | kOhms | mΩ |  |  |  | 58000 | 0.066765 | 58000 | 0.029434
| Resistance10 | Resistance | 0 | 0.011001 | 1 | kOhms | mΩ |  |  |  | 580000 | 0.017876 | 580000 | 0.0029434
| Resistance100 | Resistance | 0 | 0.011001 | 10 | kOhms | mΩ |  |  |  | 5800000 | 0.0062674 | 5800000 | 0.00029435
| Resistance1000 | Resistance | 0 | 0.011001 | 100 | kOhms | mΩ |  |  |  | 58000000 | 0.0024177 | 58000000 | 2.9392e-05
| Resistance10000 | Resistance | 0 | 0.011001 | 1000 | kOhms | mΩ |  |  |  | 580000000 | 0.00095149 | 580000000 | 0
| Resistance100000 | Resistance | 0 | 0.011001 | 10000 | kOhms | mΩ |  |  |  | 5800000000 | 0 | 5800000000 | 0
| Resistance0.0000010.011 | Resistance | 0.011 | 0.033001 | 1e-06 | kOhms | mΩ |  |  |  | 18 | 140 | 17.339 | 138.85
| Resistance0.000010.011 | Resistance | 0.011 | 0.033001 | 1e-05 | kOhms | mΩ |  |  |  | 18.263 | 133.72 | 18.258 | 133.57
| Resistance0.00010.011 | Resistance | 0.011 | 0.033001 | 0.0001 | kOhms | mΩ |  |  |  | 60.487 | 46.304 | 60.481 | 46.04
| Resistance0.0010.011 | Resistance | 0.011 | 0.033001 | 0.001 | kOhms | mΩ |  |  |  | 580.26 | 4.9938 | 580.25 | 4.8783
| Resistance0.0000010.033 | Resistance | 0.033 | 0.11001 | 1e-06 | kOhms | mΩ |  |  |  | 18 | 100 | 17.342 | 104.06
| Resistance0.000010.033 | Resistance | 0.033 | 0.11001 | 1e-05 | kOhms | mΩ |  |  |  | 18.224 | 101.51 | 18.221 | 101.28
| Resistance0.00010.033 | Resistance | 0.033 | 0.11001 | 0.0001 | kOhms | mΩ |  |  |  | 60.259 | 41.284 | 60.259 | 40.82
| Resistance0.0010.033 | Resistance | 0.033 | 0.11001 | 0.001 | kOhms | mΩ |  |  |  | 580.23 | 4.6533 | 580.23 | 4.4422
| Resistance0.0000010.11 | Resistance | 0.11 | 0.33001 | 1e-06 | kOhms | mΩ |  |  |  | 25 | 300 | 24.812 | 296.87
| Resistance0.000010.11 | Resistance | 0.11 | 0.33001 | 1e-05 | kOhms | mΩ |  |  |  | 25.358 | 294.79 | 25.35 | 294.41
| Resistance0.00010.11 | Resistance | 0.11 | 0.33001 | 0.0001 | kOhms | mΩ |  |  |  | 61.466 | 184.35 | 61.46 | 183.05
| Resistance0.0010.11 | Resistance | 0.11 | 0.33001 | 0.001 | kOhms | mΩ |  |  |  | 580.25 | 24.237 | 580.25 | 23.449
| Resistance0.000010.33 | Resistance | 0.33 | 1.1 | 1e-05 | kOhms | mΩ |  |  |  | 28 | 100 | 23.818 | 103.66
| Resistance0.00010.33 | Resistance | 0.33 | 1.1 | 0.0001 | kOhms | mΩ |  |  |  | 52.968 | 87.892 | 52.867 | 87.789
| Resistance0.0010.33 | Resistance | 0.33 | 1.1 | 0.001 | kOhms | mΩ |  |  |  | 577.2 | 17.401 | 577.18 | 17.221
| Resistance0.010.33 | Resistance | 0.33 | 1.1 | 0.01 | kOhms | mΩ |  |  |  | 5799.7 | 1.8248 | 5799.7 | 1.75
| Resistance0.000011.1 | Resistance | 1.1 | 3.3001 | 1e-05 | kOhms | Ω |  |  |  | 0.25 | 0.1 | 0.23096 | 0.10411
| Resistance0.00011.1 | Resistance | 1.1 | 3.3001 | 0.0001 | kOhms | Ω |  |  |  | 0.25 | 0.10329 | 0.23666 | 0.10325
| Resistance0.0011.1 | Resistance | 1.1 | 3.3001 | 0.001 | kOhms | Ω |  |  |  | 0.60487 | 0.064429 | 0.60441 | 0.064221
| Resistance0.011.1 | Resistance | 1.1 | 3.3001 | 0.01 | kOhms | Ω |  |  |  | 5.8014 | 0.0083761 | 5.8012 | 0.0082294
| Resistance0.00013.3 | Resistance | 3.3 | 11.001 | 0.0001 | kOhms | Ω |  |  |  | 0.17 | 0.1 | 0.12474 | 0.10353
| Resistance0.0013.3 | Resistance | 3.3 | 11.001 | 0.001 | kOhms | Ω |  |  |  | 0.46536 | 0.084151 | 0.46464 | 0.084044
| Resistance0.013.3 | Resistance | 3.3 | 11.001 | 0.01 | kOhms | Ω |  |  |  | 5.7682 | 0.015398 | 5.7682 | 0.015238
| Resistance0.13.3 | Resistance | 3.3 | 11.001 | 0.1 | kOhms | Ω |  |  |  | 57.997 | 0.0016098 | 57.997 | 0.0015438
| Resistance0.000111 | Resistance | 11 | 33.001 | 0.0001 | kOhms | Ω |  |  |  | 1.3 | 0.1 | 1.1561 | 0.10412
| Resistance0.00111 | Resistance | 11 | 33.001 | 0.001 | kOhms | Ω |  |  |  | 1.3 | 0.10253 | 1.2442 | 0.10252
| Resistance0.0111 | Resistance | 11 | 33.001 | 0.01 | kOhms | Ω |  |  |  | 5.6614 | 0.052695 | 5.6605 | 0.052621
| Resistance0.111 | Resistance | 11 | 33.001 | 0.1 | kOhms | Ω |  |  |  | 57.978 | 0.0062175 | 57.978 | 0.0061738
| Resistance0.00133 | Resistance | 33 | 110.01 | 0.001 | kOhms | Ω |  |  |  | 1.3 | 0.13 | 1.2347 | 0.12665
| Resistance0.0133 | Resistance | 33 | 110.01 | 0.01 | kOhms | Ω |  |  |  | 4.353 | 0.10805 | 4.3501 | 0.10797
| Resistance0.133 | Resistance | 33 | 110.01 | 0.1 | kOhms | Ω |  |  |  | 57.521 | 0.022173 | 57.522 | 0.022051
| Resistance133 | Resistance | 33 | 110.01 | 1 | kOhms | Ω |  |  |  | 579.95 | 0.0022971 | 579.95 | 0.0022465
| Resistance0.001110 | Resistance | 110 | 330.01 | 0.001 | kOhms | Ω |  |  |  | 12 | 0.14 | 11.565 | 0.1393
| Resistance0.01110 | Resistance | 110 | 330.01 | 0.01 | kOhms | Ω |  |  |  | 12.289 | 0.13784 | 12.284 | 0.13782
| Resistance0.1110 | Resistance | 110 | 330.01 | 0.1 | kOhms | Ω |  |  |  | 55.026 | 0.080748 | 55.024 | 0.080667
| Resistance1110 | Resistance | 110 | 330.01 | 1 | kOhms | Ω |  |  |  | 579.51 | 0.010142 | 579.51 | 0.010091
| Resistance0.01330 | Resistance | 330 | 1100.01 | 0.01 | kOhms | Ω |  |  |  | 16 | 0.17 | 12.341 | 0.17274
| Resistance0.1330 | Resistance | 330 | 1100.01 | 0.1 | kOhms | Ω |  |  |  | 38.744 | 0.15613 | 38.702 | 0.15611
| Resistance1330 | Resistance | 330 | 1100.01 | 1 | kOhms | Ω |  |  |  | 571.18 | 0.039241 | 571.18 | 0.039165
| Resistance10330 | Resistance | 330 | 1100.01 | 10 | kOhms | Ω |  |  |  | 5799.1 | 0.0040776 | 5799.1 | 0.0040446
| Resistance0.011100 | Resistance | 1100 | 3301 | 0.01 | kOhms | kΩ |  |  |  | 0.19 | 0.00017 | 0.17339 | 0.00017318
| Resistance0.11100 | Resistance | 1100 | 3301 | 0.1 | kOhms | kΩ |  |  |  | 0.19 | 0.00017215 | 0.1791 | 0.00017213
| Resistance11100 | Resistance | 1100 | 3301 | 1 | kOhms | kΩ |  |  |  | 0.55516 | 0.00011804 | 0.55501 | 0.00011789
| Resistance101100 | Resistance | 1100 | 3301 | 10 | kOhms | kΩ |  |  |  | 5.7933 | 1.659e-05 | 5.7933 | 1.6469e-05
| Resistance0.13300 | Resistance | 3300 | 11001 | 0.1 | kOhms | kΩ |  |  |  | 0.33 | 0.00069 | 0.29033 | 0.00069272
| Resistance13300 | Resistance | 3300 | 11001 | 1 | kOhms | kΩ |  |  |  | 0.37289 | 0.00068715 | 0.37252 | 0.00068716
| Resistance103300 | Resistance | 3300 | 11001 | 10 | kOhms | kΩ |  |  |  | 4.8626 | 0.00044972 | 4.8623 | 0.00044966
| Resistance1003300 | Resistance | 3300 | 11001 | 100 | kOhms | kΩ |  |  |  | 57.852 | 6.2356e-05 | 57.852 | 6.2304e-05
| Resistance0.111000 | Resistance | 11000 | 33001 | 0.1 | kOhms | kΩ |  |  |  | 2.9 | 0.0012 | 2.8873 | 0.0011548
| Resistance111000 | Resistance | 11000 | 33001 | 1 | kOhms | kΩ |  |  |  | 2.9004 | 0.0011545 | 2.9 | 0.0011545
| Resistance1011000 | Resistance | 11000 | 33001 | 10 | kOhms | kΩ |  |  |  | 4.2484 | 0.0011259 | 4.2478 | 0.0011259
| Resistance10011000 | Resistance | 11000 | 33001 | 100 | kOhms | kΩ |  |  |  | 54.575 | 0.00049852 | 54.575 | 0.00049846
| Resistance133000 | Resistance | 33000 | 110010 | 1 | kOhms | kΩ |  |  |  | 3.5 | 0.0058 | 3.4663 | 0.0057735
| Resistance1033000 | Resistance | 33000 | 110010 | 10 | kOhms | kΩ |  |  |  | 3.5803 | 0.0057727 | 3.5792 | 0.0057727
| Resistance10033000 | Resistance | 33000 | 110010 | 100 | kOhms | kΩ |  |  |  | 14.463 | 0.0056975 | 14.465 | 0.0056974
| Resistance100033000 | Resistance | 33000 | 110010 | 1000 | kOhms | kΩ |  |  |  | 503.99 | 0.0032606 | 503.99 | 0.0032605
| Resistance1110000 | Resistance | 110000 | 330010 | 1 | kOhms | kΩ |  |  |  | 120 | 0.0058 | 115.47 | 0.0057735
| Resistance10110000 | Resistance | 110000 | 330010 | 10 | kOhms | kΩ |  |  |  | 120 | 0.0057734 | 115.52 | 0.0057734
| Resistance100110000 | Resistance | 110000 | 330010 | 100 | kOhms | kΩ |  |  |  | 120 | 0.0057671 | 118.42 | 0.0057671
| Resistance1000110000 | Resistance | 110000 | 330010 | 1000 | kOhms | kΩ |  |  |  | 371.66 | 0.0052445 | 371.67 | 0.0052444
| Resistance10330000 | Resistance | 330000 | 1100010 | 10 | kOhms | kΩ |  |  |  | 930 | 0.017 | 577.19 | 0.017321
| Resistance100330000 | Resistance | 330000 | 1100010 | 100 | kOhms | kΩ |  |  |  | 930 | 0.01732 | 578 | 0.01732
| Resistance1000330000 | Resistance | 330000 | 1100010 | 1000 | kOhms | kΩ |  |  |  | 930 | 0.017297 | 611.79 | 0.017297
| Resistance10000330000 | Resistance | 330000 | 1100010 | 10000 | kOhms | kΩ |  |  |  | 3454.2 | 0.015468 | 3454.1 | 0.015468
| Resistance101100000 | Resistance | 1100000 | 1101000 | 10 | kOhms | kΩ |  |  |  | 930 | 0.017 | 577.19 | 0.017321
| Resistance1001100000 | Resistance | 1100000 | 1101000 | 100 | kOhms | kΩ |  |  |  | 930 | 0.01732 | 578 | 0.01732
| Resistance10001100000 | Resistance | 1100000 | 1101000 | 1000 | kOhms | kΩ |  |  |  | 930 | 0.017297 | 611.79 | 0.017297
| Resistance100001100000 | Resistance | 1100000 | 1101000 | 10000 | kOhms | kΩ |  |  |  | 3454.2 | 0.015468 | 3454.1 | 0.015468
| ACV0.00000100.01 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 0.01 | 0.04501 | Hz | 25 | 1700 | 23.181 | 1732
| ACV0.0000100.01 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 0.01 | 0.04501 | Hz | 25 | 1718.2 | 23.854 | 1717.9
| ACV0.000100.01 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 0.01 | 0.04501 | Hz | 62.01 | 1125.5 | 61.999 | 1123.8
| ACV0.00100.01 | ACV | 0 | 0.03301 | 0.001 | V | µV | 0.01 | 0.04501 | Hz | 580.38 | 157.54 | 580.38 | 156.33
| ACV0.0100.01 | ACV | 0 | 0.03301 | 0.01 | V | µV | 0.01 | 0.04501 | Hz | 5800 | 16.202 | 5800 | 15.714
| ACV0.100.01 | ACV | 0 | 0.03301 | 0.1 | V | µV | 0.01 | 0.04501 | Hz | 58000 | 1.7665 | 58000 | 1.5715
| ACV100.01 | ACV | 0 | 0.03301 | 1 | V | µV | 0.01 | 0.04501 | Hz | 580000 | 0.23514 | 580000 | 0.15715
| ACV1000.01 | ACV | 0 | 0.03301 | 10 | V | µV | 0.01 | 0.04501 | Hz | 5800000 | 0.046912 | 5800000 | 0.015715
| ACV0.00000100.045 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 1200 | 23.185 | 1155.4
| ACV0.0000100.045 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 1143.4 | 23.87 | 1142.8
| ACV0.000100.045 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 0.045 | 10 | Hz | 62.237 | 673.84 | 62.225 | 671.86
| ACV0.00100.045 | ACV | 0 | 0.03301 | 0.001 | V | µV | 0.045 | 10 | Hz | 580.43 | 86.247 | 580.43 | 85.036
| ACV0.0100.045 | ACV | 0 | 0.03301 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800.1 | 9.0181 | 5800.1 | 8.5307
| ACV0.100.045 | ACV | 0 | 0.03301 | 0.1 | V | µV | 0.045 | 10 | Hz | 58000 | 1.0481 | 58000 | 0.8531
| ACV100.045 | ACV | 0 | 0.03301 | 1 | V | µV | 0.045 | 10 | Hz | 580000 | 0.1633 | 580000 | 0.08531
| ACV1000.045 | ACV | 0 | 0.03301 | 10 | V | µV | 0.045 | 10 | Hz | 5800000 | 0.039728 | 5800000 | 0.008531
| ACV0.000001010 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 10 | 20.01 | kHz | 25 | 1700 | 23.181 | 1732
| ACV0.00001010 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 10 | 20.01 | kHz | 25 | 1718.2 | 23.854 | 1717.9
| ACV0.0001010 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 10 | 20.01 | kHz | 62.01 | 1125.5 | 61.999 | 1123.8
| ACV0.001010 | ACV | 0 | 0.03301 | 0.001 | V | µV | 10 | 20.01 | kHz | 580.38 | 157.54 | 580.38 | 156.33
| ACV0.01010 | ACV | 0 | 0.03301 | 0.01 | V | µV | 10 | 20.01 | kHz | 5800 | 16.202 | 5800 | 15.714
| ACV0.1010 | ACV | 0 | 0.03301 | 0.1 | V | µV | 10 | 20.01 | kHz | 58000 | 1.7665 | 58000 | 1.5715
| ACV1010 | ACV | 0 | 0.03301 | 1 | V | µV | 10 | 20.01 | kHz | 580000 | 0.23514 | 580000 | 0.15715
| ACV10010 | ACV | 0 | 0.03301 | 10 | V | µV | 10 | 20.01 | kHz | 5800000 | 0.046912 | 5800000 | 0.015715
| ACV0.000001020 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 20 | 50.01 | kHz | 24 | 2300 | 23.565 | 2300.9
| ACV0.00001020 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 20 | 50.01 | kHz | 24.224 | 2286.5 | 24.216 | 2286.3
| ACV0.0001020 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 20 | 50.01 | kHz | 61.904 | 1615.7 | 61.89 | 1614.4
| ACV0.001020 | ACV | 0 | 0.03301 | 0.001 | V | µV | 20 | 50.01 | kHz | 580.34 | 247.93 | 580.33 | 246.73
| ACV0.01020 | ACV | 0 | 0.03301 | 0.01 | V | µV | 20 | 50.01 | kHz | 5800 | 25.351 | 5800 | 24.864
| ACV0.1020 | ACV | 0 | 0.03301 | 0.1 | V | µV | 20 | 50.01 | kHz | 58000 | 2.6816 | 58000 | 2.4866
| ACV1020 | ACV | 0 | 0.03301 | 1 | V | µV | 20 | 50.01 | kHz | 580000 | 0.32665 | 580000 | 0.24866
| ACV10020 | ACV | 0 | 0.03301 | 10 | V | µV | 20 | 50.01 | kHz | 5800000 | 0.056063 | 5800000 | 0.024866
| ACV0.000001050 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 50 | 100 | kHz | 40 | 4000 | 38.937 | 4027
| ACV0.00001050 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 50 | 100.01 | kHz | 40 | 4017.5 | 39.332 | 4018
| ACV0.0001050 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 50 | 100.01 | kHz | 68.858 | 3411.4 | 68.766 | 3411.7
| ACV0.001050 | ACV | 0 | 0.03301 | 0.001 | V | µV | 50 | 100.01 | kHz | 580.9 | 730.69 | 580.86 | 729
| ACV0.01050 | ACV | 0 | 0.03301 | 0.01 | V | µV | 50 | 100.01 | kHz | 5800.1 | 75.275 | 5800.1 | 74.549
| ACV0.1050 | ACV | 0 | 0.03301 | 0.1 | V | µV | 50 | 100.01 | kHz | 58000 | 7.7474 | 58000 | 7.4566
| ACV1050 | ACV | 0 | 0.03301 | 1 | V | µV | 50 | 100.01 | kHz | 580000 | 0.86197 | 580000 | 0.74566
| ACV10050 | ACV | 0 | 0.03301 | 10 | V | µV | 50 | 100.01 | kHz | 5800000 | 0.12109 | 5800000 | 0.074566
| ACV0.0000010100 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 100 | 500 | kHz | 89 | 11000 | 72.86 | 11473
| ACV0.000010100 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 100 | 500 | kHz | 89 | 11468 | 73.063 | 11468
| ACV0.00010100 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 100 | 500 | kHz | 91.464 | 11024 | 91.322 | 11026
| ACV0.0010100 | ACV | 0 | 0.03301 | 0.001 | V | µV | 100 | 500 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
| ACV0.010100 | ACV | 0 | 0.03301 | 0.01 | V | µV | 100 | 500 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
| ACV0.10100 | ACV | 0 | 0.03301 | 0.1 | V | µV | 100 | 500 | kHz | 58000 | 53.659 | 58000 | 52.994
| ACV10100 | ACV | 0 | 0.03301 | 1 | V | µV | 100 | 500 | kHz | 580000 | 5.5657 | 580000 | 5.2995
| ACV0.0000010500 | ACV | 0 | 0.03301 | 1e-06 | V | µV | 500 | 501 | kHz | 89 | 11000 | 72.86 | 11473
| ACV0.000010500 | ACV | 0 | 0.03301 | 1e-05 | V | µV | 500 | 501 | kHz | 89 | 11468 | 73.063 | 11468
| ACV0.00010500 | ACV | 0 | 0.03301 | 0.0001 | V | µV | 500 | 501 | kHz | 91.464 | 11024 | 91.322 | 11026
| ACV0.0010500 | ACV | 0 | 0.03301 | 0.001 | V | µV | 500 | 501 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
| ACV0.010500 | ACV | 0 | 0.03301 | 0.01 | V | µV | 500 | 501 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
| ACV0.10500 | ACV | 0 | 0.03301 | 0.1 | V | µV | 500 | 501 | kHz | 58000 | 53.659 | 58000 | 52.994
| ACV10500 | ACV | 0 | 0.03301 | 1 | V | µV | 500 | 501 | kHz | 580000 | 5.5657 | 580000 | 5.2995
| ACV100100 | ACV | 0 | 0.03301 | 10 | V | µV | 100 | 500.01 | kHz | 5800000 | 0.63642 | 5800000 | 0.52995
| ACV0.0000010.0330.01 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 0.01 | 0.04501 | Hz | 33 | 550 | 23.236 | 577.95
| ACV0.000010.0330.01 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 0.01 | 0.04501 | Hz | 33 | 550.18 | 23.663 | 576.89
| ACV0.00010.0330.01 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 0.01 | 0.04501 | Hz | 55.234 | 504.47 | 55.135 | 504.68
| ACV0.0010.0330.01 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 0.01 | 0.04501 | Hz | 577.51 | 123.44 | 577.47 | 123.45
| ACV0.010.0330.01 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 0.01 | 0.04501 | Hz | 5799.8 | 12.763 | 577.47 | 12.763
| ACV0.0000010.0330.045 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 350 | 23.254 | 347.52
| ACV0.000010.0330.045 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 346.2 | 23.771 | 346.32
| ACV0.00010.0330.045 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 0.045 | 10 | Hz | 58.586 | 276.03 | 58.483 | 276.21
| ACV0.0010.0330.045 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 0.045 | 10 | Hz | 579.4 | 50.965 | 579.36 | 50.969
| ACV0.010.0330.045 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800 | 5.1718 | 5799.9 | 5.1718
| ACV0.0000010.03310 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 10 | 20.01 | kHz | 24 | 810 | 23.218 | 808.67
| ACV0.000010.03310 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 10 | 20.01 | kHz | 24 | 807.65 | 23.583 | 807.74
| ACV0.00010.03310 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 10 | 20.01 | kHz | 52.238 | 738.2 | 52.145 | 738.41
| ACV0.0010.03310 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 10 | 20.01 | kHz | 574.82 | 223.39 | 574.77 | 223.41
| ACV0.010.03310 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 10 | 20.01 | kHz | 5799.5 | 23.686 | 5799.4 | 23.686
| ACV0.0000010.03320 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 20 | 50.01 | kHz | 47 | 1200 | 46.389 | 1155.4
| ACV0.000010.03320 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 20 | 50.01 | kHz | 47 | 1154.8 | 46.618 | 1154.8
| ACV0.00010.03320 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 20 | 50.01 | kHz | 66.007 | 1107.9 | 65.939 | 1108
| ACV0.0010.03320 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 20 | 50.01 | kHz | 571.22 | 452.98 | 571.18 | 452.85
| ACV0.010.03320 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 20 | 50.01 | kHz | 5799 | 51.02 | 5798.9 | 50.942
| ACV0.0000010.03350 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 50 | 100.01 | kHz | 200 | 2700 | 196.5 | 2656.4
| ACV0.000010.03350 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 50 | 100.01 | kHz | 200 | 2656.2 | 196.56 | 2656.3
| ACV0.00010.03350 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 50 | 100.01 | kHz | 202.89 | 2641.9 | 202.83 | 2642
| ACV0.0010.03350 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 50 | 100.01 | kHz | 582.19 | 1932.6 | 582.1 | 1932.5
| ACV0.010.03350 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 50 | 100.01 | kHz | 5796.8 | 308.19 | 5796.8 | 308.02
| ACV0.0000010.033100 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 100 | 500 | kHz | 390 | 5800 | 381.75 | 5779.4
| ACV0.000010.033100 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 100 | 500 | kHz | 390 | 5779.3 | 381.8 | 5779.3
| ACV0.00010.033100 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 100 | 500 | kHz | 390 | 5771.9 | 384.93 | 5772
| ACV0.0010.033100 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 100 | 500 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
| ACV0.010.033100 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 100 | 500 | kHz | 5783 | 1371.2 | 5783 | 1370.9
| ACV0.0000010.033500 | ACV | 0.033 | 0.33001 | 1e-06 | V | µV | 500 | 501 | kHz | 390 | 5800 | 381.75 | 5779.4
| ACV0.000010.033500 | ACV | 0.033 | 0.33001 | 1e-05 | V | µV | 500 | 501 | kHz | 390 | 5779.3 | 381.8 | 5779.3
| ACV0.00010.033500 | ACV | 0.033 | 0.33001 | 0.0001 | V | µV | 500 | 501 | kHz | 390 | 5771.9 | 384.93 | 5772
| ACV0.0010.033500 | ACV | 0.033 | 0.33001 | 0.001 | V | µV | 500 | 501 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
| ACV0.010.033500 | ACV | 0.033 | 0.33001 | 0.01 | V | µV | 500 | 501 | kHz | 5783 | 1371.2 | 5783 | 1370.9
| ACV0.000010.330.01 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 0.01 | 0.04501 | Hz | 0.11 | 0.56 | 0.10391 | 0.56221
| ACV0.00010.330.01 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 0.01 | 0.04501 | Hz | 0.11 | 0.56033 | 0.10921 | 0.56037
| ACV0.0010.330.01 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 0.48274 | 0.46448 | 0.4824 | 0.4645
| ACV0.010.330.01 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 5.7719 | 0.10338 | 5.7717 | 0.10332
| ACV0.10.330.01 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 57.997 | 0.010643 | 57.997 | 0.010612
| ACV0.000010.330.045 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 0.045 | 10 | Hz | 0.07 | 0.35 | 0.069477 | 0.34756
| ACV0.00010.330.045 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.079443 | 0.34499 | 0.079136 | 0.34505
| ACV0.0010.330.045 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.52675 | 0.24887 | 0.5264 | 0.24885
| ACV0.010.330.045 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.7894 | 0.041576 | 5.7892 | 0.041503
| ACV0.10.330.045 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.999 | 0.0042263 | 57.999 | 0.004196
| ACV0.000010.3310 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 10 | 20.01 | kHz | 0.07 | 0.81 | 0.069435 | 0.80878
| ACV0.00010.3310 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 10 | 20.01 | kHz | 0.075008 | 0.80729 | 0.074829 | 0.80733
| ACV0.0010.3310 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 10 | 20.01 | kHz | 0.43427 | 0.71669 | 0.43391 | 0.71674
| ACV0.010.3310 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 10 | 20.01 | kHz | 5.7428 | 0.2035 | 5.7426 | 0.20344
| ACV0.10.3310 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 10 | 20.01 | kHz | 57.994 | 0.021456 | 57.994 | 0.021426
| ACV0.000010.3320 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 20 | 50.01 | kHz | 0.07 | 1.2 | 0.069506 | 1.1559
| ACV0.00010.3320 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 20 | 50.01 | kHz | 0.073711 | 1.1548 | 0.073371 | 1.1549
| ACV0.0010.3320 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 20 | 50.01 | kHz | 0.38065 | 1.0748 | 0.37987 | 1.0749
| ACV0.010.3320 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 20 | 50.01 | kHz | 5.6886 | 0.39166 | 5.6883 | 0.39154
| ACV0.10.3320 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 20 | 50.01 | kHz | 57.988 | 0.043219 | 57.988 | 0.043148
| ACV0.000010.3350 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 50 | 100.01 | kHz | 0.24 | 2.7 | 0.23108 | 2.6571
| ACV0.00010.3350 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 50 | 100.01 | kHz | 0.24 | 2.6566 | 0.23256 | 2.6567
| ACV0.0010.3350 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 50 | 100.01 | kHz | 0.38846 | 2.6152 | 0.38733 | 2.6154
| ACV0.010.3350 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 50 | 100.01 | kHz | 5.3723 | 1.6168 | 5.3715 | 1.6167
| ACV0.10.3350 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 50 | 100.01 | kHz | 57.935 | 0.23029 | 57.935 | 0.23012
| ACV0.000010.33100 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 100 | 500 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
| ACV0.00010.33100 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 100 | 500 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
| ACV0.0010.33100 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 100 | 500 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
| ACV0.010.33100 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 100 | 500 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
| ACV0.10.33100 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 100 | 500 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
| ACV0.000010.33500 | ACV | 0.33 | 3.3001 | 1e-05 | V | mV | 500 | 501 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
| ACV0.00010.33500 | ACV | 0.33 | 3.3001 | 0.0001 | V | mV | 500 | 501 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
| ACV0.0010.33500 | ACV | 0.33 | 3.3001 | 0.001 | V | mV | 500 | 501 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
| ACV0.010.33500 | ACV | 0.33 | 3.3001 | 0.01 | V | mV | 500 | 501 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
| ACV0.10.33500 | ACV | 0.33 | 3.3001 | 0.1 | V | mV | 500 | 501 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
| ACV0.00013.30.01 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 0.01 | 0.04501 | Hz | 0.93 | 0.58 | 0.92292 | 0.5783
| ACV0.0013.30.01 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 0.98726 | 0.57661 | 0.98668 | 0.57662
| ACV0.013.30.01 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 4.8574 | 0.48407 | 4.8565 | 0.48406
| ACV0.13.30.01 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 57.704 | 0.11064 | 57.704 | 0.11059
| ACV13.30.01 | ACV | 3.3 | 33.001 | 1 | V | mV | 0.01 | 0.04501 | Hz | 579.97 | 0.011401 | 579.97 | 0.011382
| ACV0.00013.30.045 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.7 | 0.35 | 0.69218 | 0.34795
| ACV0.0013.30.045 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.78957 | 0.34543 | 0.78891 | 0.34543
| ACV0.013.30.045 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.2633 | 0.24921 | 5.2624 | 0.24918
| ACV0.13.30.045 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.892 | 0.04162 | 57.892 | 0.041575
| ACV13.30.045 | ACV | 3.3 | 33.001 | 1 | V | mV | 0.045 | 10 | Hz | 579.99 | 0.0042219 | 579.99 | 0.0042034
| ACV0.00013.310 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 10 | 20.01 | kHz | 0.7 | 0.81 | 0.69304 | 0.80896
| ACV0.0013.310 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 10 | 20.01 | kHz | 0.74744 | 0.8075 | 0.74699 | 0.80751
| ACV0.013.310 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 10 | 20.01 | kHz | 4.3392 | 0.7169 | 4.3383 | 0.7169
| ACV0.13.310 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 10 | 20.01 | kHz | 57.426 | 0.20355 | 57.426 | 0.20351
| ACV13.310 | ACV | 3.3 | 33.001 | 1 | V | mV | 10 | 20.01 | kHz | 579.94 | 0.021452 | 579.94 | 0.021433
| ACV0.00013.320 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 20 | 50.01 | kHz | 0.7 | 1.2 | 0.69122 | 1.1565
| ACV0.0013.320 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 20 | 50.01 | kHz | 0.73466 | 1.1553 | 0.73165 | 1.1554
| ACV0.013.320 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 20 | 50.01 | kHz | 3.8002 | 1.0754 | 3.7971 | 1.0754
| ACV0.13.320 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 20 | 50.01 | kHz | 56.883 | 0.3919 | 56.882 | 0.39182
| ACV13.320 | ACV | 3.3 | 33.001 | 1 | V | mV | 20 | 50.01 | kHz | 579.88 | 0.043223 | 579.88 | 0.043182
| ACV0.00013.350 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 50 | 100.01 | kHz | 2.3 | 2.7 | 2.3079 | 2.6577
| ACV0.0013.350 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 50 | 100.01 | kHz | 2.3265 | 2.6572 | 2.3234 | 2.6573
| ACV0.013.350 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 50 | 100.01 | kHz | 3.8797 | 2.6158 | 3.8711 | 2.616
| ACV0.13.350 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 50 | 100.01 | kHz | 53.719 | 1.6172 | 53.713 | 1.6172
| ACV13.350 | ACV | 3.3 | 33.001 | 1 | V | mV | 50 | 100.01 | kHz | 579.35 | 0.23028 | 579.35 | 0.23021
| ACV0.00013.3100 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3364 | 2.6995
| ACV0.0013.3100 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 2.3501 | 2.699
| ACV0.013.3100 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 100 | 500 | kHz | 13 | 4.6 | 3.831 | 2.6581
| ACV0.13.3100 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 100 | 500 | kHz | 47.329 | 3.5597 | 53.608 | 1.6549
| ACV13.3100 | ACV | 3.3 | 33.001 | 1 | V | mV | 100 | 500 | kHz | 579.33 | 0.23723 | 579.33 | 0.23717
| ACV0.00013.3500 | ACV | 3.3 | 33.001 | 0.0001 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 2.3364 | 2.6995
| ACV0.0013.3500 | ACV | 3.3 | 33.001 | 0.001 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 2.3501 | 2.699
| ACV0.013.3500 | ACV | 3.3 | 33.001 | 0.01 | V | mV | 500 | 501 | kHz | 13 | 4.6 | 3.831 | 2.6581
| ACV0.13.3500 | ACV | 3.3 | 33.001 | 0.1 | V | mV | 500 | 501 | kHz | 47.329 | 3.5597 | 53.608 | 1.6549
| ACV13.3500 | ACV | 3.3 | 33.001 | 1 | V | mV | 500 | 501 | kHz | 579.33 | 0.23723 | 579.33 | 0.23717
| ACV0.001330.01 | ACV | 33 | 330.01 | 0.001 | V | mV | 0.01 | 0.04501 | Hz | 3.5 | 0.58 | 3.4705 | 0.57868
| ACV0.01330.01 | ACV | 33 | 330.01 | 0.01 | V | mV | 0.01 | 0.04501 | Hz | 4.2872 | 0.57647 | 4.2606 | 0.57652
| ACV0.1330.01 | ACV | 33 | 330.01 | 0.1 | V | mV | 0.01 | 0.04501 | Hz | 46.641 | 0.47359 | 46.602 | 0.47361
| ACV1330.01 | ACV | 33 | 330.01 | 1 | V | mV | 0.01 | 0.04501 | Hz | 576.98 | 0.10542 | 576.96 | 0.10533
| ACV10330.01 | ACV | 33 | 330.01 | 10 | V | mV | 0.01 | 0.04501 | Hz | 5799.7 | 0.010862 | 5799.7 | 0.010822
| ACV0.001330.045 | ACV | 33 | 330.01 | 0.001 | V | mV | 0.045 | 10 | Hz | 12 | 0.92 | 10.393 | 0.92459
| ACV0.01330.045 | ACV | 33 | 330.01 | 0.01 | V | mV | 0.045 | 10 | Hz | 12 | 0.92331 | 10.834 | 0.9234
| ACV0.1330.045 | ACV | 33 | 330.01 | 0.1 | V | mV | 0.045 | 10 | Hz | 43.309 | 0.841 | 43.21 | 0.84115
| ACV1330.045 | ACV | 33 | 330.01 | 1 | V | mV | 0.045 | 10 | Hz | 572.72 | 0.26556 | 572.68 | 0.26538
| ACV10330.045 | ACV | 33 | 330.01 | 10 | V | mV | 0.045 | 10 | Hz | 5799.2 | 0.028474 | 5799.2 | 0.028386
| ACV0.0013310 | ACV | 33 | 330.01 | 0.001 | V | mV | 10 | 20.01 | kHz | 24 | 1 | 10.388 | 1.04
| ACV0.013310 | ACV | 33 | 330.01 | 0.01 | V | mV | 10 | 20.01 | kHz | 24 | 1.0332 | 10.796 | 1.0389
| ACV0.13310 | ACV | 33 | 330.01 | 0.1 | V | mV | 10 | 20.01 | kHz | 41.651 | 0.95969 | 41.554 | 0.95985
| ACV13310 | ACV | 33 | 330.01 | 1 | V | mV | 10 | 20.01 | kHz | 570.92 | 0.32864 | 570.88 | 0.32848
| ACV103310 | ACV | 33 | 330.01 | 10 | V | mV | 10 | 20.01 | kHz | 5799 | 0.035762 | 5799 | 0.035674
| ACV0.0013320 | ACV | 33 | 330.01 | 0.001 | V | mV | 20 | 50.01 | kHz | 22 | 1.4 | 10.266 | 1.3912
| ACV0.013320 | ACV | 33 | 330.01 | 0.01 | V | mV | 20 | 50.01 | kHz | 22 | 1.4 | 10.577 | 1.3904
| ACV0.13320 | ACV | 33 | 330.01 | 0.1 | V | mV | 20 | 50.01 | kHz | 36.02 | 1.3575 | 37.167 | 1.3205
| ACV13320 | ACV | 33 | 330.01 | 1 | V | mV | 20 | 50.01 | kHz | 564.6 | 0.55034 | 564.56 | 0.55021
| ACV103320 | ACV | 33 | 330.01 | 10 | V | mV | 20 | 50.01 | kHz | 5798.2 | 0.063012 | 5798.2 | 0.062925
| ACV0.0013350 | ACV | 33 | 330.01 | 0.001 | V | mV | 50 | 100 | kHz | 92 | 2.8 | 91.937 | 2.7856
| ACV0.013350 | ACV | 33 | 330.01 | 0.01 | V | mV | 50 | 100 | kHz | 92.063 | 2.7853 | 92.033 | 2.7854
| ACV0.13350 | ACV | 33 | 330.01 | 0.1 | V | mV | 50 | 100 | kHz | 101.77 | 2.761 | 101.69 | 2.7611
| ACV13350 | ACV | 33 | 330.01 | 1 | V | mV | 50 | 100 | kHz | 546.63 | 1.8765 | 546.54 | 1.8763
| ACV103350 | ACV | 33 | 330.01 | 10 | V | mV | 50 | 100 | kHz | 5793.6 | 0.28499 | 5793.5 | 0.28476
| ACV0.00133100 | ACV | 33 | 330.01 | 0.001 | V | mV | 100 | 100.1 | kHz | 92 | 2.8 | 91.937 | 2.7856
| ACV0.0133100 | ACV | 33 | 330.01 | 0.01 | V | mV | 100 | 100.1 | kHz | 92.063 | 2.7853 | 92.033 | 2.7854
| ACV0.133100 | ACV | 33 | 330.01 | 0.1 | V | mV | 100 | 100.1 | kHz | 101.77 | 2.761 | 101.69 | 2.7611
| ACV133100 | ACV | 33 | 330.01 | 1 | V | mV | 100 | 100.1 | kHz | 546.63 | 1.8765 | 546.54 | 1.8763
| ACV1033100 | ACV | 33 | 330.01 | 10 | V | mV | 100 | 100.1 | kHz | 5793.6 | 0.28499 | 5793.5 | 0.28476
| ACV0.0013300.045 | ACV | 330 | 1020.01 | 0.001 | V | mV | 0.045 | 1 | Hz | 24 | 0.58 | 23.335 | 0.57834
| ACV0.013300.045 | ACV | 330 | 1020.01 | 0.01 | V | mV | 0.045 | 1 | Hz | 33.354 | 0.57134 | 33.321 | 0.57119
| ACV0.13300.045 | ACV | 330 | 1020.01 | 0.1 | V | mV | 0.045 | 1 | Hz | 510.18 | 0.32785 | 510.26 | 0.32724
| ACV13300.045 | ACV | 330 | 1020.01 | 1 | V | mV | 0.045 | 1 | Hz | 5790.3 | 0.041506 | 5790.4 | 0.041123
| ACV103300.045 | ACV | 330 | 1020.01 | 10 | V | mV | 0.045 | 1 | Hz | 57999 | 0.0042793 | 57999 | 0.004125
| ACV0.0013301 | ACV | 330 | 1020.01 | 0.001 | V | mV | 1 | 5.01 | kHz | 28 | 0.92 | 23.262 | 0.92437
| ACV0.013301 | ACV | 330 | 1020.01 | 0.01 | V | mV | 1 | 5.01 | kHz | 29.895 | 0.91967 | 29.873 | 0.91957
| ACV0.13301 | ACV | 330 | 1020.01 | 0.1 | V | mV | 1 | 5.01 | kHz | 446.23 | 0.66766 | 446.26 | 0.66723
| ACV13301 | ACV | 330 | 1020.01 | 1 | V | mV | 1 | 5.01 | kHz | 5775.4 | 0.10274 | 5775.5 | 0.10236
| ACV103301 | ACV | 330 | 1020.01 | 10 | V | mV | 1 | 5.01 | kHz | 57998 | 0.010468 | 57998 | 0.010314
| ACV0.0013305 | ACV | 330 | 1020.01 | 0.001 | V | mV | 5 | 10.01 | kHz | 64 | 1 | 23.23 | 1.0398
| ACV0.013305 | ACV | 330 | 1020.01 | 0.01 | V | mV | 5 | 10.01 | kHz | 64 | 1.0355 | 29.165 | 1.0355
| ACV0.13305 | ACV | 330 | 1020.01 | 0.1 | V | mV | 5 | 10.01 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
| ACV13305 | ACV | 330 | 1020.01 | 1 | V | mV | 5 | 10.01 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
| ACV103305 | ACV | 330 | 1020.01 | 10 | V | mV | 5 | 10.01 | kHz | 57997 | 0.013152 | 57997 | 0.012998
| ACV0.00133010 | ACV | 330 | 1020.01 | 0.001 | V | mV | 10 | 10.01 | kHz | 64 | 1 | 23.23 | 1.0398
| ACV0.0133010 | ACV | 330 | 1020.01 | 0.01 | V | mV | 10 | 10.01 | kHz | 64 | 1.0355 | 29.165 | 1.0355
| ACV0.133010 | ACV | 330 | 1020.01 | 0.1 | V | mV | 10 | 10.01 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
| ACV133010 | ACV | 330 | 1020.01 | 1 | V | mV | 10 | 10.01 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
| ACV1033010 | ACV | 330 | 1020.01 | 10 | V | mV | 10 | 10.01 | kHz | 57997 | 0.013152 | 57997 | 0.012998
| Frequency0.0000010 | Frequency | 0 | 0.12 | 1e-06 | kHz | Hz |  |  |  | 0.006 | 0.013 | 0.0012889 | 0.028049
| Frequency0.000010 | Frequency | 0 | 0.12 | 1e-05 | kHz | Hz |  |  |  | 0.0060576 | 0.015807 | 0.005881 | 0.012606
| Frequency0.00010 | Frequency | 0 | 0.12 | 0.0001 | kHz | Hz |  |  |  | 0.057819 | 0.0033129 | 0.057746 | 0.0014481
| Frequency0.0010 | Frequency | 0 | 0.12 | 0.001 | kHz | Hz |  |  |  | 0.57739 | 0.00089636 | 0.57736 | 0.00014506
| Frequency0.010 | Frequency | 0 | 0.12 | 0.01 | kHz | Hz |  |  |  | 5.7736 | 0.00031508 | 5.7736 | 1.4506e-05
| Frequency0.10 | Frequency | 0 | 0.12 | 0.1 | kHz | Hz |  |  |  | 57.736 | 0.00012168 | 57.736 | 1.4506e-06
| Frequency10 | Frequency | 0 | 0.12 | 1 | kHz | Hz |  |  |  | 577.36 | 4.8236e-05 | 577.36 | 1.4506e-07
| Frequency0.000010.12 | Frequency | 0.12 | 1.2 | 1e-05 | kHz | Hz |  |  |  | 0.058 | 0.0095 | 0.0041865 | 0.026727
| Frequency0.00010.12 | Frequency | 0.12 | 1.2 | 0.0001 | kHz | Hz |  |  |  | 0.058519 | 0.010244 | 0.056808 | 0.0092702
| Frequency0.0010.12 | Frequency | 0.12 | 1.2 | 0.001 | kHz | Hz |  |  |  | 0.57795 | 0.0015973 | 0.57725 | 0.0010094
| Frequency0.010.12 | Frequency | 0.12 | 1.2 | 0.01 | kHz | Hz |  |  |  | 5.7739 | 0.00033757 | 5.7735 | 0.00010104
| Frequency0.00011.2 | Frequency | 1.2 | 12.001 | 0.0001 | kHz | Hz |  |  |  | 0.58 | 0.0091 | 0.036333 | 0.026333
| Frequency0.00011.2 | Frequency | 1.2 | 12.001 | 0.0001 | kHz | Hz |  |  |  | 0.58497 | 0.0098286 | 0.56786 | 0.0088368
| Frequency0.0011.2 | Frequency | 1.2 | 12.001 | 0.001 | kHz | Hz |  |  |  | 5.7795 | 0.0015456 | 5.7725 | 0.00095754
| Frequency0.011.2 | Frequency | 1.2 | 12.001 | 0.01 | kHz | Hz |  |  |  | 57.739 | 0.00033237 | 57.735 | 9.584e-05
| Frequency0.00112 | Frequency | 12 | 120.01 | 0.001 | kHz | Hz |  |  |  | 19 | 0.028 | 17.33 | 0.028854
| Frequency0.0112 | Frequency | 12 | 120.01 | 0.01 | kHz | Hz |  |  |  | 19 | 0.028 | 18.255 | 0.027641
| Frequency0.112 | Frequency | 12 | 120.01 | 0.1 | kHz | Hz |  |  |  | 60.592 | 0.011287 | 60.269 | 0.0091178
| Frequency112 | Frequency | 12 | 120.01 | 1 | kHz | Hz |  |  |  | 577.75 | 0.0019059 | 577.61 | 0.00096075
| Frequency0.01120 | Frequency | 120 | 1200.01 | 0.01 | kHz | Hz |  |  |  | 60 | 0.016 | 18.159 | 0.028435
| Frequency0.1120 | Frequency | 120 | 1200.01 | 0.1 | kHz | Hz |  |  |  | 60.876 | 0.018461 | 59.55 | 0.015104
| Frequency1120 | Frequency | 120 | 1200.01 | 1 | kHz | Hz |  |  |  | 578.01 | 0.0038671 | 577.51 | 0.0018144
| Frequency10120 | Frequency | 120 | 1200.01 | 10 | kHz | Hz |  |  |  | 5773.8 | 0.0010097 | 5773.6 | 0.00018186
| Frequency0.11200 | Frequency | 1200 | 2000 | 0.1 | kHz | kHz |  |  |  | 0.43 | 0.13 | -0.35824 | 0.36326
| Frequency11200 | Frequency | 1200 | 2000 | 1 | kHz | kHz |  |  |  | 0.43 | 0.14137 | 0.42566 | 0.12836
| Frequency101200 | Frequency | 1200 | 2000 | 10 | kHz | kHz |  |  |  | 5.7549 | 0.021951 | 5.757 | 0.014015
| Frequency1001200 | Frequency | 1200 | 2000 | 100 | kHz | kHz |  |  |  | 57.733 | 0.0045961 | 57.734 | 0.0014029
| Frequency0.12000 | Frequency | 2000 | 2001 | 0.1 | kHz | kHz |  |  |  | 0.43 | 0.13 | -0.35824 | 0.36326
| Frequency12000 | Frequency | 2000 | 2001 | 1 | kHz | kHz |  |  |  | 0.43 | 0.14137 | 0.42566 | 0.12836
| Frequency102000 | Frequency | 2000 | 2001 | 10 | kHz | kHz |  |  |  | 5.7549 | 0.021951 | 5.757 | 0.014015
| Frequency1002000 | Frequency | 2000 | 2001 | 100 | kHz | kHz |  |  |  | 57.733 | 0.0045961 | 57.734 | 0.0014029
| ACA0.0000000100.01 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 0.13 | 2300 | 0.11565 | 2314.2
| ACA0.000000100.01 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 0.13 | 2300 | 0.12522 | 2290.9
| ACA0.00000100.01 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 0.56558 | 1474.9 | 0.56533 | 1474.5
| ACA0.0000100.01 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 5.7994 | 124.28 | 5.7993 | 123.73
| ACA0.000100.01 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 58.025 | -857.45 | 58.025 | -857.67
| ACA0.00100.01 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.01 | 0.02001 | kHz | 580.26 | -8787.4 | 580.03 | -8787.5
| ACA0.0100.01 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.01 | 0.02001 | kHz | 5802.6 | -87896 | 5800 | -87896
| ACA0.100.01 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.01 | 0.02001 | kHz | 58026 | -878960 | 58000 | -878960
| ACA100.01 | ACA | 0 | 0.00033001 | 1 | A | µA | 0.01 | 0.02001 | kHz | 580260 | -8789600 | 580000 | -8789600
| ACA1000.01 | ACA | 0 | 0.00033001 | 10 | A | µA | 0.01 | 0.02001 | kHz | 5802600 | -87896000 | 5800000 | -87896000
| ACA0.0000000100.02 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 0.13 | 1700 | 0.11582 | 1731.6
| ACA0.000000100.02 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 0.13 | 1700 | 0.1262 | 1707.4
| ACA0.00000100.02 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 0.57237 | 986.73 | 0.57213 | 986.13
| ACA0.0000100.02 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 5.7723 | 128.02 | 5.7722 | 127.48
| ACA0.000100.02 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 57.735 | 13.015 | 57.735 | 12.795
| ACA0.00100.02 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.02 | 0.04501 | kHz | 577.36 | 1.3674 | 577.35 | 1.2795
| ACA0.0100.02 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.02 | 0.04501 | kHz | 5773.6 | 0.16308 | 5773.6 | 0.12796
| ACA0.100.02 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.02 | 0.04501 | kHz | 57736 | 0.026847 | 57736 | 0.012796
| ACA100.02 | ACA | 0 | 0.00033001 | 1 | A | µA | 0.02 | 0.04501 | kHz | 577360 | 0.0068998 | 577360 | 0.0012794
| ACA1000.02 | ACA | 0 | 0.00033001 | 10 | A | µA | 0.02 | 0.04501 | kHz | 5773600 | 0.0023794 | 5773600 | 0.00012995
| ACA0.0000000100.045 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.14 | 1400 | 0.11585 | 1442.9
| ACA0.000000100.045 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.14 | 1400 | 0.12669 | 1418.5
| ACA0.00000100.045 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.5767 | 759.64 | 0.57647 | 758.92
| ACA0.0000100.045 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.7731 | 93.964 | 5.773 | 93.42
| ACA0.000100.045 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.736 | 9.5875 | 57.735 | 9.3679
| ACA0.00100.045 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 0.045 | 1 | kHz | 577.36 | 1.0246 | 577.36 | 0.93682
| ACA0.0100.045 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 0.045 | 1 | kHz | 5773.6 | 0.12881 | 5773.6 | 0.093682
| ACA0.100.045 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 0.045 | 1 | kHz | 57736 | 0.02342 | 57736 | 0.0093682
| ACA100.045 | ACA | 0 | 0.00033001 | 1 | A | µA | 0.045 | 1 | kHz | 577360 | 0.0065572 | 577360 | 0.00093674
| ACA1000.045 | ACA | 0 | 0.00033001 | 10 | A | µA | 0.045 | 1 | kHz | 5773600 | 0.0023422 | 5773600 | 9.5917e-05
| ACA0.0000000101 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 0.18 | 3500 | 0.17344 | 3463.7
| ACA0.000000101 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 0.18 | 3500 | 0.17984 | 3448.1
| ACA0.00000101 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 0.56231 | 2652.9 | 0.56207 | 2652.8
| ACA0.0000101 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 5.7665 | 471.27 | 5.7664 | 470.74
| ACA0.000101 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 57.735 | 47.918 | 57.735 | 47.699
| ACA0.00101 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 1 | 5.01 | kHz | 577.36 | 4.8583 | 577.35 | 4.7705
| ACA0.0101 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 1 | 5.01 | kHz | 5773.6 | 0.51218 | 5773.6 | 0.47705
| ACA0.101 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 1 | 5.01 | kHz | 57736 | 0.061756 | 57736 | 0.047705
| ACA101 | ACA | 0 | 0.00033001 | 1 | A | µA | 1 | 5.01 | kHz | 577360 | 0.010391 | 577360 | 0.0047703
| ACA1001 | ACA | 0 | 0.00033001 | 10 | A | µA | 1 | 5.01 | kHz | 5773600 | 0.0027259 | 5773600 | 0.00047649
| ACA0.0000000105 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 0.25 | 9200 | 0.23106 | 9237.4
| ACA0.000000105 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 0.25 | 9200 | 0.23463 | 9228.1
| ACA0.00000105 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 0.5163 | 8526.7 | 0.51576 | 8527.5
| ACA0.0000105 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 5.7139 | 2807.7 | 5.7137 | 2806.8
| ACA0.000105 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 57.729 | 302.48 | 57.729 | 302.01
| ACA0.00105 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 5 | 10.01 | kHz | 577.35 | 30.413 | 577.35 | 30.226
| ACA0.0105 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 5 | 10.01 | kHz | 5773.6 | 3.0976 | 5773.5 | 3.0226
| ACA0.105 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 5 | 10.01 | kHz | 57736 | 0.33224 | 57736 | 0.30226
| ACA105 | ACA | 0 | 0.00033001 | 1 | A | µA | 5 | 10.01 | kHz | 577360 | 0.04222 | 577360 | 0.030226
| ACA1005 | ACA | 0 | 0.00033001 | 10 | A | µA | 5 | 10.01 | kHz | 5773600 | 0.0078188 | 5773600 | 0.0030229
| ACA0.00000001010 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46198 | 18475
| ACA0.0000001010 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46389 | 18470
| ACA0.000001010 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 10 | 30 | kHz | 0.6298 | 18044 | 0.62958 | 18044
| ACA0.00001010 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 10 | 30 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
| ACA0.0001010 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 10 | 30 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
| ACA0.001010 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 10 | 30 | kHz | 577.35 | 121.09 | 577.32 | 120.9
| ACA0.01010 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 10 | 30 | kHz | 5773.6 | 12.165 | 5773.5 | 12.09
| ACA0.1010 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 10 | 30 | kHz | 57736 | 1.239 | 57735 | 1.209
| ACA1010 | ACA | 0 | 0.00033001 | 1 | A | µA | 10 | 30 | kHz | 577360 | 0.1329 | 577360 | 0.1209
| ACA10010 | ACA | 0 | 0.00033001 | 10 | A | µA | 10 | 30 | kHz | 5773600 | 0.016888 | 5773600 | 0.012089
| ACA0.00000001030 | ACA | 0 | 0.00033001 | 1e-08 | A | µA | 30 | 30.1 | kHz | 0.62 | 18000 | 0.46198 | 18475
| ACA0.0000001030 | ACA | 0 | 0.00033001 | 1e-07 | A | µA | 30 | 30.1 | kHz | 0.62 | 18000 | 0.46389 | 18470
| ACA0.000001030 | ACA | 0 | 0.00033001 | 1e-06 | A | µA | 30 | 30.1 | kHz | 0.6298 | 18044 | 0.62958 | 18044
| ACA0.00001030 | ACA | 0 | 0.00033001 | 1e-05 | A | µA | 30 | 30.1 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
| ACA0.0001030 | ACA | 0 | 0.00033001 | 0.0001 | A | µA | 30 | 30.1 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
| ACA0.001030 | ACA | 0 | 0.00033001 | 0.001 | A | µA | 30 | 30.1 | kHz | 577.35 | 121.09 | 577.32 | 120.9
| ACA0.01030 | ACA | 0 | 0.00033001 | 0.01 | A | µA | 30 | 30.1 | kHz | 5773.6 | 12.165 | 5773.5 | 12.09
| ACA0.1030 | ACA | 0 | 0.00033001 | 0.1 | A | µA | 30 | 30.1 | kHz | 57736 | 1.239 | 57735 | 1.209
| ACA1030 | ACA | 0 | 0.00033001 | 1 | A | µA | 30 | 30.1 | kHz | 577360 | 0.1329 | 577360 | 0.1209
| ACA10030 | ACA | 0 | 0.00033001 | 10 | A | µA | 30 | 30.1 | kHz | 5773600 | 0.016888 | 5773600 | 0.012089
| ACA0.000000010.000330.01 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 0.23 | 2300 | 0.17326 | 2315.4
| ACA0.00000010.000330.01 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 0.23 | 2300 | 0.17481 | 2314.9
| ACA0.0000010.000330.01 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 0.35294 | 2267.3 | 0.35227 | 2267.5
| ACA0.000010.000330.01 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 5.4199 | 1301.8 | 5.4195 | 1301.8
| ACA0.00010.000330.01 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 57.686 | 174.69 | 57.685 | 174.66
| ACA0.000000010.000330.02 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 0.32 | 1400 | 0.17327 | 1443.8
| ACA0.00000010.000330.02 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 0.32 | 1400 | 0.17604 | 1443.1
| ACA0.0000010.000330.02 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 0.41399 | 1381.1 | 0.41349 | 1381.2
| ACA0.000010.000330.02 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 5.6118 | 601.75 | 5.6114 | 601.72
| ACA0.00010.000330.02 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 57.716 | 69.766 | 57.716 | 69.735
| ACA0.000000010.000330.045 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17375 | 1155.2
| ACA0.00000010.000330.045 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17651 | 1154.4
| ACA0.0000010.000330.045 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.43013 | 1124.2 | 0.44201 | 1086.4
| ACA0.000010.000330.045 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.6654 | 409.29 | 5.6651 | 409.25
| ACA0.00010.000330.045 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.723 | 45.397 | 57.723 | 45.366
| ACA0.000000010.000331 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 0.27 | 2300 | 0.23112 | 2309.7
| ACA0.00000010.000331 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 0.27 | 2300 | 0.2328 | 2309.2
| ACA0.0000010.000331 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 0.40201 | 2264.3 | 0.40156 | 2264.4
| ACA0.000010.000331 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 5.4266 | 1309.3 | 5.4263 | 1309.3
| ACA0.00010.000331 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 57.686 | 176.15 | 57.686 | 176.12
| ACA0.000000010.000335 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 0.35 | 5800 | 0.34654 | 5773.6
| ACA0.00000010.000335 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 0.35 | 5800 | 0.34724 | 5773.4
| ACA0.0000010.000335 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 0.41769 | 5779.5 | 0.4264 | 5752
| ACA0.000010.000335 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 4.6373 | 4728.3 | 4.6369 | 4728.3
| ACA0.00010.000335 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 57.432 | 1053.3 | 57.432 | 1053.2
| ACA0.000000010.0003310 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69332 | 11547
| ACA0.00000010.0003310 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69339 | 11547
| ACA0.0000010.0003310 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.7339 | 11536
| ACA0.000010.0003310 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 10 | 30 | kHz | 3.6592 | 11103 | 3.7773 | 10742
| ACA0.00010.0003310 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 10 | 30 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
| ACA0.000000010.0003330 | ACA | 0.00033 | 0.0033001 | 1e-08 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.69332 | 11547
| ACA0.00000010.0003330 | ACA | 0.00033 | 0.0033001 | 1e-07 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.69339 | 11547
| ACA0.0000010.0003330 | ACA | 0.00033 | 0.0033001 | 1e-06 | A | µA | 30 | 30.1 | kHz | 0.7 | 12000 | 0.7339 | 11536
| ACA0.000010.0003330 | ACA | 0.00033 | 0.0033001 | 1e-05 | A | µA | 30 | 30.1 | kHz | 3.6592 | 11103 | 3.7773 | 10742
| ACA0.00010.0003330 | ACA | 0.00033 | 0.0033001 | 0.0001 | A | µA | 30 | 30.1 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
| ACA0.000000010.00330.01 | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.01 | 0.02001 | kHz | 2.4 | 2100 | 2.3005 | 2087.6
| ACA0.00000010.00330.01 | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.01 | 0.02001 | kHz | 2.4 | 2100 | 2.3157 | 2087.1
| ACA0.0000010.00330.01 | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.01 | 0.02001 | kHz | 4.0897 | 2048.8 | 4.1185 | 2039.5
| ACA0.000010.00330.01 | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.01 | 0.02001 | kHz | 54.774 | 1117.8 | 54.773 | 1117.7
| ACA0.00010.00330.01 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.01 | 0.02001 | kHz | 576.95 | 144.81 | 576.95 | 144.75
| ACA0.000000010.00330.02 | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.02 | 0.04501 | kHz | 3.8 | 1000 | 2.2968 | 1043.6
| ACA0.00000010.00330.02 | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.02 | 0.04501 | kHz | 3.8 | 1000 | 2.3282 | 1042.8
| ACA0.0000010.00330.02 | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.02 | 0.04501 | kHz | 4.917 | 977.92 | 4.9145 | 977.93
| ACA0.000010.00330.02 | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.02 | 0.04501 | kHz | 56.865 | 350.66 | 56.863 | 350.55
| ACA0.00010.00330.02 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.02 | 0.04501 | kHz | 577.26 | 38.407 | 577.26 | 38.349
| ACA0.000000010.00330.045 | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 0.045 | 1 | kHz | 2.4 | 470 | 2.2998 | 465.62
| ACA0.00000010.00330.045 | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 0.045 | 1 | kHz | 2.4 | 470 | 2.3464 | 464.49
| ACA0.0000010.00330.045 | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 0.045 | 1 | kHz | 5.6396 | 392.41 | 5.6369 | 392.36
| ACA0.000010.00330.045 | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 0.045 | 1 | kHz | 57.584 | 84.81 | 57.583 | 84.672
| ACA0.00010.00330.045 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 0.045 | 1 | kHz | 577.34 | 8.7266 | 577.34 | 8.6681
| ACA0.000000010.00331 | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 1 | 5.01 | kHz | 2.4 | 930 | 2.3048 | 925.75
| ACA0.00000010.00331 | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 1 | 5.01 | kHz | 2.4 | 930 | 2.3386 | 924.88
| ACA0.0000010.00331 | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 1 | 5.01 | kHz | 5.0491 | 857.92 | 5.0466 | 857.92
| ACA0.000010.00331 | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 1 | 5.01 | kHz | 57.047 | 284.47 | 57.045 | 284.35
| ACA0.00010.00331 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 1 | 5.01 | kHz | 577.28 | 30.67 | 577.28 | 30.612
| ACA0.000000010.00335 | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 5 | 10.01 | kHz | 3.8 | 2300 | 3.4644 | 2309.4
| ACA0.00000010.00335 | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 5 | 10.01 | kHz | 3.8 | 2300 | 3.4807 | 2309
| ACA0.0000010.00335 | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 5 | 10.01 | kHz | 5.0186 | 2268.7 | 5.0129 | 2268.8
| ACA0.000010.00335 | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 5 | 10.01 | kHz | 54.397 | 1333.5 | 54.39 | 1333.5
| ACA0.00010.00335 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 5 | 10.01 | kHz | 576.87 | 180.75 | 576.87 | 180.65
| ACA0.000000010.003310 | ACA | 0.0033 | 0.033001 | 1e-08 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
| ACA0.00000010.003310 | ACA | 0.0033 | 0.033001 | 1e-07 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
| ACA0.0000010.003310 | ACA | 0.0033 | 0.033001 | 1e-06 | A | µA | 10 | 30 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
| ACA0.000010.003310 | ACA | 0.0033 | 0.033001 | 1e-05 | A | µA | 10 | 30 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
| ACA0.00010.003310 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 10 | 30 | kHz | 575.41 | 695.3 | 575.4 | 695.06
| ACA0.00010.003330 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
| ACA0.00010.003330 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
| ACA0.00010.003330 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
| ACA0.00010.003330 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
| ACA0.00010.003330 | ACA | 0.0033 | 0.033001 | 0.0001 | A | µA | 30 | 30.01 | kHz | 575.41 | 695.3 | 575.4 | 695.06
| ACA0.0000010.0330.01 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.01 | 0.02001 | kHz | 0.023 | 2.1 | 0.022917 | 2.0907
| ACA0.000010.0330.01 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.01 | 0.02001 | kHz | 0.023 | 2.1 | 0.023129 | 2.0901
| ACA0.00010.0330.01 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.01 | 0.02001 | kHz | 0.041052 | 2.0453 | 0.041148 | 2.0425
| ACA0.0010.0330.01 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.01 | 0.02001 | kHz | 0.54767 | 1.1202 | 0.54766 | 1.1202
| ACA0.010.0330.01 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.01 | 0.02001 | kHz | 5.7695 | 0.14517 | 5.7695 | 0.14515
| ACA0.0000010.0330.02 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.02 | 0.04501 | kHz | 0.038 | 1 | 0.023094 | 1.0436
| ACA0.000010.0330.02 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.02 | 0.04501 | kHz | 0.038 | 1 | 0.023358 | 1.0428
| ACA0.00010.0330.02 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.02 | 0.04501 | kHz | 0.049203 | 0.97804 | 0.049192 | 0.97804
| ACA0.0010.0330.02 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.02 | 0.04501 | kHz | 0.56864 | 0.35073 | 0.56864 | 0.35068
| ACA0.010.0330.02 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.02 | 0.04501 | kHz | 5.7726 | 0.03839 | 5.7726 | 0.038365
| ACA0.0000010.0330.045 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.02282 | 0.47204
| ACA0.000010.0330.045 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.023285 | 0.47091
| ACA0.00010.0330.045 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.056197 | 0.39853 | 0.056185 | 0.39851
| ACA0.0010.0330.045 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.57578 | 0.086645 | 0.57577 | 0.086584
| ACA0.010.0330.045 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7734 | 0.0088941 | 5.7734 | 0.0088682
| ACA0.0000010.0331 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 1 | 5.01 | kHz | 0.058 | 1.2 | 0.057624 | 1.1588
| ACA0.000010.0331 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 1 | 5.01 | kHz | 0.058 | 1.2 | 0.057825 | 1.1583
| ACA0.00010.0331 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 1 | 5.01 | kHz | 0.073934 | 1.1517 | 0.075029 | 1.1175
| ACA0.0010.0331 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 1 | 5.01 | kHz | 0.56965 | 0.47368 | 0.56963 | 0.47361
| ACA0.010.0331 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 1 | 5.01 | kHz | 5.7726 | 0.053742 | 5.7726 | 0.053697
| ACA0.0000010.0335 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 5 | 10.01 | kHz | 0.12 | 2.3 | 0.11543 | 2.3114
| ACA0.000010.0335 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 5 | 10.01 | kHz | 0.12 | 2.3 | 0.11552 | 2.3112
| ACA0.00010.0335 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 5 | 10.01 | kHz | 0.12471 | 2.2891 | 0.12466 | 2.2892
| ACA0.0010.0335 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 5 | 10.01 | kHz | 0.55922 | 1.4904 | 0.55916 | 1.4904
| ACA0.010.0335 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 5 | 10.01 | kHz | 5.7697 | 0.21298 | 5.7697 | 0.21288
| ACA0.0000010.03310 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
| ACA0.000010.03310 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
| ACA0.00010.03310 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
| ACA0.0010.03310 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 10 | 30 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
| ACA0.010.03310 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 10 | 30 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
| ACA0.0000010.03330 | ACA | 0.033 | 0.33001 | 1e-06 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
| ACA0.000010.03330 | ACA | 0.033 | 0.33001 | 1e-05 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
| ACA0.00010.03330 | ACA | 0.033 | 0.33001 | 0.0001 | A | mA | 30 | 30.01 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
| ACA0.0010.03330 | ACA | 0.033 | 0.33001 | 0.001 | A | mA | 30 | 30.01 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
| ACA0.010.03330 | ACA | 0.033 | 0.33001 | 0.01 | A | mA | 30 | 30.01 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
| ACA0.000010.330.01 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11617 | 2.0784
| ACA0.00010.330.01 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11882 | 2.0766
| ACA0.0010.330.01 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.35291 | 1.9255 | 0.35284 | 1.9254
| ACA0.010.330.01 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 5.6471 | 0.55145 | 5.6471 | 0.55121
| ACA0.10.330.01 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 57.722 | 0.057755 | 57.722 | 0.057649
| ACA0.000010.330.045 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.12 | 0.58 | 0.11734 | 0.57739
| ACA0.00010.330.045 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.12267 | 0.57758 | 0.124 | 0.57333
| ACA0.0010.330.045 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.52825 | 0.38235 | 0.52823 | 0.382
| ACA0.010.330.045 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7643 | 0.053023 | 5.7643 | 0.05276
| ACA0.10.330.045 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.735 | 0.005408 | 57.735 | 0.0053018
| ACA0.000010.331 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 1 | 5.01 | kHz | 1.2 | 6.9 | 1.1549 | 6.9282
| ACA0.00010.331 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 1 | 5.01 | kHz | 1.2 | 6.9 | 1.1555 | 6.9278
| ACA0.0010.331 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 1 | 5.01 | kHz | 1.2164 | 6.8903 | 1.2155 | 6.8904
| ACA0.010.331 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 1 | 5.01 | kHz | 5.1011 | 4.9149 | 5.0998 | 4.9136
| ACA0.10.331 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 1 | 5.01 | kHz | 57.598 | 0.72965 | 57.598 | 0.72819
| ACA0.000010.335 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7734 | 28.868
| ACA0.00010.335 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7742 | 28.867
| ACA0.0010.335 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7874 | 28.859
| ACA0.010.335 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 5 | 10 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
| ACA0.10.335 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 5 | 10 | kHz | 55.816 | 11.862 | 55.815 | 11.86
| ACA0.000010.3310 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7734 | 28.868
| ACA0.00010.3310 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7742 | 28.867
| ACA0.0010.3310 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7874 | 28.859
| ACA0.010.3310 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 10 | 10.01 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
| ACA0.10.3310 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 10 | 10.01 | kHz | 55.816 | 11.862 | 55.815 | 11.86
| ACA0.000011.10.01 | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11589 | 2.0787
| ACA0.00011.10.01 | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 0.12 | 2.1 | 0.11676 | 2.0785
| ACA0.0011.10.01 | ACA | 1.1 | 3.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.18892 | 2.077 | 0.20868 | 2.0565
| ACA0.011.10.01 | ACA | 1.1 | 3.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 4.9071 | 1.2281 | 4.9043 | 1.2265
| ACA0.11.10.01 | ACA | 1.1 | 3.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 57.613 | 0.15838 | 57.613 | 0.15706
| ACA0.000011.10.045 | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11619 | 0.69369
| ACA0.00011.10.045 | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11868 | 0.6931
| ACA0.0011.10.045 | ACA | 1.1 | 3.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.35192 | 0.64234 | 0.34551 | 0.6421
| ACA0.011.10.045 | ACA | 1.1 | 3.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.6463 | 0.18057 | 5.6448 | 0.17758
| ACA0.11.10.045 | ACA | 1.1 | 3.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.722 | 0.019812 | 57.722 | 0.018474
| ACA0.000011.11 | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1548 | 6.9283
| ACA0.00011.11 | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1552 | 6.9282
| ACA0.0011.11 | ACA | 1.1 | 3.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 1.3 | 6.9 | 1.1804 | 6.9223
| ACA0.011.11 | ACA | 1.1 | 3.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 3.4547 | 6.4116 | 3.4521 | 6.4115
| ACA0.11.11 | ACA | 1.1 | 3.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 56.451 | 1.7722 | 56.451 | 1.771
| ACA0.000011.15 | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.773 | 28.868
| ACA0.00011.15 | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7731 | 28.868
| ACA0.0011.15 | ACA | 1.1 | 3.001 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.78 | 28.866
| ACA0.011.15 | ACA | 1.1 | 3.001 | 0.01 | A | mA | 5 | 10 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
| ACA0.11.15 | ACA | 1.1 | 3.001 | 0.1 | A | mA | 5 | 10 | kHz | 45.662 | 21.093 | 45.659 | 21.092
| ACA0.000011.110 | ACA | 1.1 | 3.001 | 1e-05 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.773 | 28.868
| ACA0.00011.110 | ACA | 1.1 | 3.001 | 0.0001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.7731 | 28.868
| ACA0.0011.110 | ACA | 1.1 | 3.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 5.8 | 29 | 5.78 | 28.866
| ACA0.011.110 | ACA | 1.1 | 3.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
| ACA0.11.110 | ACA | 1.1 | 3.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 45.662 | 21.093 | 45.659 | 21.092
| ACA0.000130.045 | ACA | 3 | 11.001 | 0.0001 | A | mA | 0.045 | 0.10001 | kHz | 2.4 | 0.69 | 2.3298 | 0.69373
| ACA0.00130.045 | ACA | 3 | 11.001 | 0.001 | A | mA | 0.045 | 0.10001 | kHz | 2.4 | 0.69 | 2.3709 | 0.69112
| ACA0.0130.045 | ACA | 3 | 11.001 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 5.6644 | 0.53041 | 5.6622 | 0.53026
| ACA0.130.045 | ACA | 3 | 11.001 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 57.647 | 0.085551 | 57.646 | 0.085343
| ACA130.045 | ACA | 3 | 11.001 | 1 | A | mA | 0.045 | 0.10001 | kHz | 577.35 | 0.0086938 | 577.35 | 0.0086087
| ACA0.000130.1 | ACA | 3 | 11.001 | 0.0001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.324 | 1.1553
| ACA0.00130.1 | ACA | 3 | 11.001 | 0.001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.3551 | 1.1531
| ACA0.0130.1 | ACA | 3 | 11.001 | 0.01 | A | mA | 0.1 | 1 | kHz | 5.1972 | 0.99041 | 5.1951 | 0.99036
| ACA0.130.1 | ACA | 3 | 11.001 | 0.1 | A | mA | 0.1 | 1 | kHz | 57.412 | 0.20433 | 57.411 | 0.20413
| ACA130.1 | ACA | 3 | 11.001 | 1 | A | mA | 0.1 | 1 | kHz | 577.32 | 0.020884 | 577.32 | 0.020799
| ACA0.000131 | ACA | 3 | 11.001 | 0.0001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3103 | 34.641
| ACA0.00131 | ACA | 3 | 11.001 | 0.001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3112 | 34.641
| ACA0.0131 | ACA | 3 | 11.001 | 0.01 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.5084 | 34.627
| ACA0.131 | ACA | 3 | 11.001 | 0.1 | A | mA | 1 | 5 | kHz | 20.888 | 33.346 | 20.867 | 33.347
| ACA131 | ACA | 3 | 11.001 | 1 | A | mA | 1 | 5 | kHz | 547.31 | 13.25 | 547.3 | 13.249
| ACA0.000135 | ACA | 3 | 11.001 | 0.0001 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.3103 | 34.641
| ACA0.00135 | ACA | 3 | 11.001 | 0.001 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.3112 | 34.641
| ACA0.0135 | ACA | 3 | 11.001 | 0.01 | A | mA | 5 | 5.01 | kHz | 2.5 | 35 | 2.5084 | 34.627
| ACA0.135 | ACA | 3 | 11.001 | 0.1 | A | mA | 5 | 5.01 | kHz | 20.888 | 33.346 | 20.867 | 33.347
| ACA135 | ACA | 3 | 11.001 | 1 | A | mA | 5 | 5.01 | kHz | 547.31 | 13.25 | 547.3 | 13.249
| ACA0.0001110.045 | ACA | 11 | 20.501 | 0.0001 | A | mA | 0.045 | 0.10001 | kHz | 5.8 | 1.4 | 5.7745 | 1.386
| ACA0.001110.045 | ACA | 11 | 20.501 | 0.001 | A | mA | 0.045 | 0.10001 | kHz | 5.8 | 1.4 | 5.7855 | 1.3857
| ACA0.01110.045 | ACA | 11 | 20.501 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 6.9249 | 1.3567 | 6.9028 | 1.3548
| ACA0.1110.045 | ACA | 11 | 20.501 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 55.153 | 0.59996 | 55.168 | 0.5931
| ACA1110.045 | ACA | 11 | 20.501 | 1 | A | mA | 0.045 | 0.10001 | kHz | 579.65 | 0.069317 | 579.66 | 0.065887
| ACA0.0001110.1 | ACA | 11 | 20.501 | 0.0001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7754 | 1.7324
| ACA0.001110.1 | ACA | 11 | 20.501 | 0.001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7852 | 1.7321
| ACA0.01110.1 | ACA | 11 | 20.501 | 0.01 | A | mA | 0.1 | 1 | kHz | 6.7671 | 1.7062 | 6.7466 | 1.7047
| ACA0.1110.1 | ACA | 11 | 20.501 | 0.1 | A | mA | 0.1 | 1 | kHz | 53.699 | 0.85925 | 53.709 | 0.85296
| ACA1110.1 | ACA | 11 | 20.501 | 1 | A | mA | 0.1 | 1 | kHz | 579.44 | 0.102 | 579.45 | 0.098572
| ACA0.0001111 | ACA | 11 | 20.501 | 0.0001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7742 | 34.641
| ACA0.001111 | ACA | 11 | 20.501 | 0.001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7745 | 34.641
| ACA0.01111 | ACA | 11 | 20.501 | 0.01 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.8395 | 34.639
| ACA0.1111 | ACA | 11 | 20.501 | 0.1 | A | mA | 1 | 5 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
| ACA1111 | ACA | 11 | 20.501 | 1 | A | mA | 1 | 5 | kHz | 437.55 | 23.603 | 437.54 | 23.602
| ACA0.0001115 | ACA | 11 | 20.501 | 0.0001 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.7742 | 34.641
| ACA0.001115 | ACA | 11 | 20.501 | 0.001 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.7745 | 34.641
| ACA0.01115 | ACA | 11 | 20.501 | 0.01 | A | mA | 5 | 5.01 | kHz | 5.9 | 35 | 5.8395 | 34.639
| ACA0.1115 | ACA | 11 | 20.501 | 0.1 | A | mA | 5 | 5.01 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
| ACA1115 | ACA | 11 | 20.501 | 1 | A | mA | 5 | 5.01 | kHz | 437.55 | 23.603 | 437.54 | 23.602
| ACA 10 turn0.000110.01 | ACA 10 turn | 1 | 11.001 | 0.0001 | A | mA | 0.01 | 0.04501 | kHz | 3.3 | 2.7 | 0.97318 | 2.873
| ACA 10 turn0.00110.01 | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 4.3265 | 2.6149 | 4.3204 | 2.6148
| ACA 10 turn0.0110.01 | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 57.287 | 0.84059 | 57.285 | 0.83952
| ACA 10 turn0.110.01 | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 579.92 | 0.090643 | 579.92 | 0.090136
| ACA 10 turn110.01 | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 5800 | 0.0092241 | 5800 | 0.0090207
| ACA 10 turn1010.01 | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 58000 | 0.00098346 | 58000 | 0.00090208
| ACA 10 turn0.00110.045 | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.66 | 2.1 | 0.65508 | 2.0543
| ACA 10 turn0.0110.045 | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 4.6183 | 1.7587 | 4.6169 | 1.7583
| ACA 10 turn0.110.045 | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.618 | 0.44323 | 57.618 | 0.44244
| ACA 10 turn110.045 | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 0.045 | 1 | kHz | 579.96 | 0.046302 | 579.96 | 0.045959
| ACA 10 turn1010.045 | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 0.045 | 1 | kHz | 5800 | 0.004735 | 5800 | 0.0045977
| ACA 10 turn0.00111 | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.405 | 7.1882
| ACA 10 turn0.0111 | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 12.361 | 7.1187 | 12.349 | 7.1191
| ACA 10 turn0.111 | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 56.262 | 4.6574 | 56.248 | 4.6564
| ACA 10 turn111 | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 1 | 5.01 | kHz | 579.63 | 0.67288 | 579.63 | 0.67163
| ACA 10 turn1011 | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 1 | 5.01 | kHz | 5800 | 0.068091 | 5800 | 0.067582
| ACA 10 turn0.00115 | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.703 | 28.929
| ACA 10 turn0.0115 | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 28.914 | 57.905 | 28.915
| ACA 10 turn0.115 | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 5 | 10 | kHz | 76.668 | 27.611 | 76.638 | 27.612
| ACA 10 turn115 | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 5 | 10 | kHz | 575.98 | 10.476 | 575.96 | 10.474
| ACA 10 turn1015 | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 5 | 10 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
| ACA 10 turn0.001110 | ACA 10 turn | 1 | 11.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.703 | 28.929
| ACA 10 turn0.01110 | ACA 10 turn | 1 | 11.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 58 | 28.914 | 57.905 | 28.915
| ACA 10 turn0.1110 | ACA 10 turn | 1 | 11.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 76.668 | 27.611 | 76.638 | 27.612
| ACA 10 turn1110 | ACA 10 turn | 1 | 11.001 | 1 | A | mA | 10 | 10.01 | kHz | 575.98 | 10.476 | 575.96 | 10.474
| ACA 10 turn10110 | ACA 10 turn | 1 | 11.001 | 10 | A | mA | 10 | 10.01 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
| ACA 10 turn0.001110.01 | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | 0.87 | 2.9 | 0.85389 | 2.8839
| ACA 10 turn0.01110.01 | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | 1.4281 | 2.8814 | 1.5429 | 2.8673
| ACA 10 turn0.1110.01 | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 44.394 | 2.0188 | 44.317 | 2.0184
| ACA 10 turn1110.01 | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 577.68 | 0.29774 | 577.66 | 0.2963
| ACA 10 turn10110.01 | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 5799.8 | 0.030408 | 5799.8 | 0.029819
| ACA 10 turn0.001110.045 | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.9 | 2.1 | 0.42297 | 2.1155
| ACA 10 turn0.01110.045 | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 1.3908 | 2.0927 | 1.3648 | 2.0928
| ACA 10 turn0.1110.045 | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 48.986 | 1.2457 | 48.967 | 1.244
| ACA 10 turn1110.045 | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 0.045 | 1 | kHz | 578.74 | 0.1605 | 578.73 | 0.15918
| ACA 10 turn10110.045 | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 0.045 | 1 | kHz | 5799.9 | 0.016507 | 5799.9 | 0.015972
| ACA 10 turn0.001111 | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.174 | 7.2092
| ACA 10 turn0.01111 | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 12 | 7.2 | 11.422 | 7.2034
| ACA 10 turn0.1111 | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 33.827 | 6.6978 | 33.792 | 6.6979
| ACA 10 turn1111 | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 1 | 5.01 | kHz | 566.19 | 1.8955 | 566.18 | 1.8942
| ACA 10 turn10111 | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 1 | 5.01 | kHz | 5798.5 | 0.19808 | 5798.5 | 0.1975
| ACA 10 turn0.001115 | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.627 | 28.936
| ACA 10 turn0.01115 | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.682 | 28.935
| ACA 10 turn0.1115 | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 5 | 10 | kHz | 63.353 | 28.822 | 63.613 | 28.797
| ACA 10 turn1115 | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 5 | 10 | kHz | 458.92 | 21.118 | 458.89 | 21.117
| ACA 10 turn10115 | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 5 | 10 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
| ACA 10 turn0.0011110 | ACA 10 turn | 11 | 30.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.627 | 28.936
| ACA 10 turn0.011110 | ACA 10 turn | 11 | 30.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 58 | 29 | 57.682 | 28.935
| ACA 10 turn0.11110 | ACA 10 turn | 11 | 30.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 63.353 | 28.822 | 63.613 | 28.797
| ACA 10 turn11110 | ACA 10 turn | 11 | 30.001 | 1 | A | mA | 10 | 10.01 | kHz | 458.92 | 21.118 | 458.89 | 21.117
| ACA 10 turn101110 | ACA 10 turn | 11 | 30.001 | 10 | A | mA | 10 | 10.01 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
| ACA 10 turn0.01300.045 | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 12 | 2.1 | 11.984 | 2.086
| ACA 10 turn0.1300.045 | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 36.579 | 1.9251 | 36.551 | 1.9245
| ACA 10 turn1300.045 | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 568.43 | 0.54491 | 568.44 | 0.54347
| ACA 10 turn10300.045 | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5798.8 | 0.057436 | 5798.8 | 0.056801
| ACA 10 turn100300.045 | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 58000 | 0.0059371 | 58000 | 0.0056828
| ACA 10 turn0.01300.1 | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 15 | 2.3 | 14.872 | 2.2866
| ACA 10 turn0.1300.1 | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 37.348 | 2.1392 | 37.319 | 2.1387
| ACA 10 turn1300.1 | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 0.1 | 1 | kHz | 566.33 | 0.65492 | 566.34 | 0.65351
| ACA 10 turn10300.1 | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 0.1 | 1 | kHz | 5798.5 | 0.069561 | 5798.5 | 0.068926
| ACA 10 turn100300.1 | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0071509 | 58000 | 0.0068966
| ACA 10 turn0.01301 | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 1 | 5 | kHz | 24 | 35 | 23.064 | 34.699
| ACA 10 turn0.1301 | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 1 | 5 | kHz | 24 | 35 | 25.112 | 34.684
| ACA 10 turn1301 | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 1 | 5 | kHz | 210.04 | 33.396 | 210.02 | 33.395
| ACA 10 turn10301 | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 1 | 5 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
| ACA 10 turn100301 | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 1 | 5 | kHz | 57966 | 1.4661 | 57966 | 1.4652
| ACA 10 turn0.01305 | ACA 10 turn | 30 | 110.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 24 | 35 | 23.064 | 34.699
| ACA 10 turn0.1305 | ACA 10 turn | 30 | 110.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 24 | 35 | 25.112 | 34.684
| ACA 10 turn1305 | ACA 10 turn | 30 | 110.01 | 1 | A | mA | 5 | 5.01 | kHz | 210.04 | 33.396 | 210.02 | 33.395
| ACA 10 turn10305 | ACA 10 turn | 30 | 110.01 | 10 | A | mA | 5 | 5.01 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
| ACA 10 turn100305 | ACA 10 turn | 30 | 110.01 | 100 | A | mA | 5 | 5.01 | kHz | 57966 | 1.4661 | 57966 | 1.4652
| ACA 10 turn0.011100.045 | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 42 | 2.4 | 38.553 | 2.4159
| ACA 10 turn0.11100.045 | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 46.787 | 2.3924 | 46.65 | 2.3916
| ACA 10 turn11100.045 | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 500.59 | 1.4072 | 500.63 | 1.403
| ACA 10 turn101100.045 | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5788.7 | 0.17673 | 5788.8 | 0.17407
| ACA 10 turn1001100.045 | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57999 | 0.018526 | 57999 | 0.017455
| ACA 10 turn0.011100.1 | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 42.202 | 2.6324
| ACA 10 turn0.11100.1 | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 49.765 | 2.6108 | 49.649 | 2.61
| ACA 10 turn11100.1 | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 0.1 | 1 | kHz | 490.55 | 1.6184 | 490.58 | 1.6144
| ACA 10 turn101100.1 | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 0.1 | 1 | kHz | 5786.7 | 0.2093 | 5786.8 | 0.20664
| ACA 10 turn1001100.1 | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 0.1 | 1 | kHz | 57999 | 0.021803 | 57999 | 0.020732
| ACA 10 turn0.011101 | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 1 | 5 | kHz | 58 | 35 | 57.612 | 34.699
| ACA 10 turn0.11101 | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 1 | 5 | kHz | 58 | 35 | 58.262 | 34.697
| ACA 10 turn11101 | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 1 | 5 | kHz | 79.07 | 34.897 | 123.66 | 34.491
| ACA 10 turn101101 | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 1 | 5 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
| ACA 10 turn1001101 | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 1 | 5 | kHz | 57768 | 3.2889 | 57768 | 3.2878
| ACA 10 turn0.011105 | ACA 10 turn | 110 | 205.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 58 | 35 | 57.612 | 34.699
| ACA 10 turn0.11105 | ACA 10 turn | 110 | 205.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 58 | 35 | 58.262 | 34.697
| ACA 10 turn11105 | ACA 10 turn | 110 | 205.01 | 1 | A | mA | 5 | 5.01 | kHz | 79.07 | 34.897 | 123.66 | 34.491
| ACA 10 turn101105 | ACA 10 turn | 110 | 205.01 | 10 | A | mA | 5 | 5.01 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
| ACA 10 turn1001105 | ACA 10 turn | 110 | 205.01 | 100 | A | mA | 5 | 5.01 | kHz | 57768 | 3.2889 | 57768 | 3.2878
| ACA 20 turn0.001220.01 | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 0.01 | 0.04501 | kHz | -5.2 | 3 | -6.3515 | 3.0358
| ACA 20 turn0.01220.01 | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 0.01 | 0.04501 | kHz | -5.2 | 3.0309 | -6.0166 | 3.031
| ACA 20 turn0.1220.01 | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 0.01 | 0.04501 | kHz | 24.504 | 2.6609 | 24.465 | 2.6608
| ACA 20 turn1220.01 | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 0.01 | 0.04501 | kHz | 569.88 | 0.59872 | 569.88 | 0.59787
| ACA 20 turn10220.01 | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 0.01 | 0.04501 | kHz | 5799 | 0.061619 | 5799 | 0.061255
| ACA 20 turn0.001220.045 | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -5.078 | 2.227
| ACA 20 turn0.01220.045 | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -4.615 | 2.2203
| ACA 20 turn0.1220.045 | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 0.045 | 1 | kHz | 25.632 | 2.1228 | 32.901 | 1.7903
| ACA 20 turn1220.045 | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 0.045 | 1 | kHz | 574.48 | 0.32468 | 574.48 | 0.3238
| ACA 20 turn10220.045 | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 0.045 | 1 | kHz | 5799.4 | 0.033173 | 5799.4 | 0.032809
| ACA 20 turn0.001221 | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 1 | 5.01 | kHz | 2.3 | 7.6 | 2.2215 | 7.5886
| ACA 20 turn0.01221 | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 1 | 5.01 | kHz | 2.3482 | 7.5869 | 2.3433 | 7.5869
| ACA 20 turn0.1221 | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 1 | 5.01 | kHz | 14.11 | 7.4273 | 14.091 | 7.4273
| ACA 20 turn1221 | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 1 | 5.01 | kHz | 526.05 | 3.5184 | 526.05 | 3.5178
| ACA 20 turn10221 | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 1 | 5.01 | kHz | 5793.5 | 0.40619 | 5793.5 | 0.40583
| ACA 20 turn0.001225 | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 5 | 10 | kHz | 52 | 30 | 34.529 | 30.459
| ACA 20 turn0.01225 | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 5 | 10 | kHz | 52 | 30.458 | 34.545 | 30.459
| ACA 20 turn0.1225 | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 5 | 10 | kHz | 52 | 30.419 | 37.46 | 30.419
| ACA 20 turn1225 | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 5 | 10 | kHz | 304.91 | 27.255 | 304.84 | 27.255
| ACA 20 turn10225 | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 5 | 10 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
| ACA 20 turn0.0012210 | ACA 20 turn | 22 | 60.001 | 0.001 | A | mA | 10 | 10.01 | kHz | 52 | 30 | 34.529 | 30.459
| ACA 20 turn0.012210 | ACA 20 turn | 22 | 60.001 | 0.01 | A | mA | 10 | 10.01 | kHz | 52 | 30.458 | 34.545 | 30.459
| ACA 20 turn0.12210 | ACA 20 turn | 22 | 60.001 | 0.1 | A | mA | 10 | 10.01 | kHz | 52 | 30.419 | 37.46 | 30.419
| ACA 20 turn12210 | ACA 20 turn | 22 | 60.001 | 1 | A | mA | 10 | 10.01 | kHz | 304.91 | 27.255 | 304.84 | 27.255
| ACA 20 turn102210 | ACA 20 turn | 22 | 60.001 | 10 | A | mA | 10 | 10.01 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
| ACA 20 turn0.01600.045 | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 24 | 2.1 | 23.534 | 2.0875
| ACA 20 turn0.1600.045 | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 37.146 | 2.0415 | 37.093 | 2.0415
| ACA 20 turn1600.045 | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 540.32 | 0.97471 | 540.3 | 0.97417
| ACA 20 turn10600.045 | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5795.1 | 0.11376 | 5795.1 | 0.11344
| ACA 20 turn100600.045 | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 58000 | 0.011493 | 58000 | 0.011365
| ACA 20 turn0.01600.1 | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 30 | 2.3 | 29.364 | 2.2879
| ACA 20 turn0.1600.1 | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 41.544 | 2.2475 | 41.547 | 2.2467
| ACA 20 turn1600.1 | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 0.1 | 1 | kHz | 534.45 | 1.1508 | 534.42 | 1.1503
| ACA 20 turn10600.1 | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 0.1 | 1 | kHz | 5794.1 | 0.13793 | 5794.1 | 0.13761
| ACA 20 turn100600.1 | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.01392 | 58000 | 0.013793
| ACA 20 turn0.01601 | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 1 | 5 | kHz | 47 | 35 | 46.103 | 34.699
| ACA 20 turn0.1601 | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 1 | 5 | kHz | 47 | 35 | 47.146 | 34.695
| ACA 20 turn1601 | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 1 | 5 | kHz | 127.73 | 34.633 | 144.7 | 34.35
| ACA 20 turn10601 | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 1 | 5 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
| ACA 20 turn100601 | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 1 | 5 | kHz | 57864 | 2.9205 | 57864 | 2.9201
| ACA 20 turn0.01605 | ACA 20 turn | 60 | 220.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 47 | 35 | 46.103 | 34.699
| ACA 20 turn0.1605 | ACA 20 turn | 60 | 220.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 47 | 35 | 47.146 | 34.695
| ACA 20 turn1605 | ACA 20 turn | 60 | 220.01 | 1 | A | mA | 5 | 5.01 | kHz | 127.73 | 34.633 | 144.7 | 34.35
| ACA 20 turn10605 | ACA 20 turn | 60 | 220.01 | 10 | A | mA | 5 | 5.01 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
| ACA 20 turn100605 | ACA 20 turn | 60 | 220.01 | 100 | A | mA | 5 | 5.01 | kHz | 57864 | 2.9205 | 57864 | 2.9201
| ACA 20 turn0.012200.045 | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 84 | 2.4 | 76.978 | 2.4161
| ACA 20 turn0.12200.045 | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 84 | 2.4069 | 81.077 | 2.4099
| ACA 20 turn12200.045 | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 407.38 | 1.9712 | 407.21 | 1.9701
| ACA 20 turn102200.045 | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5755.8 | 0.34673 | 5755.9 | 0.34526
| ACA 20 turn1002200.045 | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57996 | 0.03551 | 57996 | 0.034907
| ACA 20 turn0.012200.1 | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 98 | 2.6 | 84.28 | 2.6326
| ACA 20 turn0.12200.1 | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 98 | 2.6065 | 88.046 | 2.6269
| ACA 20 turn12200.1 | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 0.1 | 1 | kHz | 396.27 | 2.2061 | 396.12 | 2.205
| ACA 20 turn102200.1 | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 0.1 | 1 | kHz | 5747.7 | 0.41071 | 5747.8 | 0.40925
| ACA 20 turn1002200.1 | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 0.1 | 1 | kHz | 57995 | 0.042063 | 57995 | 0.04146
| ACA 20 turn0.012201 | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.22 | 34.699
| ACA 20 turn0.12201 | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.69 | 34.698
| ACA 20 turn12201 | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 148.62 | 34.646
| ACA 20 turn102201 | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 1 | 5 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
| ACA 20 turn1002201 | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 1 | 5 | kHz | 57090 | 6.4831 | 57090 | 6.4816
| ACA 20 turn0.012205 | ACA 20 turn | 220 | 410.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.22 | 34.699
| ACA 20 turn0.12205 | ACA 20 turn | 220 | 410.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.69 | 34.698
| ACA 20 turn12205 | ACA 20 turn | 220 | 410.01 | 1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 148.62 | 34.646
| ACA 20 turn102205 | ACA 20 turn | 220 | 410.01 | 10 | A | mA | 5 | 5.01 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
| ACA 20 turn1002205 | ACA 20 turn | 220 | 410.01 | 100 | A | mA | 5 | 5.01 | kHz | 57090 | 6.4831 | 57090 | 6.4816
| ACA 50 turn0.001550.01 | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8 | 4.2301 | 2.8841
| ACA 50 turn0.01550.01 | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.7673 | 4.3733 | 2.8834
| ACA 50 turn0.1550.01 | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8076 | 17.829 | 2.819
| ACA 50 turn1550.01 | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.01 | 0.045 | kHz | 530.85 | 1.3025 | 530.83 | 1.3019
| ACA 50 turn10550.01 | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.01 | 0.045 | kHz | 5794.1 | 0.14924 | 5794.1 | 0.14886
| ACA 50 turn100550.01 | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.01506 | 57992 | 0.01491
| ACA 50 turn0.001550.045 | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
| ACA 50 turn0.01550.045 | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
| ACA 50 turn0.1550.045 | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 20.257 | 2.0293
| ACA 50 turn1550.045 | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.045 | 0.1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
| ACA 50 turn10550.045 | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.045 | 0.1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
| ACA 50 turn100550.045 | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
| ACA 50 turn0.001550.1 | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
| ACA 50 turn0.01550.1 | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
| ACA 50 turn0.1550.1 | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 20.257 | 2.0293
| ACA 50 turn1550.1 | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 0.1 | 1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
| ACA 50 turn10550.1 | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 0.1 | 1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
| ACA 50 turn100550.1 | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
| ACA 50 turn0.001551 | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 1 | 5.01 | kHz | 58 | 7.2 | 55.853 | 7.2093
| ACA 50 turn0.01551 | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 1 | 5.01 | kHz | 58 | 7.209 | 55.901 | 7.2091
| ACA 50 turn0.1551 | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 1 | 5.01 | kHz | 60.896 | 7.1858 | 60.843 | 7.1859
| ACA 50 turn1551 | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 1 | 5.01 | kHz | 422.41 | 5.6956 | 422.3 | 5.6954
| ACA 50 turn10551 | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 1 | 5.01 | kHz | 5763.9 | 0.97779 | 5763.9 | 0.97715
| ACA 50 turn100551 | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 1 | 5.01 | kHz | 57996 | 0.099047 | 57948 | 0.098783
| ACA 50 turn0.001555 | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 5 | 10 | kHz | 290 | 29 | 288.13 | 28.936
| ACA 50 turn0.01555 | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 5 | 10 | kHz | 290 | 28.936 | 288.14 | 28.936
| ACA 50 turn0.1555 | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 5 | 10 | kHz | 290 | 28.931 | 289.4 | 28.93
| ACA 50 turn1555 | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 5 | 10 | kHz | 405.34 | 28.398 | 405.23 | 28.397
| ACA 50 turn10555 | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 5 | 10 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
| ACA 50 turn100555 | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 5 | 10 | kHz | 57941 | 1.6237 | 57264 | 1.6205
| ACA 50 turn0.0015510 | ACA 50 turn | 55 | 150.01 | 0.001 | A | mA | 10 | 10.01 | kHz | 290 | 29 | 288.13 | 28.936
| ACA 50 turn0.015510 | ACA 50 turn | 55 | 150.01 | 0.01 | A | mA | 10 | 10.01 | kHz | 290 | 28.936 | 288.14 | 28.936
| ACA 50 turn0.15510 | ACA 50 turn | 55 | 150.01 | 0.1 | A | mA | 10 | 10.01 | kHz | 290 | 28.931 | 289.4 | 28.93
| ACA 50 turn15510 | ACA 50 turn | 55 | 150.01 | 1 | A | mA | 10 | 10.01 | kHz | 405.34 | 28.398 | 405.23 | 28.397
| ACA 50 turn105510 | ACA 50 turn | 55 | 150.01 | 10 | A | mA | 10 | 10.01 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
| ACA 50 turn1005510 | ACA 50 turn | 55 | 150.01 | 100 | A | mA | 10 | 10.01 | kHz | 57941 | 1.6237 | 57264 | 1.6205
| ACA 50 turn0.011500.045 | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 59 | 2.1 | 58.542 | 2.0879
| ACA 50 turn0.11500.045 | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 62.205 | 2.0942 | 64.156 | 2.0802
| ACA 50 turn11500.045 | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 445.36 | 1.6257 | 445.1 | 1.6253
| ACA 50 turn101500.045 | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5769.8 | 0.28152 | 5769.8 | 0.28084
| ACA 50 turn1001500.045 | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57997 | 0.02869 | 57997 | 0.028411
| ACA 50 turn0.011500.1 | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 74 | 2.3 | 73.171 | 2.2882
| ACA 50 turn0.11500.1 | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 76.076 | 2.2962 | 78.162 | 2.2814
| ACA 50 turn11500.1 | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 0.1 | 1 | kHz | 437.31 | 1.8466 | 437.03 | 1.8463
| ACA 50 turn101500.1 | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 0.1 | 1 | kHz | 5764 | 0.34064 | 5763.9 | 0.33997
| ACA 50 turn1001500.1 | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 0.1 | 1 | kHz | 57996 | 0.034758 | 57996 | 0.034478
| ACA 50 turn0.011501 | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.24 | 34.699
| ACA 50 turn0.11501 | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.76 | 34.698
| ACA 50 turn11501 | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 155.32 | 34.642
| ACA 50 turn101501 | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 1 | 5 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
| ACA 50 turn1001501 | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 1 | 5 | kHz | 57174 | 7.1298 | 57174 | 7.1293
| ACA 50 turn0.011505 | ACA 50 turn | 150 | 550.01 | 0.01 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.24 | 34.699
| ACA 50 turn0.11505 | ACA 50 turn | 150 | 550.01 | 0.1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 115.76 | 34.698
| ACA 50 turn11505 | ACA 50 turn | 150 | 550.01 | 1 | A | mA | 5 | 5.01 | kHz | 120 | 35 | 155.32 | 34.642
| ACA 50 turn101505 | ACA 50 turn | 150 | 550.01 | 10 | A | mA | 5 | 5.01 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
| ACA 50 turn1001505 | ACA 50 turn | 150 | 550.01 | 100 | A | mA | 5 | 5.01 | kHz | 57174 | 7.1298 | 57174 | 7.1293
| ACA 50 turn0.015500.045 | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 0.045 | 0.10001 | kHz | 210 | 2.4 | 192.41 | 2.4161
| ACA 50 turn0.15500.045 | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 0.045 | 0.10001 | kHz | 210 | 2.4 | 193.98 | 2.4152
| ACA 50 turn15500.045 | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 0.045 | 0.10001 | kHz | 350.89 | 2.3236 | 350.74 | 2.3224
| ACA 50 turn105500.045 | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 0.045 | 0.10001 | kHz | 5545.6 | 0.82078 | 5546.4 | 0.81772
| ACA 50 turn1005500.045 | ACA 50 turn | 550 | 1025.01 | 100 | A | mA | 0.045 | 0.10001 | kHz | 57972 | 0.088607 | 57972 | 0.087215
| ACA 50 turn0.015500.1 | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
| ACA 50 turn0.15500.1 | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
| ACA 50 turn15500.1 | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 0.1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
| ACA 50 turn105500.1 | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 0.1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
| ACA 50 turn0.015501 | ACA 50 turn | 550 | 1025.01 | 0.01 | A | mA | 1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
| ACA 50 turn0.15501 | ACA 50 turn | 550 | 1025.01 | 0.1 | A | mA | 1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
| ACA 50 turn15501 | ACA 50 turn | 550 | 1025.01 | 1 | A | mA | 1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
| ACA 50 turn105501 | ACA 50 turn | 550 | 1025.01 | 10 | A | mA | 1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
| Capacitance0.00000000010.00000022 | Capacitance | 2.2e-07 | 3.3001e-06 | 1e-10 | mF | nF |  |  |  | 0.012 | 5800 | 0.01155 | 5775.2
| Capacitance0.0000000010.00000022 | Capacitance | 2.2e-07 | 3.3001e-06 | 1e-09 | mF | nF |  |  |  | 0.012 | 5800 | 0.011563 | 5772.8
| Capacitance0.000000010.00000022 | Capacitance | 2.2e-07 | 3.3001e-06 | 1e-08 | mF | nF |  |  |  | 0.012845 | 5547.8 | 0.012838 | 5547.8
| Capacitance0.00000010.00000022 | Capacitance | 2.2e-07 | 3.3001e-06 | 1e-07 | mF | nF |  |  |  | 0.058705 | 2016.9 | 0.058698 | 2014.5
| Capacitance0.0000010.00000022 | Capacitance | 2.2e-07 | 3.3001e-06 | 1e-06 | mF | nF |  |  |  | 0.57745 | 218.19 | 0.57745 | 217
| Capacitance0.0000000010.0000033 | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-09 | mF | nF |  |  |  | 0.012 | 2900 | 0.011576 | 2885.2
| Capacitance0.000000010.0000033 | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-08 | mF | nF |  |  |  | 0.012482 | 2856.2 | 0.012511 | 2834.7
| Capacitance0.00000010.0000033 | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-07 | mF | nF |  |  |  | 0.056911 | 1398.7 | 0.056878 | 1390.6
| Capacitance0.0000010.0000033 | Capacitance | 3.3e-06 | 1.1001e-05 | 1e-06 | mF | nF |  |  |  | 0.57722 | 165.28 | 0.57721 | 160.66
| Capacitance0.0000000010.000011 | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-09 | mF | nF |  |  |  | 0.12 | 2900 | 0.11547 | 2886.7
| Capacitance0.000000010.000011 | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-08 | mF | nF |  |  |  | 0.12 | 2900 | 0.11561 | 2885.2
| Capacitance0.00000010.000011 | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-07 | mF | nF |  |  |  | 0.12801 | 2745.1 | 0.12797 | 2743.5
| Capacitance0.0000010.000011 | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-06 | mF | nF |  |  |  | 0.58647 | 856.77 | 0.58644 | 853.69
| Capacitance0.000010.000011 | Capacitance | 1.1e-05 | 3.3001e-05 | 1e-05 | mF | nF |  |  |  | 5.7745 | 853.27 | 5.7744 | 89.445
| Capacitance0.000000010.000033 | Capacitance | 3.3e-05 | 0.00011001 | 1e-08 | mF | nF |  |  |  | 0.12 | 2900 | 0.11565 | 2886.2
| Capacitance0.00000010.000033 | Capacitance | 3.3e-05 | 0.00011001 | 1e-07 | mF | nF |  |  |  | 0.12459 | 2858.3 | 0.125 | 2835.7
| Capacitance0.0000010.000033 | Capacitance | 3.3e-05 | 0.00011001 | 1e-06 | mF | nF |  |  |  | 0.56898 | 1398 | 0.56874 | 1390.9
| Capacitance0.000010.000033 | Capacitance | 3.3e-05 | 0.00011001 | 1e-05 | mF | nF |  |  |  | 5.7722 | 164.7 | 5.7721 | 160.69
| Capacitance0.000000010.00011 | Capacitance | 0.00011 | 0.00033001 | 1e-08 | mF | nF |  |  |  | 0.35 | 2900 | 0.34649 | 2886.7
| Capacitance0.00000010.00011 | Capacitance | 0.00011 | 0.00033001 | 1e-07 | mF | nF |  |  |  | 0.35 | 2900 | 0.34958 | 2881.2
| Capacitance0.0000010.00011 | Capacitance | 0.00011 | 0.00033001 | 1e-06 | mF | nF |  |  |  | 0.60931 | 2463.1 | 0.60906 | 2462.3
| Capacitance0.000010.00011 | Capacitance | 0.00011 | 0.00033001 | 1e-05 | mF | nF |  |  |  | 5.7585 | 484.5 | 5.7585 | 483.14
| Capacitance0.00000010.00033 | Capacitance | 0.00033 | 0.0011001 | 1e-07 | mF | nF |  |  |  | 1.2 | 2900 | 1.1566 | 2886.2
| Capacitance0.0000010.00033 | Capacitance | 0.00033 | 0.0011001 | 1e-06 | mF | nF |  |  |  | 1.2438 | 2860.2 | 1.2501 | 2835.6
| Capacitance0.000010.00033 | Capacitance | 0.00033 | 0.0011001 | 1e-05 | mF | nF |  |  |  | 5.6891 | 1394 | 5.6874 | 1390.9
| Capacitance0.00010.00033 | Capacitance | 0.00033 | 0.0011001 | 0.0001 | mF | nF |  |  |  | 57.722 | 162.48 | 57.721 | 160.69
| Capacitance0.00000010.0011 | Capacitance | 0.0011 | 0.0033001 | 1e-07 | mF | nF |  |  |  | 3.5 | 2900 | 3.4649 | 2886.7
| Capacitance0.0000010.0011 | Capacitance | 0.0011 | 0.0033001 | 1e-06 | mF | nF |  |  |  | 3.5 | 2900 | 3.4958 | 2881.2
| Capacitance0.000010.0011 | Capacitance | 0.0011 | 0.0033001 | 1e-05 | mF | nF |  |  |  | 6.0937 | 2463.2 | 6.0906 | 2462.3
| Capacitance0.00010.0011 | Capacitance | 0.0011 | 0.0033001 | 0.0001 | mF | nF |  |  |  | 57.585 | 484.68 | 57.585 | 483.14
| Capacitance0.0000010.0033 | Capacitance | 0.0033 | 0.011001 | 1e-06 | mF | nF |  |  |  | 12 | 2900 | 11.566 | 2886.2
| Capacitance0.000010.0033 | Capacitance | 0.0033 | 0.011001 | 1e-05 | mF | nF |  |  |  | 12.429 | 2861.1 | 12.501 | 2835.6
| Capacitance0.00010.0033 | Capacitance | 0.0033 | 0.011001 | 0.0001 | mF | nF |  |  |  | 56.88 | 1394.3 | 56.874 | 1390.9
| Capacitance0.0010.0033 | Capacitance | 0.0033 | 0.011001 | 0.001 | mF | nF |  |  |  | 577.21 | 162.54 | 577.21 | 160.69
| Capacitance0.0000010.011 | Capacitance | 0.011 | 0.033001 | 1e-06 | mF | nF |  |  |  | 36 | 4600 | 34.648 | 4618.7
| Capacitance0.000010.011 | Capacitance | 0.011 | 0.033001 | 1e-05 | mF | nF |  |  |  | 36 | 4600 | 34.894 | 4613.9
| Capacitance0.00010.011 | Capacitance | 0.011 | 0.033001 | 0.0001 | mF | nF |  |  |  | 56.814 | 4211.3 | 56.808 | 4211
| Capacitance0.0010.011 | Capacitance | 0.011 | 0.033001 | 0.001 | mF | nF |  |  |  | 572.02 | 1057.8 | 572.02 | 1057.2
| Capacitance0.000010.033 | Capacitance | 0.033 | 0.11001 | 1e-05 | mF | nF |  |  |  | 120 | 5200 | 115.61 | 5195.6
| Capacitance0.00010.033 | Capacitance | 0.033 | 0.11001 | 0.0001 | mF | nF |  |  |  | 121.85 | 5183.2 | 122.72 | 5152.8
| Capacitance0.0010.033 | Capacitance | 0.033 | 0.11001 | 0.001 | mF | nF |  |  |  | 536.63 | 3284.2 | 536.45 | 3282
| Capacitance0.010.033 | Capacitance | 0.033 | 0.11001 | 0.01 | mF | nF |  |  |  | 5766.3 | 438.5 | 5766.3 | 436.53
| Capacitance0.000010.11 | Capacitance | 0.11 | 0.33001 | 1e-05 | mF | µF |  |  |  | 0.35 | 5.2 | 0.34613 | 5.1998
| Capacitance0.00010.11 | Capacitance | 0.11 | 0.33001 | 0.0001 | mF | µF |  |  |  | 0.35 | 5.2 | 0.34836 | 5.1953
| Capacitance0.0010.11 | Capacitance | 0.11 | 0.33001 | 0.001 | mF | µF |  |  |  | 0.55623 | 4.8038 | 0.55608 | 4.8037
| Capacitance0.010.11 | Capacitance | 0.11 | 0.33001 | 0.01 | mF | µF |  |  |  | 5.7038 | 1.2944 | 5.7038 | 1.2938
| Capacitance0.00010.33 | Capacitance | 0.33 | 1.1001 | 0.0001 | mF | µF |  |  |  | 1.2 | 5.2 | 1.1558 | 5.1983
| Capacitance0.0010.33 | Capacitance | 0.33 | 1.1001 | 0.001 | mF | µF |  |  |  | 1.2193 | 5.1825 | 1.2268 | 5.1555
| Capacitance0.010.33 | Capacitance | 0.33 | 1.1001 | 0.01 | mF | µF |  |  |  | 5.3658 | 3.2865 | 5.3641 | 3.2842
| Capacitance0.10.33 | Capacitance | 0.33 | 1.1001 | 0.1 | mF | µF |  |  |  | 57.663 | 0.43886 | 57.663 | 0.43689
| Capacitance0.00011.1 | Capacitance | 1.1 | 3.3001 | 0.0001 | mF | µF |  |  |  | 3.5 | 5.2 | 3.4636 | 5.1989
| Capacitance0.0011.1 | Capacitance | 1.1 | 3.3001 | 0.001 | mF | µF |  |  |  | 3.5 | 5.2 | 3.4866 | 5.1944
| Capacitance0.011.1 | Capacitance | 1.1 | 3.3001 | 0.01 | mF | µF |  |  |  | 5.5761 | 4.8039 | 5.5634 | 4.8029
| Capacitance0.11.1 | Capacitance | 1.1 | 3.3001 | 0.1 | mF | µF |  |  |  | 57.043 | 1.2991 | 57.038 | 1.2937
| Capacitance0.0013.3 | Capacitance | 3.3 | 11.001 | 0.001 | mF | µF |  |  |  | 13 | 5.2 | 12.076 | 5.1983
| Capacitance0.013.3 | Capacitance | 3.3 | 11.001 | 0.01 | mF | µF |  |  |  | 13 | 5.2 | 12.713 | 5.1556
| Capacitance0.13.3 | Capacitance | 3.3 | 11.001 | 0.1 | mF | µF |  |  |  | 53.666 | 3.2879 | 53.64 | 3.2843
| Capacitance13.3 | Capacitance | 3.3 | 11.001 | 1 | mF | µF |  |  |  | 576.63 | 0.43999 | 576.63 | 0.4369
| Capacitance0.00111 | Capacitance | 11 | 33.001 | 0.001 | mF | µF |  |  |  | 35 | 8.7 | 34.642 | 8.6619
| Capacitance0.0111 | Capacitance | 11 | 33.001 | 0.01 | mF | µF |  |  |  | 35 | 8.7 | 34.806 | 8.6585
| Capacitance0.111 | Capacitance | 11 | 33.001 | 0.1 | mF | µF |  |  |  | 50.508 | 8.3397 | 50.437 | 8.3396
| Capacitance111 | Capacitance | 11 | 33.001 | 1 | mF | µF |  |  |  | 557.55 | 3.1193 | 557.52 | 3.1158
| Capacitance0.0133 | Capacitance | 33 | 110.01 | 0.01 | mF | µF* |  |  |  | 120 | 13 | 115.63 | 12.701
| Capacitance0.133 | Capacitance | 33 | 110.01 | 0.1 | mF | µF* |  |  |  | 120 | 13 | 119.5 | 12.676
| Capacitance133 | Capacitance | 33 | 110.01 | 1 | mF | µF* |  |  |  | 430.99 | 10.809 | 430.28 | 10.808
| Capacitance1033 | Capacitance | 33 | 110.01 | 10 | mF | µF* |  |  |  | 5725.6 | 2.2156 | 5725.3 | 2.2101
| Capacitance0.01110 | Capacitance | 110 | 110.1 | 0.01 | mF | µF |  |  |  | 120 | 13 | 115.63 | 12.701
| Capacitance0.1110 | Capacitance | 110 | 110.1 | 0.1 | mF | µF |  |  |  | 120 | 13 | 119.5 | 12.676
| Capacitance1110 | Capacitance | 110 | 110.1 | 1 | mF | µF |  |  |  | 430.99 | 10.809 | 430.28 | 10.808
| Capacitance10110 | Capacitance | 110 | 110.1 | 10 | mF | µF |  |  |  | 5725.6 | 2.2156 | 5725.3 | 2.2101
| CDE0.1-250 | CDE | -250 | -99.9 | 0.1 | °C | °C |  |  |  | 0.58423 | 0 | 0.58258 | 0
| CDE0.1-100 | CDE | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.19859 | 0 | 0.19765 | 0
| CDE0.1-25 | CDE | -25 | 350.1 | 0.1 | °C | °C |  |  |  | 0.17588 | 0 | 0.17482 | 0
| CDE0.1350 | CDE | 350 | 650.1 | 0.1 | °C | °C |  |  |  | 0.19683 | 0 | 0.19628 | 0
| CDE0.1650 | CDE | 650 | 1000 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
| CDE0.11000 | CDE | 1000 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
| CDE1-250 | CDE | -250 | -99.9 | 1 | °C | °C |  |  |  | 0.82286 | 0 | 0.81817 | 0
| CDE1-100 | CDE | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.60874 | 0 | 0.60751 | 0
| CDE1-25 | CDE | -25 | 350.1 | 1 | °C | °C |  |  |  | 0.60171 | 0 | 0.60047 | 0
| CDE1350 | CDE | 350 | 650.1 | 1 | °C | °C |  |  |  | 0.60777 | 0 | 0.60706 | 0
| CDE1650 | CDE | 650 | 1000 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
| CDE11000 | CDE | 1000 | 1000.1 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
| CDJ0.1-210 | CDJ | -210 | -99.9 | 0.1 | °C | °C |  |  |  | 0.32482 | 0 | 0.31951 | 0
| CDJ0.1-100 | CDJ | -100 | -29.9 | 0.1 | °C | °C |  |  |  | 0.20266 | 0 | 0.19649 | 0
| CDJ0.1-30 | CDJ | -30 | 150.1 | 0.1 | °C | °C |  |  |  | 0.18177 | 0 | 0.17487 | 0
| CDJ0.1150 | CDJ | 150 | 760.1 | 0.1 | °C | °C |  |  |  | 0.2089 | 0 | 0.20722 | 0
| CDJ0.1760 | CDJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
| CDJ0.11200 | CDJ | 1200 | 1200.1 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
| CDJ1-210 | CDJ | -210 | -99.9 | 1 | °C | °C |  |  |  | 0.66767 | 0 | 0.65733 | 0
| CDJ1-100 | CDJ | -100 | -29.9 | 1 | °C | °C |  |  |  | 0.61519 | 0 | 0.60713 | 0
| CDJ1-30 | CDJ | -30 | 150.1 | 1 | °C | °C |  |  |  | 0.60863 | 0 | 0.60048 | 0
| CDJ1150 | CDJ | 150 | 760.1 | 1 | °C | °C |  |  |  | 0.61296 | 0 | 0.61069 | 0
| CDJ1760 | CDJ | 760 | 1200 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
| CDJ11200 | CDJ | 1200 | 1200.1 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
| CDK0.1-200 | CDK | -200 | -99.9 | 0.1 | °C | °C |  |  |  | 0.39349 | 0 | 0.38767 | 0
| CDK0.1-100 | CDK | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.22869 | 0 | 0.21853 | 0
| CDK0.1-25 | CDK | -25 | 120.1 | 0.1 | °C | °C |  |  |  | 0.20428 | 0 | 0.19653 | 0
| CDK0.1120 | CDK | 120 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.31026 | 0 | 0.30758 | 0
| CDK0.11000 | CDK | 1000 | 1372 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
| CDK0.11372 | CDK | 1372 | 1372.1 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
| CDK1-200 | CDK | -200 | -99.9 | 1 | °C | °C |  |  |  | 0.70603 | 0 | 0.69303 | 0
| CDK1-100 | CDK | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.62924 | 0 | 0.61462 | 0
| CDK1-25 | CDK | -25 | 120.1 | 1 | °C | °C |  |  |  | 0.61728 | 0 | 0.60715 | 0
| CDK1120 | CDK | 120 | 1000.1 | 1 | °C | °C |  |  |  | 0.65669 | 0 | 0.65162 | 0
| CDK11000 | CDK | 1000 | 1372 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
| CDK11372 | CDK | 1372 | 1372.1 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
| CDN0.1-200 | CDN | -200 | -99.9 | 0.1 | °C | °C |  |  |  | 0.47201 | 0 | 0.46682 | 0
| CDN0.1-100 | CDN | -100 | -24.9 | 0.1 | °C | °C |  |  |  | 0.26968 | 0 | 0.26325 | 0
| CDN0.1-25 | CDN | -25 | 120.1 | 0.1 | °C | °C |  |  |  | 0.23706 | 0 | 0.22973 | 0
| CDN0.1120 | CDN | 120 | 410.1 | 0.1 | °C | °C |  |  |  | 0.22303 | 0 | 0.21855 | 0
| CDN0.1410 | CDN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
| CDN0.11300 | CDN | 1300 | 1300.1 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
| CDN1-200 | CDN | -200 | -99.9 | 1 | °C | °C |  |  |  | 0.75326 | 0 | 0.74022 | 0
| CDN1-100 | CDN | -100 | -24.9 | 1 | °C | °C |  |  |  | 0.64265 | 0 | 0.6319 | 0
| CDN1-25 | CDN | -25 | 120.1 | 1 | °C | °C |  |  |  | 0.62966 | 0 | 0.61869 | 0
| CDN1120 | CDN | 120 | 410.1 | 1 | °C | °C |  |  |  | 0.62103 | 0 | 0.61463 | 0
| CDN1410 | CDN | 410 | 1300 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
| CDN11300 | CDN | 1300 | 1300.1 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
| CDR0.10 | CDR | 0 | 250.1 | 0.1 | °C | °C |  |  |  | 0.66953 | 0 | 0.66592 | 0
| CDR0.1250 | CDR | 250 | 400.1 | 0.1 | °C | °C |  |  |  | 0.41654 | 0 | 0.41163 | 0
| CDR0.1400 | CDR | 400 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.39417 | 0 | 0.38897 | 0
| CDR0.11000 | CDR | 1000 | 1767 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
| CDR0.11767 | CDR | 1767 | 1767.1 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
| CDR10 | CDR | 0 | 250.1 | 1 | °C | °C |  |  |  | 0.89034 | 0 | 0.87946 | 0
| CDR1250 | CDR | 250 | 400.1 | 1 | °C | °C |  |  |  | 0.71813 | 0 | 0.70671 | 0
| CDR1400 | CDR | 400 | 1000.1 | 1 | °C | °C |  |  |  | 0.70538 | 0 | 0.69376 | 0
| CDR11000 | CDR | 1000 | 1767 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
| CDR11767 | CDR | 1767 | 1767.1 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
| CDS0.10 | CDS | 0 | 250.1 | 0.1 | °C | °C |  |  |  | 0.55632 | 0 | 0.55207 | 0
| CDS0.1250 | CDS | 250 | 1000.1 | 0.1 | °C | °C |  |  |  | 0.42876 | 0 | 0.42399 | 0
| CDS0.11000 | CDS | 1000 | 1400.1 | 0.1 | °C | °C |  |  |  | 0.4386 | 0 | 0.43394 | 0
| CDS0.11400 | CDS | 1400 | 1767 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
| CDS0.11767 | CDS | 1767 | 1767.1 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
| CDS10 | CDS | 0 | 250.1 | 1 | °C | °C |  |  |  | 0.80847 | 0 | 0.79673 | 0
| CDS1250 | CDS | 250 | 1000.1 | 1 | °C | °C |  |  |  | 0.72528 | 0 | 0.71398 | 0
| CDS11000 | CDS | 1000 | 1400.1 | 1 | °C | °C |  |  |  | 0.73114 | 0 | 0.71993 | 0
| CDS11400 | CDS | 1400 | 1767 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
| CDS11767 | CDS | 1767 | 1767.1 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
| CDT0.1-250 | CDT | -250 | -149.9 | 0.1 | °C | °C |  |  |  | 0.55536 | 0 | 0.55058 | 0
| CDT0.1-150 | CDT | -150 | 0.1 | 0.1 | °C | °C |  |  |  | 0.29308 | 0 | 0.28549 | 0
| CDT0.10 | CDT | 0 | 120.1 | 0.1 | °C | °C |  |  |  | 0.20015 | 0 | 0.19651 | 0
| CDT0.1120 | CDT | 120 | 400 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
| CDT0.1400 | CDT | 400 | 400.1 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
| CDT1-250 | CDT | -250 | -149.9 | 1 | °C | °C |  |  |  | 0.8089 | 0 | 0.7957 | 0
| CDT1-150 | CDT | -150 | 0.1 | 1 | °C | °C |  |  |  | 0.65502 | 0 | 0.64149 | 0
| CDT10 | CDT | 0 | 120.1 | 1 | °C | °C |  |  |  | 0.61188 | 0 | 0.60714 | 0
| CDT1120 | CDT | 120 | 400 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0
| CDT1400 | CDT | 400 | 400.1 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0
| CDPT1000.1-200 | CDPT100 | -200 | 0.1 | 0.1 | °C | °C |  |  |  | 0.13102 | 0 | 0.11627 | 0
| CDPT1000.10 | CDPT100 | 0 | 100.1 | 0.1 | °C | °C |  |  |  | 0.16571 | 0 | 0.15431 | 0
| CDPT1000.1100 | CDPT100 | 100 | 300.1 | 0.1 | °C | °C |  |  |  | 0.16571 | 0 | 0.15431 | 0
| CDPT1000.1300 | CDPT100 | 300 | 400.1 | 0.1 | °C | °C |  |  |  | 0.22203 | 0 | 0.21366 | 0
| CDPT1000.1400 | CDPT100 | 400 | 630.1 | 0.1 | °C | °C |  |  |  | 0.26 | 0 | 0.25288 | 0
| CDPT1000.1630 | CDPT100 | 630 | 800 | 0.1 | °C | °C |  |  |  | 0.43497 | 0 | 0.43075 | 0
| CDPT1000.1800 | CDPT100 | 800 | 800.1 | 0.1 | °C | °C |  |  |  | 0.43497 | 0 | 0.43075 | 0
| CDPT1001-200 | CDPT100 | -200 | 0.1 | 1 | °C | °C |  |  |  | 0.60095 | 0 | 0.58869 | 0
| CDPT10010 | CDPT100 | 0 | 100.1 | 1 | °C | °C |  |  |  | 0.60946 | 0 | 0.59736 | 0
| CDPT1001100 | CDPT100 | 100 | 300.1 | 1 | °C | °C |  |  |  | 0.60946 | 0 | 0.59736 | 0
| CDPT1001300 | CDPT100 | 300 | 400.1 | 1 | °C | °C |  |  |  | 0.62712 | 0 | 0.61537 | 0
| CDPT1001400 | CDPT100 | 400 | 630.1 | 1 | °C | °C |  |  |  | 0.64154 | 0 | 0.63006 | 0
| CDPT1001630 | CDPT100 | 630 | 800 | 1 | °C | °C |  |  |  | 0.73019 | 0 | 0.72012 | 0
| CDPT1001800 | CDPT100 | 800 | 800.1 | 1 | °C | °C |  |  |  | 0.73019 | 0 | 0.72012 | 0
| CDPT10000.1-200 | CDPT1000 | -200 | 0.1 | 0.1 | °C | °C |  |  |  | 0.10309 | 0 | 0.083548 | 0
| CDPT10000.10 | CDPT1000 | 0 | 100.1 | 0.1 | °C | °C |  |  |  | 0.11707 | 0 | 0.10029 | 0
| CDPT10000.1100 | CDPT1000 | 100 | 260.1 | 0.1 | °C | °C |  |  |  | 0.11707 | 0 | 0.10029 | 0
| CDPT10000.1260 | CDPT1000 | 260 | 300.1 | 0.1 | °C | °C |  |  |  | 0.15328 | 0 | 0.14088 | 0
| CDPT10000.1300 | CDPT1000 | 300 | 400.1 | 0.1 | °C | °C |  |  |  | 0.17 | 0 | 0.1589 | 0
| CDPT10000.1400 | CDPT1000 | 400 | 600.1 | 0.1 | °C | °C |  |  |  | 0.17204 | 0 | 0.16109 | 0
| CDPT10000.1600 | CDPT1000 | 600 | 630 | 0.1 | °C | °C |  |  |  | 0.47301 | 0 | 0.46914 | 0
| CDPT10000.1630 | CDPT1000 | 630 | 630.1 | 0.1 | °C | °C |  |  |  | 0.47301 | 0 | 0.46914 | 0
| CDPT10001-200 | CDPT1000 | -200 | 0.1 | 1 | °C | °C |  |  |  | 0.59549 | 0 | 0.58311 | 0
| CDPT100010 | CDPT1000 | 0 | 100.1 | 1 | °C | °C |  |  |  | 0.59806 | 0 | 0.58574 | 0
| CDPT10001100 | CDPT1000 | 100 | 260.1 | 1 | °C | °C |  |  |  | 0.59806 | 0 | 0.58574 | 0
| CDPT10001260 | CDPT1000 | 260 | 300.1 | 1 | °C | °C |  |  |  | 0.60619 | 0 | 0.59404 | 0
| CDPT10001300 | CDPT1000 | 300 | 400.1 | 1 | °C | °C |  |  |  | 0.61063 | 0 | 0.59857 | 0
| CDPT10001400 | CDPT1000 | 400 | 600.1 | 1 | °C | °C |  |  |  | 0.61121 | 0 | 0.59915 | 0
| CDPT10001600 | CDPT1000 | 600 | 630 | 1 | °C | °C |  |  |  | 0.75347 | 0 | 0.74372 | 0
| CDPT10001630 | CDPT1000 | 630 | 630.1 | 1 | °C | °C |  |  |  | 0.75347 | 0 | 0.74372 | 0
| FDE0.1-418 | FDE | -418 | -147.9 | 0.1 | °F | °F |  |  |  | 1.0462 | 0 | 1.0433 | 0
| FDE0.1-148 | FDE | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.3466 | 0 | 0.34484 | 0
| FDE0.1-13 | FDE | -13 | 662.1 | 0.1 | °F | °F |  |  |  | 0.30437 | 0 | 0.30237 | 0
| FDE0.1662 | FDE | 662 | 1202.1 | 0.1 | °F | °F |  |  |  | 0.34352 | 0 | 0.34251 | 0
| FDE0.11202 | FDE | 1202 | 1832 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
| FDE0.11832 | FDE | 1832 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
| FDE1-418 | FDE | -418 | -147.9 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
| FDE1-148 | FDE | -148 | -12.9 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
| FDE1-13 | FDE | -13 | 662.1 | 1 | °F | °F |  |  |  | 0.6529 | 0 | 0.64917 | 0
| FDE1662 | FDE | 662 | 1202.1 | 1 | °F | °F |  |  |  | 0.67089 | 0 | 0.66882 | 0
| FDE11202 | FDE | 1202 | 1832 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
| FDE11832 | FDE | 1832 | 1832.1 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
| FDJ0.1-346 | FDJ | -346 | -147.9 | 0.1 | °F | °F |  |  |  | 0.57727 | 0 | 0.56758 | 0
| FDJ0.1-148 | FDJ | -148 | -21.9 | 0.1 | °F | °F |  |  |  | 0.35408 | 0 | 0.34263 | 0
| FDJ0.1-22 | FDJ | -22 | 302.1 | 0.1 | °F | °F |  |  |  | 0.31531 | 0 | 0.3024 | 0
| FDJ0.1302 | FDJ | 302 | 1400.1 | 0.1 | °F | °F |  |  |  | 0.36584 | 0 | 0.36275 | 0
| FDJ0.11400 | FDJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
| FDJ0.12192 | FDJ | 2192 | 2192.1 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
| FDJ1-346 | FDJ | -346 | -147.9 | 1 | °F | °F |  |  |  | 0.83458 | 0 | 0.80755 | 0
| FDJ1-148 | FDJ | -148 | -21.9 | 1 | °F | °F |  |  |  | 0.69231 | 0 | 0.66888 | 0
| FDJ1-22 | FDJ | -22 | 302.1 | 1 | °F | °F |  |  |  | 0.67331 | 0 | 0.64919 | 0
| FDJ1302 | FDJ | 302 | 1400.1 | 1 | °F | °F |  |  |  | 0.686 | 0 | 0.6794 | 0
| FDJ11400 | FDJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
| FDJ12192 | FDJ | 2192 | 2192.1 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
| FDK0.1-328 | FDK | -328 | -147.9 | 0.1 | °F | °F |  |  |  | 0.70186 | 0 | 0.69128 | 0
| FDK0.1-148 | FDK | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.40198 | 0 | 0.38321 | 0
| FDK0.1-13 | FDK | -13 | 248.1 | 0.1 | °F | °F |  |  |  | 0.35702 | 0 | 0.34265 | 0
| FDK0.1248 | FDK | 248 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.55155 | 0 | 0.54665 | 0
| FDK0.11832 | FDK | 1832 | 2500 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
| FDK0.12500 | FDK | 2500 | 2500.1 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
| FDK1-328 | FDK | -328 | -147.9 | 1 | °F | °F |  |  |  | 0.93103 | 0 | 0.89881 | 0
| FDK1-148 | FDK | -148 | -12.9 | 1 | °F | °F |  |  |  | 0.73198 | 0 | 0.69054 | 0
| FDK1-13 | FDK | -13 | 248.1 | 1 | °F | °F |  |  |  | 0.69831 | 0 | 0.66889 | 0
| FDK1248 | FDK | 248 | 1832.1 | 1 | °F | °F |  |  |  | 0.80643 | 0 | 0.79299 | 0
| FDK11832 | FDK | 1832 | 2500 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
| FDK12500 | FDK | 2500 | 2500.1 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
| FDN0.1-328 | FDN | -328 | -147.9 | 0.1 | °F | °F |  |  |  | 0.84492 | 0 | 0.83553 | 0
| FDN0.1-148 | FDN | -148 | -12.9 | 0.1 | °F | °F |  |  |  | 0.47674 | 0 | 0.46496 | 0
| FDN0.1-13 | FDN | -13 | 248.1 | 0.1 | °F | °F |  |  |  | 0.41715 | 0 | 0.40364 | 0
| FDN0.1248 | FDN | 248 | 770.1 | 0.1 | °F | °F |  |  |  | 0.3915 | 0 | 0.38322 | 0
| FDN0.1770 | FDN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
| FDN0.12372 | FDN | 2372 | 2372.1 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
| FDN1-328 | FDN | -328 | -147.9 | 1 | °F | °F |  |  |  | 1.0446 | 0 | 1.014 | 0
| FDN1-148 | FDN | -148 | -12.9 | 1 | °F | °F |  |  |  | 0.76848 | 0 | 0.73904 | 0
| FDN1-13 | FDN | -13 | 248.1 | 1 | °F | °F |  |  |  | 0.73301 | 0 | 0.70208 | 0
| FDN1248 | FDN | 248 | 770.1 | 1 | °F | °F |  |  |  | 0.70888 | 0 | 0.69055 | 0
| FDN1770 | FDN | 770 | 2372 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
| FDN12372 | FDN | 2372 | 2372.1 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
| FDR0.132 | FDR | 32 | 482.1 | 0.1 | °F | °F |  |  |  | 1.1966 | 0 | 1.19 | 0
| FDR0.1482 | FDR | 482 | 752.1 | 0.1 | °F | °F |  |  |  | 0.74217 | 0 | 0.73324 | 0
| FDR0.1752 | FDR | 752 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.70146 | 0 | 0.69201 | 0
| FDR0.11832 | FDR | 1832 | 3213 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
| FDR0.13213 | FDR | 3213 | 3213.1 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
| FDR132 | FDR | 32 | 482.1 | 1 | °F | °F |  |  |  | 1.3448 | 0 | 1.3214 | 0
| FDR1482 | FDR | 482 | 752.1 | 1 | °F | °F |  |  |  | 0.95934 | 0 | 0.93147 | 0
| FDR1752 | FDR | 752 | 1832.1 | 1 | °F | °F |  |  |  | 0.92821 | 0 | 0.89938 | 0
| FDR11832 | FDR | 1832 | 3213 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
| FDR13213 | FDR | 3213 | 3213.1 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
| FDS0.132 | FDS | 32 | 482.1 | 0.1 | °F | °F |  |  |  | 0.99102 | 0 | 0.98329 | 0
| FDS0.1482 | FDS | 482 | 1832.1 | 0.1 | °F | °F |  |  |  | 0.76311 | 0 | 0.75443 | 0
| FDS0.11832 | FDS | 1832 | 2552.1 | 0.1 | °F | °F |  |  |  | 0.78274 | 0 | 0.77428 | 0
| FDS0.12552 | FDS | 2552 | 3213 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
| FDS0.13213 | FDS | 3213 | 3213.1 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
| FDS132 | FDS | 32 | 482.1 | 1 | °F | °F |  |  |  | 1.1653 | 0 | 1.1388 | 0
| FDS1482 | FDS | 482 | 1832.1 | 1 | °F | °F |  |  |  | 0.97706 | 0 | 0.94824 | 0
| FDS11832 | FDS | 1832 | 2552.1 | 1 | °F | °F |  |  |  | 0.99106 | 0 | 0.96411 | 0
| FDS12552 | FDS | 2552 | 3213 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
| FDS13213 | FDS | 3213 | 3213.1 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
| FDT0.1-418 | FDT | -418 | -237.9 | 0.1 | °F | °F |  |  |  | 0.55536 | 0 | 0.55058 | 0
| FDT0.1-238 | FDT | -238 | 32.1 | 0.1 | °F | °F |  |  |  | 0.29308 | 0 | 0.28549 | 0
| FDT0.132 | FDT | 32 | 248.1 | 0.1 | °F | °F |  |  |  | 0.20015 | 0 | 0.19651 | 0
| FDT0.1248 | FDT | 248 | 752 | 0.1 | °F | °F |  |  |  | 0.17895 | 0 | 0.17487 | 0
| FDT0.1752 | FDT | 752 | 752.1 | 0.1 | °F | °F |  |  |  | 0.17895 | 0 | 0.17487 | 0
| FDT1-418 | FDT | -418 | -237.9 | 1 | °F | °F |  |  |  | 0.8089 | 0 | 0.7957 | 0
| FDT1-238 | FDT | -238 | 32.1 | 1 | °F | °F |  |  |  | 0.65502 | 0 | 0.64149 | 0
| FDT132 | FDT | 32 | 248.1 | 1 | °F | °F |  |  |  | 0.61188 | 0 | 0.60714 | 0
| FDT1248 | FDT | 248 | 752 | 1 | °F | °F |  |  |  | 0.60528 | 0 | 0.60048 | 0
| FDT1752 | FDT | 752 | 752.1 | 1 | °F | °F |  |  |  | 0.60528 | 0 | 0.60048 | 0
| FDPT1000.1-328 | FDPT100 | -328 | 32.1 | 0.1 | °F | °F |  |  |  | 0.21928 | 0 | 0.19044 | 0
| FDPT1000.132 | FDPT100 | 32 | 212.1 | 0.1 | °F | °F |  |  |  | 0.28537 | 0 | 0.26385 | 0
| FDPT1000.1212 | FDPT100 | 212 | 572.1 | 0.1 | °F | °F |  |  |  | 0.28537 | 0 | 0.26385 | 0
| FDPT1000.1572 | FDPT100 | 572 | 752.1 | 0.1 | °F | °F |  |  |  | 0.39012 | 0 | 0.37466 | 0
| FDPT1000.1752 | FDPT100 | 752 | 1166.1 | 0.1 | °F | °F |  |  |  | 0.45987 | 0 | 0.44683 | 0
| FDPT1000.11166 | FDPT100 | 1166 | 1472 | 0.1 | °F | °F |  |  |  | 0.77811 | 0 | 0.77048 | 0
| FDPT1000.11472 | FDPT100 | 1472 | 1472.1 | 0.1 | °F | °F |  |  |  | 0.77811 | 0 | 0.77048 | 0
| FDPT1001-328 | FDPT100 | -328 | 32.1 | 1 | °F | °F |  |  |  | 0.60123 | 0 | 0.58897 | 0
| FDPT100132 | FDPT100 | 32 | 212.1 | 1 | °F | °F |  |  |  | 0.61101 | 0 | 0.59895 | 0
| FDPT1001212 | FDPT100 | 212 | 572.1 | 1 | °F | °F |  |  |  | 0.61101 | 0 | 0.59895 | 0
| FDPT1001572 | FDPT100 | 572 | 752.1 | 1 | °F | °F |  |  |  | 0.63118 | 0 | 0.61951 | 0
| FDPT1001752 | FDPT100 | 752 | 1166.1 | 1 | °F | °F |  |  |  | 0.64673 | 0 | 0.63534 | 0
| FDPT10011166 | FDPT100 | 1166 | 1472 | 1 | °F | °F |  |  |  | 0.739 | 0 | 0.72906 | 0
| FDPT10011472 | FDPT100 | 1472 | 1472.1 | 1 | °F | °F |  |  |  | 0.739 | 0 | 0.72906 | 0
| FDPT10000.1-328 | FDPT1000 | -328 | 32.1 | 0.1 | °F | °F |  |  |  | 0.16401 | 0 | 0.1228 | 0
| FDPT10000.132 | FDPT1000 | 32 | 212.1 | 0.1 | °F | °F |  |  |  | 0.19201 | 0 | 0.15827 | 0
| FDPT10000.1212 | FDPT1000 | 212 | 500.1 | 0.1 | °F | °F |  |  |  | 0.19201 | 0 | 0.15827 | 0
| FDPT10000.1500 | FDPT1000 | 500 | 572.1 | 0.1 | °F | °F |  |  |  | 0.2619 | 0 | 0.23826 | 0
| FDPT10000.1572 | FDPT1000 | 572 | 752.1 | 0.1 | °F | °F |  |  |  | 0.29342 | 0 | 0.27254 | 0
| FDPT10000.1752 | FDPT1000 | 752 | 1112.1 | 0.1 | °F | °F |  |  |  | 0.29725 | 0 | 0.27666 | 0
| FDPT10000.11112 | FDPT1000 | 1112 | 1166 | 0.1 | °F | °F |  |  |  | 0.84699 | 0 | 0.83998 | 0
| FDPT10000.11166 | FDPT1000 | 1166 | 1166.1 | 0.1 | °F | °F |  |  |  | 0.84699 | 0 | 0.83998 | 0
| FDPT10001-328 | FDPT1000 | -328 | 32.1 | 1 | °F | °F |  |  |  | 0.59551 | 0 | 0.58313 | 0
| FDPT10001212 | FDPT1000 | 212 | 500.1 | 1 | °F | °F |  |  |  | 0.59861 | 0 | 0.58629 | 0
| FDPT10001500 | FDPT1000 | 500 | 572.1 | 1 | °F | °F |  |  |  | 0.61003 | 0 | 0.59795 | 0
| FDPT10001572 | FDPT1000 | 572 | 752.1 | 1 | °F | °F |  |  |  | 0.61481 | 0 | 0.60282 | 0
| FDPT10001752 | FDPT1000 | 752 | 1112.1 | 1 | °F | °F |  |  |  | 0.61665 | 0 | 0.6047 | 0
| FDPT100011112 | FDPT1000 | 1112 | 1166 | 1 | °F | °F |  |  |  | 1.0262 | 0 | 1.0191 | 0
| FDPT100011166 | FDPT1000 | 1166 | 1166.1 | 1 | °F | °F |  |  |  | 1.0262 | 0 | 1.0191 | 0
| DCW0.00000000010 | DCW | 0 | 0.000109 | 1e-10 | WDC | WDC |  |  |  | 0 | 0.0036 | 0 | 0
| DCW0.0000000010 | DCW | 0 | 0.000109 | 1e-09 | WDC | WDC |  |  |  | 0 | 0.0036 | 0 | 0
| DCW0.000000010 | DCW | 0 | 0.000109 | 1e-08 | WDC | WDC |  |  |  | 0 | 0.0068024 | 0 | 0
| DCW0.00000010 | DCW | 0 | 0.000109 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.05811 | 0 | 0
| DCW0.0000010 | DCW | 0 | 0.000109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.58001 | 0 | 0
| DCW0.000010 | DCW | 0 | 0.000109 | 1e-05 | WDC | WDC |  |  |  | 0 | 5.8 | 0 | 0
| DCW0.00010 | DCW | 0 | 0.000109 | 0.0001 | WDC | WDC |  |  |  | 0 | 58 | 0 | 0
| DCW0.0010 | DCW | 0 | 0.000109 | 0.001 | WDC | WDC |  |  |  | 0 | 580 | 0 | 0
| DCW0.010 | DCW | 0 | 0.000109 | 0.01 | WDC | WDC |  |  |  | 0 | 5800 | 0 | 0
| DCW0.000000010.000109 | DCW | 0.000109 | 1.0901e-06 | 1e-08 | WDC | WDC |  |  |  | 0 | 0.00035 | 0 | 0
| DCW0.00000010.000109 | DCW | 0.000109 | 1.0901e-06 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.00063401 | 0 | 0
| DCW0.0000010.000109 | DCW | 0.000109 | 1.0901e-06 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.0053337 | 0 | 0
| DCW0.000010.000109 | DCW | 0.000109 | 1.0901e-06 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.053213 | 0 | 0
| DCW0.00010.000109 | DCW | 0.000109 | 1.0901e-06 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.53211 | 0 | 0
| DCW0.00000010.00109 | DCW | 0.00109 | 0.0109 | 1e-07 | WDC | WDC |  |  |  | 0 | 0.00024 | 0 | 0
| DCW0.0000010.00109 | DCW | 0.00109 | 0.0109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.00058093 | 0 | 0
| DCW0.000010.00109 | DCW | 0.00109 | 0.0109 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.0053277 | 0 | 0
| DCW0.00010.00109 | DCW | 0.00109 | 0.0109 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.053212 | 0 | 0
| DCW0.0010.00109 | DCW | 0.00109 | 0.0109 | 0.001 | WDC | WDC |  |  |  | 0 | 0.53211 | 0 | 0
| DCW0.0000010.0109 | DCW | 0.0109 | 0.109 | 1e-06 | WDC | WDC |  |  |  | 0 | 0.00019 | 0 | 0
| DCW0.000010.0109 | DCW | 0.0109 | 0.109 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.00056395 | 0 | 0
| DCW0.00010.0109 | DCW | 0.0109 | 0.109 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.0053308 | 0 | 0
| DCW0.0010.0109 | DCW | 0.0109 | 0.109 | 0.001 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
| DCW0.010.0109 | DCW | 0.0109 | 0.109 | 0.01 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
| DCW0.000010.109 | DCW | 0.109 | 1.0901 | 1e-05 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
| DCW0.00010.109 | DCW | 0.109 | 1.0901 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.00057856 | 0 | 0
| DCW0.0010.109 | DCW | 0.109 | 1.0901 | 0.001 | WDC | WDC |  |  |  | 0 | 0.0053324 | 0 | 0
| DCW0.010.109 | DCW | 0.109 | 1.0901 | 0.01 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
| DCW0.10.109 | DCW | 0.109 | 1.0901 | 0.1 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
| DCW0.00011.09 | DCW | 1.09 | 10.9001 | 0.0001 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
| DCW0.0011.09 | DCW | 1.09 | 10.9001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00057825 | 0 | 0
| DCW0.011.09 | DCW | 1.09 | 10.9001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.0053322 | 0 | 0
| DCW0.11.09 | DCW | 1.09 | 10.9001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
| DCW11.09 | DCW | 1.09 | 10.9001 | 1 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
| DCW0.00110.9 | DCW | 10.9 | 109.0001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00023 | 0 | 0
| DCW0.0110.9 | DCW | 10.9 | 109.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.00057887 | 0
| DCW0.110.9 | DCW | 10.9 | 109.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.0053323 | 0 | 0
| DCW110.9 | DCW | 10.9 | 109.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.053261 | 0 | 0
| DCW1010.9 | DCW | 10.9 | 109.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.5326 | 0 | 0
| DCW0.001109 | DCW | 109 | 330.0001 | 0.001 | WDC | WDC |  |  |  | 0 | 0.00015 | 0 | 0
| DCW0.01109 | DCW | 109 | 330.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.00015255 | 0 | 0
| DCW0.1109 | DCW | 109 | 330.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.00055202 | 0 | 0
| DCW1109 | DCW | 109 | 330.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.0053282 | 0 | 0
| DCW10109 | DCW | 109 | 330.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.05326 | 0 | 0
| DCW0.01330 | DCW | 330 | 3000.0001 | 0.01 | WDC | WDC |  |  |  | 0 | 0.0006 | 0 | 0
| DCW0.1330 | DCW | 330 | 3000.0001 | 0.1 | WDC | WDC |  |  |  | 0 | 0.00062246 | 0 | 0
| DCW1330 | DCW | 330 | 3000.0001 | 1 | WDC | WDC |  |  |  | 0 | 0.0018565 | 0 | 0
| DCW10330 | DCW | 330 | 3000.0001 | 10 | WDC | WDC |  |  |  | 0 | 0.017586 | 0 | 0
| DCW100330 | DCW | 330 | 3000.0001 | 100 | WDC | WDC |  |  |  | 0 | 0.17576 | 0 | 0
| DCW103000 | DCW | 3000 | 11000.1 | 10 | WDC | WDC |  |  |  | 0 | 0.00091 | 0 | 0
| DCW1003000 | DCW | 3000 | 11000.1 | 100 | WDC | WDC |  |  |  | 0 | 0.00091 | 0 | 0
| DCW10003000 | DCW | 3000 | 11000.1 | 1000 | WDC | WDC |  |  |  | 0 | 0.00092998 | 0 | 0
| DCW100003000 | DCW | 3000 | 11000.1 | 10000 | WDC | WDC |  |  |  | 0 | 0.0021367 | 0 | 0
| DCW1000003000 | DCW | 3000 | 11000.1 | 100000 | WDC | WDC |  |  |  | 0 | 0.019355 | 0 | 0
| DCW1011000 | DCW | 11000 | 24500 | 10 | WDC | WDC |  |  |  | 0 | 0.0014 | 0 | 0
| DCW10011000 | DCW | 11000 | 24500 | 100 | WDC | WDC |  |  |  | 0 | 0.0014 | 0 | 0
| DCW100011000 | DCW | 11000 | 24500 | 1000 | WDC | WDC |  |  |  | 0 | 0.0014429 | 0 | 0
| DCW1000011000 | DCW | 11000 | 24500 | 10000 | WDC | WDC |  |  |  | 0 | 0.0054411 | 0 | 0
| DCW10000011000 | DCW | 11000 | 24500 | 100000 | WDC | WDC |  |  |  | 0 | 0.052745 | 0 | 0
| DCW1024500 | DCW | 24500 | 24501 | 10 | WAC | WAC |  |  |  | 0 | 0.0014 | 0 | 0
| DCW10024500 | DCW | 24500 | 24501 | 100 | WAC | WAC |  |  |  | 0 | 0.0014 | 0 | 0
| DCW100024500 | DCW | 24500 | 24501 | 1000 | WAC | WAC |  |  |  | 0 | 0.0014429 | 0 | 0
| DCW1000024500 | DCW | 24500 | 24501 | 10000 | WAC | WAC |  |  |  | 0 | 0.0054411 | 0 | 0
| DCW10000024500 | DCW | 24500 | 24501 | 100000 | WAC | WAC |  |  |  | 0 | 0.052745 | 0 | 0
| ACW0.00000000010 | ACW | 0 | 1.09e-05 | 1e-10 | WAC | WAC |  |  |  | 0 | 0.025 | 0 | 0
| ACW0.0000000010 | ACW | 0 | 1.09e-05 | 1e-09 | WAC | WAC |  |  |  | 0 | 0.025178 | 0 | 0
| ACW0.000000010 | ACW | 0 | 1.09e-05 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.062989 | 0 | 0
| ACW0.00000010 | ACW | 0 | 1.09e-05 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.58053 | 0 | 0
| ACW0.0000010 | ACW | 0 | 1.09e-05 | 1e-06 | WAC | WAC |  |  |  | 0 | 5.8001 | 0 | 0
| ACW0.0000000010.0000109 | ACW | 1.09e-05 | 0.000109 | 1e-09 | WAC | WAC |  |  |  | 0 | 0.0026 | 0 | 0
| ACW0.000000010.0000109 | ACW | 1.09e-05 | 0.000109 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.0026 | 0 | 0
| ACW0.00000010.0000109 | ACW | 1.09e-05 | 0.000109 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0058884 | 0 | 0
| ACW0.0000010.0000109 | ACW | 1.09e-05 | 0.000109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.05332 | 0 | 0
| ACW0.000010.0000109 | ACW | 1.09e-05 | 0.000109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.53261 | 0 | 0
| ACW0.000000010.000109 | ACW | 0.000109 | 0.0010901 | 1e-08 | WAC | WAC |  |  |  | 0 | 0.0017 | 0 | 0
| ACW0.00000010.000109 | ACW | 0.000109 | 0.0010901 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0017 | 0 | 0
| ACW0.0000010.000109 | ACW | 0.000109 | 0.0010901 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.0055645 | 0 | 0
| ACW0.000010.000109 | ACW | 0.000109 | 0.0010901 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.053285 | 0 | 0
| ACW0.00010.000109 | ACW | 0.000109 | 0.0010901 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
| ACW0.00000010.00109 | ACW | 0.00109 | 0.0109 | 1e-07 | WAC | WAC |  |  |  | 0 | 0.0013 | 0 | 0
| ACW0.0000010.00109 | ACW | 0.00109 | 0.0109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.0013483 | 0 | 0
| ACW0.000010.00109 | ACW | 0.00109 | 0.0109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.0054702 | 0 | 0
| ACW0.00010.00109 | ACW | 0.00109 | 0.0109 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.053275 | 0 | 0
| ACW0.0010.00109 | ACW | 0.00109 | 0.0109 | 0.001 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
| ACW0.0000010.0109 | ACW | 0.0109 | 0.109 | 1e-06 | WAC | WAC |  |  |  | 0 | 0.00078 | 0 | 0
| ACW0.000010.0109 | ACW | 0.0109 | 0.109 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.00094316 | 0 | 0
| ACW0.00010.0109 | ACW | 0.0109 | 0.109 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.0053847 | 0 | 0
| ACW0.0010.0109 | ACW | 0.0109 | 0.109 | 0.001 | WAC | WAC |  |  |  | 0 | 0.053267 | 0 | 0
| ACW0.010.0109 | ACW | 0.0109 | 0.109 | 0.01 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
| ACW0.000010.109 | ACW | 0.109 | 1.0901 | 1e-05 | WAC | WAC |  |  |  | 0 | 0.00078 | 0 | 0
| ACW0.00010.109 | ACW | 0.109 | 1.0901 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.00094283 | 0 | 0
| ACW0.0010.109 | ACW | 0.109 | 1.0901 | 0.001 | WAC | WAC |  |  |  | 0 | 0.0053846 | 0 | 0
| ACW0.010.109 | ACW | 0.109 | 1.0901 | 0.01 | WAC | WAC |  |  |  | 0 | 0.053267 | 0 | 0
| ACW0.10.109 | ACW | 0.109 | 1.0901 | 0.1 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
| ACW0.00011.09 | ACW | 1.09 | 10.9001 | 0.0001 | WAC | WAC |  |  |  | 0 | 0.0013 | 0 | 0
| ACW0.0011.09 | ACW | 1.09 | 10.9001 | 0.001 | WAC | WAC |  |  |  | 0 | 0.0013347 | 0 | 0
| ACW0.011.09 | ACW | 1.09 | 10.9001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0054669 | 0 | 0
| ACW0.11.09 | ACW | 1.09 | 10.9001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.053275 | 0 | 0
| ACW11.09 | ACW | 1.09 | 10.9001 | 1 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
| ACW0.00110.9 | ACW | 10.9 | 109.0001 | 0.001 | WAC | WAC |  |  |  | 0 | 0.00088 | 0 | 0
| ACW0.0110.9 | ACW | 10.9 | 109.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0010256 | 0 | 0
| ACW0.110.9 | ACW | 10.9 | 109.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0053997 | 0 | 0
| ACW110.9 | ACW | 10.9 | 109.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.053268 | 0 | 0
| ACW1010.9 | ACW | 10.9 | 109.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.5326 | 0 | 0
| ACW0.01109 | ACW | 109 | 363.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.0012 | 0 | 0
| ACW0.1109 | ACW | 109 | 363.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0012284 | 0 | 0
| ACW1109 | ACW | 109 | 363.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.0054419 | 0 | 0
| ACW10109 | ACW | 109 | 363.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.053273 | 0 | 0
| ACW0.01363 | ACW | 363 | 990.0001 | 0.01 | WAC | WAC |  |  |  | 0 | 0.001 | 0 | 0
| ACW0.1363 | ACW | 363 | 990.0001 | 0.1 | WAC | WAC |  |  |  | 0 | 0.001007 | 0 | 0
| ACW1363 | ACW | 363 | 990.0001 | 1 | WAC | WAC |  |  |  | 0 | 0.0018832 | 0 | 0
| ACW10363 | ACW | 363 | 990.0001 | 10 | WAC | WAC |  |  |  | 0 | 0.01601 | 0 | 0
| ACW100363 | ACW | 363 | 990.0001 | 100 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
| ACW0.1990 | ACW | 990 | 3630.1 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
| ACW1990 | ACW | 990 | 3630.1 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
| ACW10990 | ACW | 990 | 3630.1 | 10 | WAC | WAC |  |  |  | 0 | 0.0061979 | 0 | 0
| ACW100990 | ACW | 990 | 3630.1 | 100 | WAC | WAC |  |  |  | 0 | 0.058622 | 0 | 0
| ACW1000990 | ACW | 990 | 3630.1 | 1000 | WAC | WAC |  |  |  | 0 | 0.58586 | 0 | 0
| ACW0.13630 | ACW | 3630 | 11200.1 | 0.1 | WAC | WAC |  |  |  | 0 | 0.0016 | 0 | 0
| ACW13630 | ACW | 3630 | 11200.1 | 1 | WAC | WAC |  |  |  | 0 | 0.0016 | 0 | 0
| ACW103630 | ACW | 3630 | 11200.1 | 10 | WAC | WAC |  |  |  | 0 | 0.0022012 | 0 | 0
| ACW1003630 | ACW | 3630 | 11200.1 | 100 | WAC | WAC |  |  |  | 0 | 0.016051 | 0 | 0
| ACW10003630 | ACW | 3630 | 11200.1 | 1000 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
| ACW0.111200 | ACW | 11200 | 20900 | 0.1 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
| ACW111200 | ACW | 11200 | 20900 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
| ACW1011200 | ACW | 11200 | 20900 | 10 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
| ACW10011200 | ACW | 11200 | 20900 | 100 | WAC | WAC |  |  |  | 0 | 0.0055466 | 0 | 0
| ACW100011200 | ACW | 11200 | 20900 | 1000 | WAC | WAC |  |  |  | 0 | 0.051733 | 0 | 0
| ACW0.120900 | ACW | 20900 | 20901 | 0.1 | WAC | WAC |  |  |  | 0 | 0.15979 | 0 | 0
| ACW120900 | ACW | 20900 | 20901 | 1 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
| ACW1020900 | ACW | 20900 | 20901 | 10 | WAC | WAC |  |  |  | 0 | 0.0021 | 0 | 0
| ACW10020900 | ACW | 20900 | 20901 | 100 | WAC | WAC |  |  |  | 0 | 0.0055466 | 0 | 0
| ACW100020900 | ACW | 20900 | 20901 | 1000 | WAC | WAC |  |  |  | 0 | 0.051733 | 0 | 0

## Standard2
Updated: | 2021-11-18 | 3458A
|  | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.000000010 | DCV | 0 | 0.1 | 1e-08 | V | µV |  |  |  | 0.62 | 8 | 0.60546 | 8.1092
| DCV0.00000010 | DCV | 0 | 0.1 | 1e-07 | V | µV |  |  |  | 0.62 | 8.0238 | 0.60819 | 8.0937
| DCV0.0000010 | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 0.83938 | 6.9596 | 0.83737 | 6.9316
| DCV0.000010 | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 5.8321 | 1.4327 | 5.831 | 1.3947
| DCV0.00010 | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.004 | 0.15763 | 58.003 | 0.14188
| DCV0.0010 | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.020493 | 580 | 0.014191
| DCV0.010 | DCV | 0 | 0.1 | 0.01 | V | mV |  |  |  | 5.8 | 3.94e-06 | 5.7999 | 1.4191e-06
| DCV0.10 | DCV | 0 | 0.1 | 0.1 | V | mV |  |  |  | 58 | 1.1503e-06 | 58 | 1.4191e-07
| DCV10 | DCV | 0 | 0.1 | 1 | V | mV |  |  |  | 580 | 4.1754e-07 | 580 | 1.4191e-08
| DCV0.000000010.1 | DCV | 0.1 | 1 | 1e-08 | V | µV |  |  |  | 0.44 | 6 | 0.43733 | 5.9926
| DCV0.00000010.1 | DCV | 0.1 | 1 | 1e-07 | V | µV |  |  |  | 0.44094 | 5.99 | 0.43909 | 5.9911
| DCV0.0000010.1 | DCV | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 0.60901 | 5.85 | 0.60244 | 5.8536
| DCV0.000010.1 | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 5.589 | 3.0792 | 5.5845 | 3.0749
| DCV0.00010.1 | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 57.973 | 0.3881 | 57.971 | 0.38452
| DCV0.0010.1 | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 580 | 0.040018 | 579.96 | 0.038571
| DCV0.00000011 | DCV | 1 | 10 | 1e-07 | V | µV |  |  |  | 1.2 | 5.3 | 0.75905 | 5.3437
| DCV0.0000011 | DCV | 1 | 10 | 1e-06 | V | µV |  |  |  | 1.2 | 5.3409 | 0.78898 | 5.341
| DCV0.000011 | DCV | 1 | 10 | 1e-05 | V | µV |  |  |  | 3.301 | 5.1205 | 3.2984 | 5.1207
| DCV0.00011 | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 55.981 | 2.3401 | 55.98 | 2.34
| DCV0.0011 | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 579.76 | 0.27724 | 579.76 | 0.27716
| DCV0.011 | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5800 | 0.027808 | 5799.7 | 0.027777
| DCV0.00000110 | DCV | 10 | 100 | 1e-06 | V | µV |  |  |  | 34 | 7.8 | 33.132 | 7.7866
| DCV0.0000110 | DCV | 10 | 100 | 1e-05 | V | µV |  |  |  | 34 | 7.7848 | 33.3 | 7.7851
| DCV0.000110 | DCV | 10 | 100 | 0.0001 | V | µV |  |  |  | 48.99 | 7.6505 | 48.723 | 7.6514
| DCV0.00110 | DCV | 10 | 100 | 0.001 | V | µV |  |  |  | 545.46 | 4.5283 | 545.28 | 4.5242
| DCV0.0110 | DCV | 10 | 100 | 0.01 | V | µV |  |  |  | 5795 | 0.61987 | 5794.9 | 0.61637
| DCV0.00001100 | DCV | 100 | 1000 | 1e-05 | V | mV |  |  |  | 0.11 | 0.0087 | 0.10064 | 0.0086547
| DCV0.0001100 | DCV | 100 | 1000 | 0.0001 | V | mV |  |  |  | 0.11 | 0.0086529 | 0.10253 | 0.008653
| DCV0.001100 | DCV | 100 | 1000 | 0.001 | V | mV |  |  |  | 0.27989 | 0.0084976 | 0.27707 | 0.0084975
| DCV0.01100 | DCV | 100 | 1000 | 0.01 | V | mV |  |  |  | 5.3677 | 0.0051444 | 5.3663 | 0.0051359
| DCV0.1100 | DCV | 100 | 1000 | 0.1 | V | mV |  |  |  | 57.936 | 0.00072803 | 57.936 | 0.00072118
| DCV1100 | DCV | 100 | 1000 | 1 | V | mV |  |  |  | 579.99 | 7.53e-05 | 579.93 | 7.2528e-05
| DCA0.0000000000010 | DCA | 0 | 1e-07 | 1e-12 | A | nA |  |  |  | 0.048 | 34000 | 0.046246 | 34649
| DCA0.000000000010 | DCA | 0 | 1e-07 | 1e-11 | A | nA |  |  |  | 0.048 | 34000 | 0.046601 | 34403
| DCA0.00000000010 | DCA | 0 | 1e-07 | 1e-10 | A | nA |  |  |  | 0.074257 | 24027 | 0.07397 | 22146
| DCA0.0000000010 | DCA | 0 | 1e-07 | 1e-09 | A | nA |  |  |  | 0.57935 | 3897.9 | 0.5792 | 2870.4
| DCA0.000000010 | DCA | 0 | 1e-07 | 1e-08 | A | nA |  |  |  | 5.7738 | 700.9 | 5.7737 | 288.02
| DCA0.00000010 | DCA | 0 | 1e-07 | 1e-07 | A | nA |  |  |  | 57.736 | 193.97 | 57.736 | 28.803
| DCA0.0000010 | DCA | 0 | 1e-07 | 1e-06 | A | nA |  |  |  | 577.36 | 68.945 | 577.36 | 2.8803
| DCA0.000010 | DCA | 0 | 1e-07 | 1e-05 | A | nA |  |  |  | 5773.6 | 26.714 | 5773.6 | 0.28803
| DCA0.00010 | DCA | 0 | 1e-07 | 0.0001 | A | nA |  |  |  | 57736 | 10.599 | 57736 | 0.028736
| DCA0.0010 | DCA | 0 | 1e-07 | 0.001 | A | nA |  |  |  | 577360 | 4.2309 | 577360 | 0
| DCA0.010 | DCA | 0 | 1e-07 | 0.01 | A | nA |  |  |  | 5773600 | 1.6933 | 5773600 | 0
| DCA0.10 | DCA | 0 | 1e-07 | 0.1 | A | nA |  |  |  | 57736000 | 0.75258 | 57736000 | 0
| DCA0.0000000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-12 | A | nA |  |  |  | 0.048 | 23000 | 0.046073 | 24826
| DCA0.000000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-11 | A | nA |  |  |  | 0.048 | 23271 | 0.046424 | 24708
| DCA0.00000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-10 | A | nA |  |  |  | 0.074091 | 17690 | 0.07366 | 17772
| DCA0.0000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-09 | A | nA |  |  |  | 0.57936 | 2553.2 | 0.57914 | 2554.1
| DCA0.000000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-08 | A | nA |  |  |  | 5.7738 | 256.81 | 5.7737 | 256.81
| DCA0.00000010.0000001 | DCA | 1e-07 | 1e-06 | 1e-07 | A | nA |  |  |  | 57.736 | 25.682 | 57.736 | 25.682
| DCA0.0000000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-12 | A | nA |  |  |  | 0.13 | 25000 | 0.11375 | 25625
| DCA0.000000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-11 | A | nA |  |  |  | 0.13 | 25119 | 0.11389 | 25617
| DCA0.00000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-10 | A | nA |  |  |  | 0.13 | 24832 | 0.12602 | 24846
| DCA0.0000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-09 | A | nA |  |  |  | 0.58396 | 10197 | 0.58374 | 10200
| DCA0.000000010.000001 | DCA | 1e-06 | 1e-05 | 1e-08 | A | nA |  |  |  | 5.7742 | 1129.1 | 5.7741 | 1129.1
| DCA0.00000010.000001 | DCA | 1e-06 | 1e-05 | 1e-07 | A | nA |  |  |  | 57.736 | 113.04 | 57.735 | 113.04
| DCA0.000000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-11 | A | nA |  |  |  | 0.16 | 23000 | 0.11411 | 23000
| DCA0.00000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-10 | A | nA |  |  |  | 0.16 | 23403 | 0.11928 | 23403
| DCA0.0000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-09 | A | nA |  |  |  | 0.47124 | 20574 | 0.46866 | 20574
| DCA0.000000010.00001 | DCA | 1e-05 | 0.0001 | 1e-08 | A | nA |  |  |  | 5.7304 | 5486.9 | 5.7294 | 5486.9
| DCA0.00000010.00001 | DCA | 1e-05 | 0.0001 | 1e-07 | A | nA |  |  |  | 57.731 | 581.51 | 57.731 | 581.51
| DCA0.0000010.00001 | DCA | 1e-05 | 0.0001 | 1e-06 | A | nA |  |  |  | 577.35 | 61.585 | 577.35 | 61.585
| DCA0.00000000010.0001 | DCA | 0.0001 | 0.001 | 1e-10 | A | nA |  |  |  | 6.3 | 23000 | 5.7456 | 23552
| DCA0.0000000010.0001 | DCA | 0.0001 | 0.001 | 1e-09 | A | nA |  |  |  | 6.3 | 23340 | 5.7675 | 23536
| DCA0.000000010.0001 | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 7.7388 | 22133 | 7.7351 | 22126
| DCA0.00000010.0001 | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 57.586 | 7177.2 | 57.585 | 7158.5
| DCA0.0000010.0001 | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 577.34 | 770.78 | 577.34 | 762.27
| DCA0.000010.0001 | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.6 | 79.686 | 5773.5 | 76.28
| DCA0.0000000010.001 | DCA | 0.001 | 0.01 | 1e-09 | A | µA |  |  |  | 0.064 | 23 | 0.057395 | 23.593
| DCA0.000000010.001 | DCA | 0.001 | 0.01 | 1e-08 | A | µA |  |  |  | 0.064 | 23.49 | 0.057621 | 23.576
| DCA0.00000010.001 | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 0.07733 | 22.173 | 0.077294 | 22.166
| DCA0.0000010.001 | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 0.57584 | 7.1949 | 0.57583 | 7.1763
| DCA0.000010.001 | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.7734 | 0.77277 | 5.7734 | 0.76426
| DCA0.00010.001 | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.736 | 0.079885 | 57.735 | 0.076479
| DCA0.000000010.01 | DCA | 0.01 | 0.1 | 1e-08 | A | µA |  |  |  | 0.67 | 40 | 0.57451 | 40.858
| DCA0.00000010.01 | DCA | 0.01 | 0.1 | 1e-07 | A | µA |  |  |  | 0.67 | 40.354 | 0.57629 | 40.844
| DCA0.0000010.01 | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 0.75326 | 39.474 | 0.74494 | 39.51
| DCA0.000010.01 | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 5.6886 | 17.431 | 5.683 | 17.368
| DCA0.00010.01 | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 57.726 | 2.0334 | 57.724 | 1.9935
| DCA0.0010.01 | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.35 | 0.21573 | 577.34 | 0.19968
| DCA0.00000010.1 | DCA | 0.1 | 1 | 1e-07 | A | µA |  |  |  | 12 | 130 | 11.536 | 127.22
| DCA0.0000010.1 | DCA | 0.1 | 1 | 1e-06 | A | µA |  |  |  | 12 | 130 | 11.548 | 127.21
| DCA0.000010.1 | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 12 | 130 | 12.277 | 126.6
| DCA0.00010.1 | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 52.949 | 97.342 | 52.884 | 97.405
| DCA0.0010.1 | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 576.12 | 17.674 | 576.1 | 17.7
| DCA0.010.1 | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5773.4 | 1.7853 | 5771.8 | 1.7958
| DCA0.0000011 | DCA | 1 | 20 | 1e-06 | A | mA |  |  |  | 0.068 | 0.0036 | 0.043 | 0.0047621
| DCA0.000011 | DCA | 1 | 20 | 1e-05 | A | mA |  |  |  | 0.068 | 0.0036 | 0.043318 | 0.0047475
| DCA0.00011 | DCA | 1 | 20 | 0.0001 | A | mA |  |  |  | 0.070893 | 0.0038965 | 0.068928 | 0.0038258
| DCA0.0011 | DCA | 1 | 20 | 0.001 | A | mA |  |  |  | 0.57834 | 0.00081073 | 0.57809 | 0.0006992
| DCA0.011 | DCA | 1 | 20 | 0.01 | A | mA |  |  |  | 5.7739 | 0.00011677 | 5.7736 | 7.0888e-05
| DCA0.11 | DCA | 1 | 20 | 0.1 | A | mA |  |  |  | 57.736 | 2.5449e-05 | 57.735 | 7.0898e-06
| Resistance0.000000010 | Resistance | 0 | 0.01 | 1e-08 | kOhms | µOhms |  |  |  | 65 | 17000 | 61.668 | 17321
| Resistance0.00000010 | Resistance | 0 | 0.01 | 1e-07 | kOhms | µOhms |  |  |  | 81.734 | 16018 | 81.655 | 16015
| Resistance0.0000010 | Resistance | 0 | 0.01 | 1e-06 | kOhms | µOhms |  |  |  | 578.24 | 4520.7 | 578.2 | 4507.9
| Resistance0.000010 | Resistance | 0 | 0.01 | 1e-05 | kOhms | µOhms |  |  |  | 5773.6 | 476.27 | 5773.6 | 470.59
| Resistance0.00010 | Resistance | 0 | 0.01 | 0.0001 | kOhms | µOhms |  |  |  | 57736 | 49.352 | 57736 | 47.081
| Resistance0.0010 | Resistance | 0 | 0.01 | 0.001 | kOhms | µOhms |  |  |  | 577360 | 5.6165 | 577360 | 4.7081
| Resistance0.010 | Resistance | 0 | 0.01 | 0.01 | kOhms | µOhms |  |  |  | 5773600 | 0.83417 | 5773600 | 0.47081
| Resistance0.10 | Resistance | 0 | 0.01 | 0.1 | kOhms | µOhms |  |  |  | 57736000 | 0.19242 | 57736000 | 0.04708
| Resistance10 | Resistance | 0 | 0.01 | 1 | kOhms | µOhms |  |  |  | 577360000 | 0.06285 | 577360000 | 0.0047021
| Resistance100 | Resistance | 0 | 0.01 | 10 | kOhms | µOhms |  |  |  | 5773600000 | 0.023736 | 5773600000 | 0
| Resistance1000 | Resistance | 0 | 0.01 | 100 | kOhms | µOhms |  |  |  | 57736000000 | 0.0084771 | 57736000000 | 0
| Resistance10000 | Resistance | 0 | 0.01 | 1000 | kOhms | µOhms |  |  |  | 577360000000 | 0 | 577360000000 | 0
| Resistance0.000000010.01 | Resistance | 0.01 | 0.1 | 1e-08 | kOhms | µOhms |  |  |  | 580 | 14000 | 577.27
| Resistance0.00000010.01 | Resistance | 0.01 | 0.1 | 1e-07 | kOhms | µOhms |  |  |  | 580 | 13975 | 579.78 | 13975
| Resistance0.0000010.01 | Resistance | 0.01 | 0.1 | 1e-06 | kOhms | µOhms |  |  |  | 794.35 | 12648 | 794.24 | 12648
| Resistance0.000010.01 | Resistance | 0.01 | 0.1 | 1e-05 | kOhms | µOhms |  |  |  | 5786.4 | 3162.6 | 5786.3 | 3161.7
| Resistance0.00010.01 | Resistance | 0.01 | 0.1 | 0.0001 | kOhms | µOhms |  |  |  | 57737 | 326.69 | 57737 | 326.28
| Resistance0.0010.01 | Resistance | 0.01 | 0.1 | 0.001 | kOhms | µOhms |  |  |  | 577360 | 32.802 | 577350 | 32.639
| Resistance0.00000010.1 | Resistance | 0.1 | 1 | 1e-07 | kOhms | µOhms |  |  |  | 590 | 12000 | 568.01 | 11812
| Resistance0.0000010.1 | Resistance | 0.1 | 1 | 1e-06 | kOhms | µOhms |  |  |  | 647.66 | 11942 | 669.04 | 11724
| Resistance0.000010.1 | Resistance | 0.1 | 1 | 1e-05 | kOhms | µOhms |  |  |  | 5185.8 | 8475.5 | 5185 | 8474.8
| Resistance0.00010.1 | Resistance | 0.1 | 1 | 0.0001 | kOhms | µOhms |  |  |  | 57619 | 1429.9 | 57619 | 1428.7
| Resistance0.0010.1 | Resistance | 0.1 | 1 | 0.001 | kOhms | µOhms |  |  |  | 577340 | 145.01 | 577340 | 144.51
| Resistance0.010.1 | Resistance | 0.1 | 1 | 0.01 | kOhms | µOhms |  |  |  | 5773600 | 14.653 | 5773500 | 14.453
| Resistance0.0000011 | Resistance | 1 | 10 | 1e-06 | kOhms | mΩ |  |  |  | 5.9 | 12 | 5.8778 | 11.642
| Resistance0.000011 | Resistance | 1 | 10 | 1e-05 | kOhms | mΩ |  |  |  | 6.5111 | 11.939 | 6.8827 | 11.555
| Resistance0.00011 | Resistance | 1 | 10 | 0.0001 | kOhms | mΩ |  |  |  | 52.024 | 8.3239 | 52.009 | 8.3231
| Resistance0.0011 | Resistance | 1 | 10 | 0.001 | kOhms | mΩ |  |  |  | 576.23 | 1.3955 | 576.23 | 1.3939
| Resistance0.011 | Resistance | 1 | 10 | 0.01 | kOhms | mΩ |  |  |  | 5773.4 | 0.1416 | 5773.4 | 0.14095
| Resistance0.11 | Resistance | 1 | 10 | 0.1 | kOhms | mΩ |  |  |  | 57736 | 0.014357 | 57735 | 0.014097
| Resistance0.0000110 | Resistance | 10 | 100 | 1e-05 | kOhms | mΩ |  |  |  | 58 | 12 | 53.873 | 12.027
| Resistance0.000110 | Resistance | 10 | 100 | 0.0001 | kOhms | mΩ |  |  |  | 64.108 | 11.939 | 63.986 | 11.939
| Resistance0.00110 | Resistance | 10 | 100 | 0.001 | kOhms | mΩ |  |  |  | 516.51 | 8.6656 | 516.37 | 8.6649
| Resistance0.0110 | Resistance | 10 | 100 | 0.01 | kOhms | mΩ |  |  |  | 5761.5 | 1.4742 | 5761.5 | 1.4726
| Resistance0.110 | Resistance | 10 | 100 | 0.1 | kOhms | mΩ |  |  |  | 57734 | 0.14965 | 57734 | 0.149
| Resistance110 | Resistance | 10 | 100 | 1 | kOhms | mΩ |  |  |  | 577360 | 0.015162 | 577350 | 0.014902
| Resistance0.0001100 | Resistance | 100 | 1000 | 0.0001 | kOhms | Ω |  |  |  | 2.3 | 0.018 | 2.2687 | 0.017871
| Resistance0.001100 | Resistance | 100 | 1000 | 0.001 | kOhms | Ω |  |  |  | 2.3 | 0.018 | 2.3129 | 0.017835
| Resistance0.01100 | Resistance | 100 | 1000 | 0.01 | kOhms | Ω |  |  |  | 5.5163 | 0.015438 | 5.5116 | 0.015439
| Resistance0.1100 | Resistance | 100 | 1000 | 0.1 | kOhms | Ω |  |  |  | 57.516 | 0.0036361 | 57.515 | 0.0036327
| Resistance1100 | Resistance | 100 | 1000 | 1 | kOhms | Ω |  |  |  | 577.33 | 0.00037584 | 577.33 | 0.00037433
| Resistance10100 | Resistance | 100 | 1000 | 10 | kOhms | Ω |  |  |  | 5773.6 | 3.8048e-05 | 5773.5 | 3.7445e-05
| Resistance0.0011000 | Resistance | 1000 | 10000 | 0.001 | kOhms | Ω |  |  |  | 120 | 0.059 | 115.5 | 0.057739
| Resistance0.011000 | Resistance | 1000 | 10000 | 0.01 | kOhms | Ω |  |  |  | 120 | 0.059 | 115.61 | 0.057731
| Resistance0.11000 | Resistance | 1000 | 10000 | 0.1 | kOhms | Ω |  |  |  | 124.34 | 0.058567 | 125.65 | 0.056965
| Resistance11000 | Resistance | 1000 | 10000 | 1 | kOhms | Ω |  |  |  | 569.84 | 0.033313 | 569.55 | 0.033236
| Resistance101000 | Resistance | 1000 | 10000 | 10 | kOhms | Ω |  |  |  | 5771.9 | 0.0043708 | 5771.8 | 0.0043145
| Resistance1001000 | Resistance | 1000 | 10000 | 100 | kOhms | Ω |  |  |  | 57735 | 0.00045579 | 57731 | 0.00043308
| Resistance0.0110000 | Resistance | 10000 | 100000 | 0.01 | kOhms | kOhms |  |  |  | 1.2 | 0.00058 | 1.1508 | 0.0005778
| Resistance0.110000 | Resistance | 10000 | 100000 | 0.1 | kOhms | kOhms |  |  |  | 1.2 | 0.00058 | 1.1511 | 0.0005778
| Resistance110000 | Resistance | 10000 | 100000 | 1 | kOhms | kOhms |  |  |  | 1.2 | 0.00058 | 1.1777 | 0.00057756
| Resistance1010000 | Resistance | 10000 | 100000 | 10 | kOhms | kOhms |  |  |  | 3.4529 | 0.00055768 | 3.442 | 0.00055771
| Resistance10010000 | Resistance | 10000 | 100000 | 100 | kOhms | kOhms |  |  |  | 55.449 | 0.00027074 | 55.444 | 0.00027056
| Resistance1000010000 | Resistance | 10000 | 100000 | 10000 | kOhms | kOhms |  |  |  | 5773.5 | 3.3068e-06 | 5770.9 | 3.2955e-06
| Resistance0.1100000 | Resistance | 100000 | 1000000 | 0.1 | kOhms | kOhms |  |  |  | 12 | 0.0058 | 11.592 | 0.0057735
| Resistance1100000 | Resistance | 100000 | 1000000 | 1 | kOhms | kOhms |  |  |  | 12 | 0.0058 | 11.592 | 0.0057735
| Resistance10100000 | Resistance | 100000 | 1000000 | 10 | kOhms | kOhms |  |  |  | 12 | 0.0058 | 11.595 | 0.0057707
| Resistance100100000 | Resistance | 100000 | 1000000 | 100 | kOhms | kOhms |  |  |  | 12 | 0.0058 | 14.68 | 0.0057707
| Resistance10000100000 | Resistance | 100000 | 1000000 | 10000 | kOhms | kOhms |  |  |  | 5540.2 | 0.0026331 | 5540.2 | 0.002633
| Resistance100000100000 | Resistance | 100000 | 1000000 | 100000 | kOhms | kOhms |  |  |  | 57707 | 0.00031791 | 57475 | 0.0003179
| ACV0.0000000101000 | ACV | 0 | 0.01 | 1e-08 | V | µV | 1000 | 4000 | kHz | 23 | 80000 | 22.679 | 79411
| ACV0.000000101000 | ACV | 0 | 0.01 | 1e-07 | V | µV | 1000 | 4000 | kHz | 23 | 79583 | 22.679 | 79411
| ACV0.00000101000 | ACV | 0 | 0.01 | 1e-06 | V | µV | 1000 | 4000 | kHz | 23 | 79581 | 22.688 | 79410
| ACV0.0000101000 | ACV | 0 | 0.01 | 1e-05 | V | µV | 1000 | 4000 | kHz | 23.293 | 79521 | 23.228 | 79358
| ACV0.000101000 | ACV | 0 | 0.01 | 0.0001 | V | µV | 1000 | 4000 | kHz | 57.909 | 76095 | 57.746 | 76108
| ACV0.00101000 | ACV | 0 | 0.01 | 0.001 | V | µV | 1000 | 4000 | kHz | 573.98 | 42637 | 570.56 | 42634
| ACV0.0101000 | ACV | 0 | 0.01 | 0.01 | V | µV | 1000 | 4000 | kHz | 5773.1 | 5803.4 | 5766 | 5798.9
| ACV0.101000 | ACV | 0 | 0.01 | 0.1 | V | µV | 1000 | 4000 | kHz | 57735 | 584.6 | 57728 | 582.75
| ACV101000 | ACV | 0 | 0.01 | 1 | V | µV | 1000 | 4000 | kHz | 577360 | 59.02 | 577350 | 58.278
| ACV0.0000000104000 | ACV | 0 | 0.01 | 1e-08 | V | µV | 4000 | 8000 | kHz | 31 | 230000 | 30.119 | 228890
| ACV0.000000104000 | ACV | 0 | 0.01 | 1e-07 | V | µV | 4000 | 8000 | kHz | 31 | 228990 | 30.12 | 228890
| ACV0.00000104000 | ACV | 0 | 0.01 | 1e-06 | V | µV | 4000 | 8000 | kHz | 31 | 228990 | 30.12 | 228890
| ACV0.0000104000 | ACV | 0 | 0.01 | 1e-05 | V | µV | 4000 | 8000 | kHz | 31 | 228940 | 30.434 | 228860
| ACV0.000104000 | ACV | 0 | 0.01 | 0.0001 | V | µV | 4000 | 8000 | kHz | 56.084 | 226390 | 55.738 | 226400
| ACV0.00104000 | ACV | 0 | 0.01 | 0.001 | V | µV | 4000 | 8000 | kHz | 561.69 | 182820 | 557.14 | 182830
| ACV0.0104000 | ACV | 0 | 0.01 | 0.01 | V | µV | 4000 | 8000 | kHz | 5769.3 | 45269 | 5751.2 | 45261
| ACV0.104000 | ACV | 0 | 0.01 | 0.1 | V | µV | 4000 | 8000 | kHz | 57735 | 4703.5 | 57713 | 4700
| ACV104000 | ACV | 0 | 0.01 | 1 | V | µV | 4000 | 8000 | kHz | 577360 | 471.57 | 577330 | 470.19
| ACV0.000000010.011000 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 1000 | 2000 | kHz | 53 | 17000 | 11.488 | 17414
| ACV0.00000010.011000 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.488 | 17414
| ACV0.0000010.011000 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.489 | 17414
| ACV0.000010.011000 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 53 | 17054 | 11.588 | 17413
| ACV0.00010.011000 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 53 | 17065 | 21.129 | 17327
| ACV0.0010.011000 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 1000 | 2000 | kHz | 469.05 | 13766 | 433.19 | 13767
| ACV0.000000010.012000 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.778 | 46223
| ACV0.00000010.012000 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.778 | 46223
| ACV0.0000010.012000 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.779 | 46223
| ACV0.000010.012000 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 80.809 | 46223
| ACV0.00010.012000 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 2000 | 4000 | kHz | 110 | 46000 | 84.139 | 46193
| ACV0.0010.012000 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 2000 | 4000 | kHz | 354.54 | 43842 | 330.66 | 43842
| ACV0.000000010.014000 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.256 | 46261
| ACV0.00000010.014000 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.256 | 46261
| ACV0.0000010.014000 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.257 | 46261
| ACV0.000010.014000 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 92.287 | 46261
| ACV0.00010.014000 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 4000 | 8000 | kHz | 120 | 46000 | 95.542 | 46232
| ACV0.0010.014000 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 4000 | 8000 | kHz | 363.53 | 43909 | 338.44 | 43920
| ACV0.000000010.018000 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 8000 | 10000 | kHz | 450 | 170000 | 115.4 | 173280
| ACV0.00000010.018000 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 8000 | 10000 | kHz | 450 | 170030 | 115.4 | 173280
| ACV0.0000010.018000 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 115.4 | 173280
| ACV0.000010.018000 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 115.41 | 173280
| ACV0.00010.018000 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 8000 | 10000 | kHz | 450 | 170040 | 116.4 | 173270
| ACV0.0010.018000 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 8000 | 10000 | kHz | 450 | 170150 | 203.58 | 172400
| ACV0.00000010.11000 | ACV | 0.1 | 1 | 1e-07 | V | µV | 1000 | 2000 | kHz | 590 | 17000 | 108.37 | 17479
| ACV0.0000010.11000 | ACV | 0.1 | 1 | 1e-06 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 108.37 | 17479
| ACV0.000010.11000 | ACV | 0.1 | 1 | 1e-05 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 108.38 | 17479
| ACV0.00010.11000 | ACV | 0.1 | 1 | 0.0001 | V | µV | 1000 | 2000 | kHz | 590 | 17152 | 109.37 | 17478
| ACV0.0010.11000 | ACV | 0.1 | 1 | 0.001 | V | µV | 1000 | 2000 | kHz | 590 | 17163 | 204.79 | 17392
| ACV0.010.11000 | ACV | 0.1 | 1 | 0.01 | V | µV | 1000 | 2000 | kHz | 4683.4 | 13828 | 4325.4 | 13829
| ACV0.00000010.12000 | ACV | 0.1 | 1 | 1e-07 | V | mV | 2000 | 4000 | kHz | 1.2 | 46 | 0.80072 | 46.326
| ACV0.0000010.12000 | ACV | 0.1 | 1 | 1e-06 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.029 | 0.80072 | 46.326
| ACV0.000010.12000 | ACV | 0.1 | 1 | 1e-05 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.029 | 0.80073 | 46.326
| ACV0.00010.12000 | ACV | 0.1 | 1 | 0.0001 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.03 | 0.8015 | 46.325
| ACV0.0010.12000 | ACV | 0.1 | 1 | 0.001 | V | mV | 2000 | 4000 | kHz | 1.2 | 46.034 | 0.83431 | 46.296
| ACV0.010.12000 | ACV | 0.1 | 1 | 0.01 | V | mV | 2000 | 4000 | kHz | 3.5375 | 43.944 | 3.2985 | 43.945
| ACV0.00000010.14000 | ACV | 0.1 | 1 | 1e-07 | V | mV | 4000 | 8000 | kHz | 1.3 | 46 | 0.91606 | 46.326
| ACV0.0000010.14000 | ACV | 0.1 | 1 | 1e-06 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91606 | 46.326
| ACV0.000010.14000 | ACV | 0.1 | 1 | 1e-05 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91607 | 46.326
| ACV0.00010.14000 | ACV | 0.1 | 1 | 0.0001 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.046 | 0.91665 | 46.325
| ACV0.0010.14000 | ACV | 0.1 | 1 | 0.001 | V | mV | 4000 | 8000 | kHz | 1.3 | 46.051 | 0.94915 | 46.296
| ACV0.010.14000 | ACV | 0.1 | 1 | 0.01 | V | mV | 4000 | 8000 | kHz | 3.6131 | 43.982 | 3.378 | 43.984
| ACV0.00000010.18000 | ACV | 0.1 | 1 | 1e-07 | V | mV | 8000 | 10000 | kHz | 4.6 | 170 | 1.146 | 173.36
| ACV0.0000010.18000 | ACV | 0.1 | 1 | 1e-06 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.146 | 173.36
| ACV0.000010.18000 | ACV | 0.1 | 1 | 1e-05 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.146 | 173.36
| ACV0.00010.18000 | ACV | 0.1 | 1 | 0.0001 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.1461 | 173.36
| ACV0.0010.18000 | ACV | 0.1 | 1 | 0.001 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.12 | 1.156 | 173.35
| ACV0.010.18000 | ACV | 0.1 | 1 | 0.01 | V | mV | 8000 | 10000 | kHz | 4.6 | 170.23 | 2.0278 | 172.49
| ACV0.00000111000 | ACV | 1 | 10 | 1e-06 | V | mV | 1000 | 2000 | kHz | 6.1 | 17 | 1.1338 | 17.487
| ACV0.0000111000 | ACV | 1 | 10 | 1e-05 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1338 | 17.487
| ACV0.000111000 | ACV | 1 | 10 | 0.0001 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1338 | 17.487
| ACV0.00111000 | ACV | 1 | 10 | 0.001 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.156 | 1.1448 | 17.486
| ACV0.0111000 | ACV | 1 | 10 | 0.01 | V | mV | 1000 | 2000 | kHz | 6.1 | 17.167 | 2.0915 | 17.401
| ACV0.111000 | ACV | 1 | 10 | 0.1 | V | mV | 1000 | 2000 | kHz | 46.847 | 13.84 | 43.262 | 13.841
| ACV0.00000112000 | ACV | 1 | 10 | 1e-06 | V | mV | 2000 | 4000 | kHz | 12 | 46 | 8.0599 | 46.328
| ACV0.0000112000 | ACV | 1 | 10 | 1e-05 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.06 | 46.328
| ACV0.000112000 | ACV | 1 | 10 | 0.0001 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.06 | 46.328
| ACV0.00112000 | ACV | 1 | 10 | 0.001 | V | mV | 2000 | 4000 | kHz | 12 | 46.038 | 8.0603 | 46.328
| ACV0.0112000 | ACV | 1 | 10 | 0.01 | V | mV | 2000 | 4000 | kHz | 12 | 46.042 | 8.3953 | 46.298
| ACV0.112000 | ACV | 1 | 10 | 0.1 | V | mV | 2000 | 4000 | kHz | 35.391 | 43.95 | 33.018 | 43.95
| ACV0.00000114000 | ACV | 1 | 10 | 1e-06 | V | mV | 4000 | 8000 | kHz | 15 | 46 | 9.1632 | 46.576
| ACV0.0000114000 | ACV | 1 | 10 | 1e-05 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1632 | 46.576
| ACV0.000114000 | ACV | 1 | 10 | 0.0001 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1632 | 46.576
| ACV0.00114000 | ACV | 1 | 10 | 0.001 | V | mV | 4000 | 8000 | kHz | 15 | 46.325 | 9.1658 | 46.576
| ACV0.0114000 | ACV | 1 | 10 | 0.01 | V | mV | 4000 | 8000 | kHz | 15 | 46.329 | 9.49 | 46.547
| ACV0.114000 | ACV | 1 | 10 | 0.1 | V | mV | 4000 | 8000 | kHz | 36.034 | 44.241 | 33.704 | 44.241
| ACV0.00000118000 | ACV | 1 | 10 | 1e-06 | V | mV | 8000 | 10000 | kHz | 46 | 170 | 11.528 | 173.36
| ACV0.0000118000 | ACV | 1 | 10 | 1e-05 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.529 | 173.36
| ACV0.000118000 | ACV | 1 | 10 | 0.0001 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.529 | 173.36
| ACV0.00118000 | ACV | 1 | 10 | 0.001 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.584 | 173.35
| ACV0.0118000 | ACV | 1 | 10 | 0.01 | V | mV | 8000 | 10000 | kHz | 46 | 170.12 | 11.629 | 173.35
| ACV0.118000 | ACV | 1 | 10 | 0.1 | V | mV | 8000 | 10000 | kHz | 46 | 170.24 | 20.344 | 172.48
| ACV0.0000110100 | ACV | 10 | 100 | 1e-05 | V | mV | 100 | 300 | kHz | 15 | 4.6 | 11.575 | 4.6268
| ACV0.000110100 | ACV | 10 | 100 | 0.0001 | V | mV | 100 | 300 | kHz | 15 | 4.6028 | 11.575 | 4.6268
| ACV0.00110100 | ACV | 10 | 100 | 0.001 | V | mV | 100 | 300 | kHz | 15 | 4.6028 | 11.576 | 4.6268
| ACV0.0110100 | ACV | 10 | 100 | 0.01 | V | mV | 100 | 300 | kHz | 15 | 4.6033 | 11.89 | 4.624
| ACV0.110100 | ACV | 10 | 100 | 0.1 | V | mV | 100 | 300 | kHz | 37.775 | 4.4 | 37.726 | 4.4003
| ACV110100 | ACV | 10 | 100 | 1 | V | mV | 100 | 300 | kHz | 561.72 | 1.8549 | 536.24 | 1.8547
| ACV0.0000110300 | ACV | 10 | 100 | 1e-05 | V | mV | 300 | 1000 | kHz | 46 | 17 | 12.317 | 17.334
| ACV0.000110300 | ACV | 10 | 100 | 0.0001 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.318 | 17.334
| ACV0.00110300 | ACV | 10 | 100 | 0.001 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.318 | 17.334
| ACV0.0110300 | ACV | 10 | 100 | 0.01 | V | mV | 300 | 1000 | kHz | 46 | 17.019 | 12.427 | 17.333
| ACV0.110300 | ACV | 10 | 100 | 0.1 | V | mV | 300 | 1000 | kHz | 46 | 17.03 | 21.972 | 17.247
| ACV110300 | ACV | 10 | 100 | 1 | V | mV | 300 | 1000 | kHz | 469.61 | 13.691 | 433.99 | 13.692
| ACV0.00011001 | ACV | 100 | 1000 | 0.0001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.055 | 0.69387
| ACV0.0011001 | ACV | 100 | 1000 | 0.001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.063 | 0.69386
| ACV0.011001 | ACV | 100 | 1000 | 0.01 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.256 | 0.69369
| ACV0.11001 | ACV | 100 | 1000 | 0.1 | V | mV | 1 | 20 | kHz | 40.883 | 0.68212 | 41.185 | 0.67806
| ACV11001 | ACV | 100 | 1000 | 1 | V | mV | 1 | 20 | kHz | 547.47 | 0.37317 | 547.4 | 0.3731
| ACV0.000110020 | ACV | 100 | 1000 | 0.0001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 22.895 | 1.3896
| ACV0.00110020 | ACV | 100 | 1000 | 0.001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 22.896 | 1.3896
| ACV0.0110020 | ACV | 100 | 1000 | 0.01 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.007 | 1.3895
| ACV0.110020 | ACV | 100 | 1000 | 0.1 | V | mV | 20 | 50 | kHz | 33.01 | 1.39 | 33.875 | 1.3798
| ACV110020 | ACV | 100 | 1000 | 1 | V | mV | 20 | 50 | kHz | 496.92 | 1.0292 | 496.68 | 1.0293
| ACV0.000110050 | ACV | 100 | 1000 | 0.0001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.116 | 3.4656
| ACV0.00110050 | ACV | 100 | 1000 | 0.001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.117 | 3.4656
| ACV0.0110050 | ACV | 100 | 1000 | 0.01 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.153 | 3.4656
| ACV0.110050 | ACV | 100 | 1000 | 0.1 | V | mV | 50 | 100 | kHz | 24.483 | 3.4985 | 28.03 | 3.4612
| ACV110050 | ACV | 100 | 1000 | 1 | V | mV | 50 | 100 | kHz | 369.25 | 3.1671 | 368.87 | 3.1673
| ACV0.0000000100.001 | ACV | 0 | 0.01 | 1e-08 | V | µV | 0.001 | 0.04 | kHz | 4.5 | 310 | 4.3849 | 312.12
| ACV0.000000100.001 | ACV | 0 | 0.01 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 4.5 | 312.16 | 4.3853 | 312.11
| ACV0.00000100.001 | ACV | 0 | 0.01 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 4.5 | 310.77 | 4.4226 | 310.57
| ACV0.0000100.001 | ACV | 0 | 0.01 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 7.2506 | 223.08 | 7.2466 | 222.31
| ACV0.000100.001 | ACV | 0 | 0.01 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 57.903 | 32.603 | 57.901 | 32.045
| ACV0.00100.001 | ACV | 0 | 0.01 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 577.37 | 3.4481 | 577.37 | 3.2224
| ACV0.0000000100.04 | ACV | 0 | 0.01 | 1e-08 | V | µV | 0.04 | 1 | kHz | 3 | 160 | 2.9725 | 162.69
| ACV0.000000100.04 | ACV | 0 | 0.01 | 1e-07 | V | µV | 0.04 | 1 | kHz | 3 | 162.76 | 2.9731 | 162.67
| ACV0.00000100.04 | ACV | 0 | 0.01 | 1e-06 | V | µV | 0.04 | 1 | kHz | 3.0304 | 161.11 | 3.028 | 160.75
| ACV0.0000100.04 | ACV | 0 | 0.01 | 1e-05 | V | µV | 0.04 | 1 | kHz | 6.4968 | 89.969 | 6.4924 | 88.922
| ACV0.000100.04 | ACV | 0 | 0.01 | 0.0001 | V | µV | 0.04 | 1 | kHz | 57.814 | 11.229 | 57.812 | 10.667
| ACV0.00100.04 | ACV | 0 | 0.01 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.36 | 1.2948 | 577.36 | 1.0691
| ACV0.000000101 | ACV | 0 | 0.01 | 1e-07 | V | µV | 1 | 20 | kHz | 3.1 | 250 | 3.0556 | 254.12
| ACV0.0000000101 | ACV | 0 | 0.01 | 1e-08 | V | µV | 1 | 20 | kHz | 3.1 | 254.16 | 3.0563 | 254.09
| ACV0.00000101 | ACV | 0 | 0.01 | 1e-06 | V | µV | 1 | 20 | kHz | 3.1119 | 251.96 | 3.1096 | 251.7
| ACV0.0000101 | ACV | 0 | 0.01 | 1e-05 | V | µV | 1 | 20 | kHz | 6.5335 | 152.12 | 6.5291 | 151.19
| ACV0.000101 | ACV | 0 | 0.01 | 0.0001 | V | µV | 1 | 20 | kHz | 57.818 | 19.6 | 57.816 | 19.039
| ACV0.00101 | ACV | 0 | 0.01 | 0.001 | V | µV | 1 | 20 | kHz | 577.36 | 2.1354 | 577.36 | 1.9097
| ACV0.00000001020 | ACV | 0 | 0.01 | 1e-08 | V | µV | 20 | 50 | kHz | 5.6 | 840 | 5.5358 | 839.05
| ACV0.0000001020 | ACV | 0 | 0.01 | 1e-07 | V | µV | 20 | 50 | kHz | 5.6 | 839.02 | 5.5362 | 839.03
| ACV0.000001020 | ACV | 0 | 0.01 | 1e-06 | V | µV | 20 | 50 | kHz | 5.6 | 837.24 | 5.5656 | 837.27
| ACV0.00001020 | ACV | 0 | 0.01 | 1e-05 | V | µV | 20 | 50 | kHz | 7.9956 | 709.13 | 7.9861 | 708.96
| ACV0.0001020 | ACV | 0 | 0.01 | 0.0001 | V | µV | 20 | 50 | kHz | 58 | 140.31 | 57.994 | 139.7
| ACV0.001020 | ACV | 0 | 0.01 | 0.001 | V | µV | 20 | 50 | kHz | 577.38 | 14.459 | 577.37 | 14.2
| ACV0.00000001050 | ACV | 0 | 0.01 | 1e-08 | V | µV | 50 | 100 | kHz | 8.8 | 5100 | 8.3103 | 5147.5
| ACV0.0000001050 | ACV | 0 | 0.01 | 1e-07 | V | µV | 50 | 100 | kHz | 8.8 | 5147.5 | 8.3103 | 5147.5
| ACV0.000001050 | ACV | 0 | 0.01 | 1e-06 | V | µV | 50 | 100 | kHz | 8.8 | 5145.7 | 8.3291 | 5145.9
| ACV0.00001050 | ACV | 0 | 0.01 | 1e-05 | V | µV | 50 | 100 | kHz | 10.062 | 5001.3 | 10.046 | 5001.8
| ACV0.0001050 | ACV | 0 | 0.01 | 0.0001 | V | µV | 50 | 100 | kHz | 58.168 | 2497.7 | 58.157 | 2495.6
| ACV0.001050 | ACV | 0 | 0.01 | 0.001 | V | µV | 50 | 100 | kHz | 577.4 | 306.39 | 577.17 | 305.02
| ACV0.000000010100 | ACV | 0 | 0.01 | 1e-08 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.8932 | 46131
| ACV0.00000010100 | ACV | 0 | 0.01 | 1e-07 | V | µV | 100 | 300 | kHz | 4.2 | 46000 | 2.8935 | 46131
| ACV0.0000010100 | ACV | 0 | 0.01 | 1e-06 | V | µV | 100 | 300 | kHz | 4.2005 | 46000 | 2.919 | 46128
| ACV0.000010100 | ACV | 0 | 0.01 | 1e-05 | V | µV | 100 | 300 | kHz | 4.9503 | 45929 | 4.8763 | 45936
| ACV0.00010100 | ACV | 0 | 0.01 | 0.0001 | V | µV | 100 | 300 | kHz | 54.13 | 41367 | 54.086 | 41369
| ACV0.0010100 | ACV | 0 | 0.01 | 0.001 | V | µV | 100 | 300 | kHz | 575.77 | 16511 | 573.27 | 16507
| ACV0.000000010300 | ACV | 0 | 0.01 | 1e-08 | V | µV | 300 | 1000 | kHz | 18 | 45000 | 12.904 | 45496
| ACV0.00000010300 | ACV | 0 | 0.01 | 1e-07 | V | µV | 300 | 1000 | kHz | 18 | 45496 | 12.904 | 45496
| ACV0.0000010300 | ACV | 0 | 0.01 | 1e-06 | V | µV | 300 | 1000 | kHz | 18 | 45495 | 12.915 | 45495
| ACV0.000010300 | ACV | 0 | 0.01 | 1e-05 | V | µV | 300 | 1000 | kHz | 18 | 45403 | 13.84 | 45406
| ACV0.00010300 | ACV | 0 | 0.01 | 0.0001 | V | µV | 300 | 1000 | kHz | 56.207 | 41523 | 56.163 | 41525
| ACV0.0010300 | ACV | 0 | 0.01 | 0.001 | V | µV | 300 | 1000 | kHz | 575.97 | 16722 | 573.47 | 16718
| ACV0.000000010.010.001 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 0.001 | 0.04 | kHz | 5.4 | 83 | 5.316 | 83.165
| ACV0.00000010.010.001 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 5.4 | 83.166 | 5.3163 | 83.163
| ACV0.0000010.010.001 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 5.4 | 83.01 | 5.3447 | 83
| ACV0.000010.010.001 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 7.7286 | 70.835 | 7.7258 | 70.789
| ACV0.00010.010.001 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 57.923 | 14.075 | 57.922 | 14.014
| ACV0.0010.010.001 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 577.37 | 1.4498 | 577.25 | 1.4244
| ACV0.000000010.010.04 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 0.04 | 1 | kHz | 3.5 | 81 | 3.4429 | 80.684
| ACV0.00000010.010.04 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 0.04 | 1 | kHz | 3.5 | 80.684 | 3.4434 | 80.682
| ACV0.0000010.010.04 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 0.04 | 1 | kHz | 3.5 | 80.42 | 3.4848 | 80.411
| ACV0.000010.010.04 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 0.04 | 1 | kHz | 6.538 | 63.487 | 6.5346 | 63.435
| ACV0.00010.010.04 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 0.04 | 1 | kHz | 57.784 | 10.953 | 57.783 | 10.891
| ACV0.0010.010.04 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.36 | 1.1266 | 577.26 | 1.1012
| ACV0.000000010.011 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 1 | 20 | kHz | 3.4 | 160 | 3.3551 | 157.72
| ACV0.00000010.011 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 1 | 20 | kHz | 3.4 | 157.72 | 3.3552 | 157.72
| ACV0.0000010.011 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 1 | 20 | kHz | 3.4 | 157.45 | 3.3911 | 157.45
| ACV0.000010.011 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 1 | 20 | kHz | 6.2204 | 137.65 | 6.2175 | 137.62
| ACV0.00010.011 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 1 | 20 | kHz | 57.628 | 32.01 | 57.626 | 31.951
| ACV0.0010.011 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 1 | 20 | kHz | 577.34 | 3.3107 | 577.06 | 3.2853
| ACV0.000000010.0120 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 20 | 50 | kHz | 4.7 | 330 | 4.6688 | 326.81
| ACV0.00000010.0120 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 20 | 50 | kHz | 4.7 | 327.36 | 4.6691 | 326.81
| ACV0.0000010.0120 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 20 | 50 | kHz | 4.7 | 327.1 | 4.6916 | 326.63
| ACV0.000010.0120 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 20 | 50 | kHz | 6.7172 | 310.87 | 6.7063 | 310.87
| ACV0.00010.0120 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 20 | 50 | kHz | 57.119 | 116.68 | 57.114 | 116.5
| ACV0.0010.0120 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 20 | 50 | kHz | 577.28 | 12.896 | 576.24 | 12.803
| ACV0.000000010.0150 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 50 | 100 | kHz | 6 | 910 | 5.9729 | 906.03
| ACV0.00000010.0150 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 50 | 100 | kHz | 6 | 906.02 | 5.973 | 906.03
| ACV0.0000010.0150 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 50 | 100 | kHz | 6 | 905.89 | 5.9854 | 905.92
| ACV0.000010.0150 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 50 | 100 | kHz | 7.1698 | 895.95 | 7.1433 | 896.05
| ACV0.00010.0150 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 50 | 100 | kHz | 53.812 | 587.62 | 53.788 | 587.3
| ACV0.0010.0150 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 50 | 100 | kHz | 576.69 | 87.306 | 571.68 | 86.954
| ACV0.000000010.01100 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.521 | 3476.1
| ACV0.00000010.01100 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.521 | 3476.1
| ACV0.0000010.01100 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.522 | 3476.1
| ACV0.000010.01100 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 100 | 300 | kHz | 12 | 3500 | 11.911 | 3472.7
| ACV0.00010.01100 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 100 | 300 | kHz | 41.904 | 3219 | 41.802 | 3219.4
| ACV0.0010.01100 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 100 | 300 | kHz | 568.06 | 1120 | 547.01 | 1119.2
| ACV0.000000010.01300 | ACV | 0.01 | 0.1 | 1e-08 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
| ACV0.00000010.01300 | ACV | 0.01 | 0.1 | 1e-07 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
| ACV0.0000010.01300 | ACV | 0.01 | 0.1 | 1e-06 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.521 | 11631
| ACV0.000010.01300 | ACV | 0.01 | 0.1 | 1e-05 | V | µV | 300 | 1000 | kHz | 12 | 12000 | 11.635 | 11630
| ACV0.00010.01300 | ACV | 0.01 | 0.1 | 0.0001 | V | µV | 300 | 1000 | kHz | 21.308 | 11907 | 25.139 | 11509
| ACV0.0010.01300 | ACV | 0.01 | 0.1 | 0.001 | V | µV | 300 | 1000 | kHz | 511.75 | 7972.6 | 476.24 | 7972.4
| ACV0.00000010.10.001 | ACV | 0.1 | 1 | 1e-07 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 80.62 | 80.62
| ACV0.0000010.10.001 | ACV | 0.1 | 1 | 1e-06 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 80.618 | 80.618
| ACV0.000010.10.001 | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.001 | 0.04 | kHz | 47 | 82 | 80.426 | 80.426
| ACV0.00010.10.001 | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.001 | 0.04 | kHz | 72.743 | 66.872 | 66.849 | 66.849
| ACV0.0010.10.001 | ACV | 0.1 | 1 | 0.001 | V | µV | 0.001 | 0.04 | kHz | 578.69 | 12.536 | 12.498 | 12.498
| ACV0.010.10.001 | ACV | 0.1 | 1 | 0.01 | V | µV | 0.001 | 0.04 | kHz | 5773.7 | 1.2831 | 1.2674 | 1.2674
| ACV0.00000010.10.04 | ACV | 0.1 | 1 | 1e-07 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 23.524 | 80.445
| ACV0.0000010.10.04 | ACV | 0.1 | 1 | 1e-06 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 23.53 | 80.441
| ACV0.000010.10.04 | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.04 | 1 | kHz | 24 | 82 | 24.088 | 80.041
| ACV0.00010.10.04 | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.04 | 1 | kHz | 59.943 | 59.049 | 59.9 | 59.024
| ACV0.0010.10.04 | ACV | 0.1 | 1 | 0.001 | V | µV | 0.04 | 1 | kHz | 577.3 | 9.3986 | 577.28 | 9.3602
| ACV0.010.10.04 | ACV | 0.1 | 1 | 0.01 | V | µV | 0.04 | 1 | kHz | 5773.6 | 0.9599 | 5772.7 | 0.94416
| ACV0.00000010.11 | ACV | 0.1 | 1 | 1e-07 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.438 | 161.34
| ACV0.0000010.11 | ACV | 0.1 | 1 | 1e-06 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.442 | 161.34
| ACV0.000010.11 | ACV | 0.1 | 1 | 1e-05 | V | µV | 1 | 20 | kHz | 26 | 160 | 23.893 | 160.98
| ACV0.00010.11 | ACV | 0.1 | 1 | 0.0001 | V | µV | 1 | 20 | kHz | 56.377 | 137.35 | 56.262 | 137.33
| ACV0.0010.11 | ACV | 0.1 | 1 | 0.001 | V | µV | 1 | 20 | kHz | 575.7 | 30.683 | 575.66 | 30.548
| ACV0.010.11 | ACV | 0.1 | 1 | 0.01 | V | µV | 1 | 20 | kHz | 5773.4 | 3.1917 | 5770.6 | 0.31603
| ACV0.00000010.120 | ACV | 0.1 | 1 | 1e-07 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.379 | 346.14
| ACV0.0000010.120 | ACV | 0.1 | 1 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.383 | 346.14
| ACV0.000010.120 | ACV | 0.1 | 1 | 1e-05 | V | µV | 20 | 50 | kHz | 24 | 350 | 23.693 | 345.87
| ACV0.00010.120 | ACV | 0.1 | 1 | 0.0001 | V | µV | 20 | 50 | kHz | 49.713 | 324.48 | 49.371 | 324.63
| ACV0.0010.120 | ACV | 0.1 | 1 | 0.001 | V | µV | 20 | 50 | kHz | 568.73 | 117.17 | 568.57 | 116.91
| ACV0.010.120 | ACV | 0.1 | 1 | 0.01 | V | µV | 20 | 50 | kHz | 5772.6 | 12.94 | 5762.2 | 12.802
| ACV0.00000010.150 | ACV | 0.1 | 1 | 1e-07 | V | µV | 50 | 100 | kHz | 27 | 940 | 23.141 | 938.86
| ACV0.0000010.150 | ACV | 0.1 | 1 | 1e-06 | V | µV | 50 | 100 | kHz | 27 | 935.56 | 23.148 | 938.85
| ACV0.000010.150 | ACV | 0.1 | 1 | 1e-05 | V | µV | 50 | 100 | kHz | 27 | 935.62 | 23.298 | 938.72
| ACV0.00010.150 | ACV | 0.1 | 1 | 0.0001 | V | µV | 50 | 100 | kHz | 38.506 | 925.39 | 37.913 | 925.82
| ACV0.0010.150 | ACV | 0.1 | 1 | 0.001 | V | µV | 50 | 100 | kHz | 530.37 | 592.16 | 529.89 | 592.07
| ACV0.010.150 | ACV | 0.1 | 1 | 0.01 | V | µV | 50 | 100 | kHz | 5766.2 | 87.392 | 5715.5 | 87.123
| ACV0.00000010.1100 | ACV | 0.1 | 1 | 1e-07 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.46 | 3463.3
| ACV0.0000010.1100 | ACV | 0.1 | 1 | 1e-06 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.47 | 3463.3
| ACV0.000010.1100 | ACV | 0.1 | 1 | 1e-05 | V | µV | 100 | 300 | kHz | 120 | 3500 | 116.55 | 3463.2
| ACV0.00010.1100 | ACV | 0.1 | 1 | 0.0001 | V | µV | 100 | 300 | kHz | 120 | 3500 | 120.41 | 3459.8
| ACV0.0010.1100 | ACV | 0.1 | 1 | 0.001 | V | µV | 100 | 300 | kHz | 420.55 | 3205.9 | 419.31 | 3206.7
| ACV0.010.1100 | ACV | 0.1 | 1 | 0.01 | V | µV | 100 | 300 | kHz | 5681.4 | 1112.8 | 5471.4 | 1112.4
| ACV0.0000010.1300 | ACV | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
| ACV0.0000010.1300 | ACV | 0.1 | 1 | 1e-06 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
| ACV0.000010.1300 | ACV | 0.1 | 1 | 1e-05 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 110.73 | 11675
| ACV0.00010.1300 | ACV | 0.1 | 1 | 0.0001 | V | µV | 300 | 1000 | kHz | 120 | 12000 | 111.87 | 11674
| ACV0.0010.1300 | ACV | 0.1 | 1 | 0.001 | V | µV | 300 | 1000 | kHz | 212.67 | 11907 | 246.94 | 11553
| ACV0.010.1300 | ACV | 0.1 | 1 | 0.01 | V | µV | 300 | 1000 | kHz | 5113.2 | 8011.6 | 4758 | 8011.9
| ACV0.00000110.001 | ACV | 1 | 10 | 1e-06 | V | mV | 0.001 | 0.04 | kHz | 0.46 | 0.092 | 0.45053 | 0.092199
| ACV0.0000110.001 | ACV | 1 | 10 | 1e-05 | V | mV | 0.001 | 0.04 | kHz | 0.46 | 0.091968 | 0.45057 | 0.092197
| ACV0.000110.001 | ACV | 1 | 10 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 0.46 | 0.092019 | 0.4538 | 0.091994
| ACV0.00110.001 | ACV | 1 | 10 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 0.71619 | 0.077546 | 0.71499 | 0.077402
| ACV0.0110.001 | ACV | 1 | 10 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 5.7845 | 0.015263 | 5.784 | 0.01505
| ACV0.110.001 | ACV | 1 | 10 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.737 | 0.0016174 | 57.723 | 0.001529
| ACV0.00000110.04 | ACV | 1 | 10 | 1e-06 | V | mV | 0.04 | 1 | kHz | 0.25 | 0.084 | 0.22625 | 0.085545
| ACV0.0000110.04 | ACV | 1 | 10 | 1e-05 | V | mV | 0.04 | 1 | kHz | 0.25 | 0.084438 | 0.22632 | 0.08554
| ACV0.000110.04 | ACV | 1 | 10 | 0.0001 | V | mV | 0.04 | 1 | kHz | 0.25 | 0.084683 | 0.23198 | 0.085127
| ACV0.00110.04 | ACV | 1 | 10 | 0.001 | V | mV | 0.04 | 1 | kHz | 0.59428 | 0.063506 | 0.59284 | 0.06333
| ACV0.0110.04 | ACV | 1 | 10 | 0.01 | V | mV | 0.04 | 1 | kHz | 5.7723 | 0.010443 | 5.7717 | 0.010227
| ACV0.110.04 | ACV | 1 | 10 | 0.1 | V | mV | 0.04 | 1 | kHz | 57.736 | 0.0011207 | 57.726 | 0.0010323
| ACV0.00000111 | ACV | 1 | 10 | 1e-06 | V | mV | 1 | 20 | kHz | 0.28 | 0.16 | 0.22835 | 0.16434
| ACV0.0000111 | ACV | 1 | 10 | 1e-05 | V | mV | 1 | 20 | kHz | 0.28 | 0.16134 | 0.22837 | 0.16434
| ACV0.000111 | ACV | 1 | 10 | 0.0001 | V | mV | 1 | 20 | kHz | 0.28 | 0.16156 | 0.23295 | 0.16397
| ACV0.00111 | ACV | 1 | 10 | 0.001 | V | mV | 1 | 20 | kHz | 0.56386 | 0.14001 | 0.55818 | 0.14006
| ACV0.0111 | ACV | 1 | 10 | 0.01 | V | mV | 1 | 20 | kHz | 5.7578 | 0.031829 | 5.7555 | 0.031388
| ACV0.111 | ACV | 1 | 10 | 0.1 | V | mV | 1 | 20 | kHz | 57.735 | 0.0034135 | 57.705 | 0.003222
| ACV0.000001120 | ACV | 1 | 10 | 1e-06 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.22967 | 0.34776
| ACV0.00001120 | ACV | 1 | 10 | 1e-05 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.22973 | 0.34775
| ACV0.0001120 | ACV | 1 | 10 | 0.0001 | V | mV | 20 | 50 | kHz | 0.24 | 0.35 | 0.23282 | 0.34749
| ACV0.001120 | ACV | 1 | 10 | 0.001 | V | mV | 20 | 50 | kHz | 0.50358 | 0.32533 | 0.49042 | 0.32615
| ACV0.01120 | ACV | 1 | 10 | 0.01 | V | mV | 20 | 50 | kHz | 5.6913 | 0.11808 | 5.6847 | 0.11766
| ACV0.1120 | ACV | 1 | 10 | 0.1 | V | mV | 20 | 50 | kHz | 57.728 | 0.013146 | 57.621 | 0.01289
| ACV0.000001150 | ACV | 1 | 10 | 1e-06 | V | mV | 50 | 100 | kHz | 0.27 | 0.94 | 0.22603 | 0.94121
| ACV0.00001150 | ACV | 1 | 10 | 1e-05 | V | mV | 50 | 100 | kHz | 0.27 | 0.93759 | 0.22604 | 0.94121
| ACV0.0001150 | ACV | 1 | 10 | 0.0001 | V | mV | 50 | 100 | kHz | 0.27 | 0.93768 | 0.22761 | 0.94107
| ACV0.001150 | ACV | 1 | 10 | 0.001 | V | mV | 50 | 100 | kHz | 0.38942 | 0.92692 | 0.37412 | 0.92813
| ACV0.01150 | ACV | 1 | 10 | 0.01 | V | mV | 50 | 100 | kHz | 5.3092 | 0.59368 | 5.2965 | 0.59386
| ACV0.1150 | ACV | 1 | 10 | 0.1 | V | mV | 50 | 100 | kHz | 57.665 | 0.087823 | 57.153 | 0.087462
| ACV0.0000011100 | ACV | 1 | 10 | 1e-06 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1359 | 3.4831
| ACV0.000011100 | ACV | 1 | 10 | 1e-05 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1359 | 3.4831
| ACV0.00011100 | ACV | 1 | 10 | 0.0001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1365 | 3.483
| ACV0.0011100 | ACV | 1 | 10 | 0.001 | V | mV | 100 | 300 | kHz | 1.2 | 3.5 | 1.1754 | 3.4796
| ACV0.011100 | ACV | 1 | 10 | 0.01 | V | mV | 100 | 300 | kHz | 4.1895 | 3.2244 | 4.1679 | 3.2259
| ACV0.11100 | ACV | 1 | 10 | 0.1 | V | mV | 100 | 300 | kHz | 56.807 | 1.1229 | 54.694 | 1.1224
| ACV0.0000011300 | ACV | 1 | 10 | 1e-06 | V | mV | 300 | 1000 | kHz | 0.83 | 12 | 1.127 | 11.688
| ACV0.000011300 | ACV | 1 | 10 | 1e-05 | V | mV | 300 | 1000 | kHz | 0.83 | 11.988 | 1.127 | 11.688
| ACV0.00011300 | ACV | 1 | 10 | 0.0001 | V | mV | 300 | 1000 | kHz | 0.83 | 11.988 | 1.1272 | 11.688
| ACV0.0011300 | ACV | 1 | 10 | 0.001 | V | mV | 300 | 1000 | kHz | 0.84663 | 11.986 | 1.141 | 11.687
| ACV0.011300 | ACV | 1 | 10 | 0.01 | V | mV | 300 | 1000 | kHz | 2.2239 | 11.849 | 2.4896 | 11.566
| ACV0.11300 | ACV | 1 | 10 | 0.1 | V | mV | 300 | 1000 | kHz | 51.132 | 8.0252 | 47.575 | 8.0258
| ACV0.00001100.001 | ACV | 10 | 20 | 1e-05 | V | mV | 0.001 | 0.04 | kHz | 4.7 | 0.23 | 4.6186 | 0.23291
| ACV0.0001100.001 | ACV | 10 | 20 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 4.7 | 0.23 | 4.619 | 0.2329
| ACV0.001100.001 | ACV | 10 | 20 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 4.7 | 0.23 | 4.6486 | 0.23231
| ACV0.01100.001 | ACV | 10 | 20 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 7.1399 | 0.18969 | 7.1403 | 0.18932
| ACV0.1100.001 | ACV | 10 | 20 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.827 | 0.032722 | 57.828 | 0.0324
| ACV1100.001 | ACV | 10 | 20 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.36 | 0.003403 | 577.07 | 0.0032722
| ACV0.00001100.04 | ACV | 10 | 20 | 1e-05 | V | mV | 0.04 | 1 | kHz | 2.4 | 0.23 | 2.298 | 0.23306
| ACV0.0001100.04 | ACV | 10 | 20 | 0.0001 | V | mV | 0.04 | 1 | kHz | 2.4 | 0.23 | 2.2985 | 0.23305
| ACV0.001100.04 | ACV | 10 | 20 | 0.001 | V | mV | 0.04 | 1 | kHz | 2.4 | 0.23 | 2.3458 | 0.23187
| ACV0.01100.04 | ACV | 10 | 20 | 0.01 | V | mV | 0.04 | 1 | kHz | 5.7569 | 0.1647 | 5.7574 | 0.16425
| ACV0.1100.04 | ACV | 10 | 20 | 0.1 | V | mV | 0.04 | 1 | kHz | 57.687 | 0.023591 | 57.688 | 0.023267
| ACV1100.04 | ACV | 10 | 20 | 1 | V | mV | 0.04 | 1 | kHz | 577.35 | 0.0024695 | 577.14 | 0.0023387
| ACV0.00001101 | ACV | 10 | 20 | 1e-05 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.298 | 0.23306
| ACV0.0001101 | ACV | 10 | 20 | 0.0001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.2985 | 0.23305
| ACV0.001101 | ACV | 10 | 20 | 0.001 | V | mV | 1 | 20 | kHz | 2.4 | 0.23 | 2.3458 | 0.23187
| ACV0.01101 | ACV | 10 | 20 | 0.01 | V | mV | 1 | 20 | kHz | 5.7569 | 0.1647 | 5.7574 | 0.16425
| ACV0.1101 | ACV | 10 | 20 | 0.1 | V | mV | 1 | 20 | kHz | 57.687 | 0.023591 | 57.688 | 0.023267
| ACV1101 | ACV | 10 | 20 | 1 | V | mV | 1 | 20 | kHz | 577.35 | 0.0024695 | 577.14 | 0.0023387
| ACV0.000011020 | ACV | 10 | 20 | 1e-05 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.2656 | 0.40924
| ACV0.00011020 | ACV | 10 | 20 | 0.0001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.266 | 0.40923
| ACV0.0011020 | ACV | 10 | 20 | 0.001 | V | mV | 20 | 50 | kHz | 2.4 | 0.41 | 2.302 | 0.40822
| ACV0.011020 | ACV | 10 | 20 | 0.01 | V | mV | 20 | 50 | kHz | 5.2425 | 0.33547 | 5.2373 | 0.3351
| ACV0.11020 | ACV | 10 | 20 | 0.1 | V | mV | 20 | 50 | kHz | 57.496 | 0.059412 | 57.495 | 0.058915
| ACV11020 | ACV | 10 | 20 | 1 | V | mV | 20 | 50 | kHz | 577.33 | 0.0061608 | 576.8 | 0.0059565
| ACV0.000011050 | ACV | 10 | 20 | 1e-05 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3767 | 1.3889
| ACV0.00011050 | ACV | 10 | 20 | 0.0001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3769 | 1.3889
| ACV0.0011050 | ACV | 10 | 20 | 0.001 | V | mV | 50 | 100 | kHz | 2.4 | 1.4 | 2.3922 | 1.3884
| ACV0.011050 | ACV | 10 | 20 | 0.01 | V | mV | 50 | 100 | kHz | 3.8217 | 1.3443 | 3.8184 | 1.3442
| ACV0.11050 | ACV | 10 | 20 | 0.1 | V | mV | 50 | 100 | kHz | 54.831 | 0.51571 | 54.83 | 0.51529
| ACV11050 | ACV | 10 | 20 | 1 | V | mV | 50 | 100 | kHz | 577.03 | 0.05599 | 572.43 | 0.055786
| ACV0.0000110100 | ACV | 10 | 20 | 1e-05 | V | mV | 100 | 300 | kHz | 13 | 4.6 | 11.573 | 4.6268
| ACV0.000110100 | ACV | 10 | 20 | 0.0001 | V | mV | 100 | 300 | kHz | 13 | 4.6109 | 11.574 | 4.6268
| ACV0.00110100 | ACV | 10 | 20 | 0.001 | V | mV | 100 | 300 | kHz | 13 | 4.6111 | 11.577 | 4.6267
| ACV0.0110100 | ACV | 10 | 20 | 0.01 | V | mV | 100 | 300 | kHz | 13 | 4.6141 | 11.988 | 4.6141
| ACV0.110100 | ACV | 10 | 20 | 0.1 | V | mV | 100 | 300 | kHz | 44.406 | 3.7323 | 44.403 | 3.7322
| ACV110100 | ACV | 10 | 20 | 1 | V | mV | 100 | 300 | kHz | 573.82 | 0.64234 | 542.92 | 0.64214
| ACV0.0000110300 | ACV | 10 | 20 | 1e-05 | V | mV | 300 | 1000 | kHz | 19 | 17 | 12.388 | 17.324
| ACV0.000110300 | ACV | 10 | 20 | 0.0001 | V | mV | 300 | 1000 | kHz | 19 | 17.087 | 12.389 | 17.324
| ACV0.00110300 | ACV | 10 | 20 | 0.001 | V | mV | 300 | 1000 | kHz | 19 | 17.087 | 12.389 | 17.324
| ACV0.0110300 | ACV | 10 | 20 | 0.01 | V | mV | 300 | 1000 | kHz | 19 | 17.092 | 12.517 | 17.32
| ACV0.110300 | ACV | 10 | 20 | 0.1 | V | mV | 300 | 1000 | kHz | 25.325 | 16.908 | 25.323 | 16.908
| ACV110300 | ACV | 10 | 20 | 1 | V | mV | 300 | 1000 | kHz | 533.13 | 7.3338 | 437.38 | 7.3336
| ACV0.00001200.001 | ACV | 20 | 100 | 1e-05 | V | mV | 0.001 | 0.04 | kHz | 5 | 0.23 | 4.6009 | 0.23351
| ACV0.0001200.001 | ACV | 20 | 100 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 5 | 0.23167 | 4.6011 | 0.23351
| ACV0.001200.001 | ACV | 20 | 100 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 5 | 0.23175 | 4.6195 | 0.23336
| ACV0.01200.001 | ACV | 20 | 100 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 6.5042 | 0.22022 | 6.5015 | 0.22022
| ACV0.1200.001 | ACV | 20 | 100 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 57.057 | 0.070854 | 57.056 | 0.070819
| ACV1200.001 | ACV | 20 | 100 | 1 | V | mV | 0.001 | 0.04 | kHz | 577.28 | 0.0075312 | 576.01 | 0.0075149
| ACV0.00001200.04 | ACV | 20 | 100 | 1e-05 | V | mV | 0.04 | 1 | kHz | 2.7 | 0.23 | 2.2758 | 0.23386
| ACV0.0001200.04 | ACV | 20 | 100 | 0.0001 | V | mV | 0.04 | 1 | kHz | 2.7 | 0.2318 | 2.276 | 0.23386
| ACV0.001200.04 | ACV | 20 | 100 | 0.001 | V | mV | 0.04 | 1 | kHz | 2.7 | 0.23188 | 2.2998 | 0.23364
| ACV0.01200.04 | ACV | 20 | 100 | 0.01 | V | mV | 0.04 | 1 | kHz | 4.7115 | 0.21577 | 4.7084 | 0.21577
| ACV0.1200.04 | ACV | 20 | 100 | 0.1 | V | mV | 0.04 | 1 | kHz | 56.896 | 0.062832 | 56.895 | 0.062796
| ACV1200.04 | ACV | 20 | 100 | 1 | V | mV | 0.04 | 1 | kHz | 577.27 | 0.0066108 | 576.14 | 0.0065945
| ACV0.00001201 | ACV | 20 | 100 | 1e-05 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8564 | 0.20787
| ACV0.0001201 | ACV | 20 | 100 | 0.0001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.8565 | 0.20787
| ACV0.001201 | ACV | 20 | 100 | 0.001 | V | mV | 1 | 20 | kHz | 4.9 | 0.21 | 4.882 | 0.20768
| ACV0.01201 | ACV | 20 | 100 | 0.01 | V | mV | 1 | 20 | kHz | 7.1093 | 0.19179 | 7.1056 | 0.1918
| ACV0.1201 | ACV | 20 | 100 | 0.1 | V | mV | 1 | 20 | kHz | 57.594 | 0.05585 | 57.592 | 0.055819
| ACV1201 | ACV | 20 | 100 | 1 | V | mV | 1 | 20 | kHz | 577.34 | 0.0058763 | 576.84 | 0.0058618
| ACV0.000012020 | ACV | 20 | 100 | 1e-05 | V | mV | 20 | 50 | kHz | 6.8 | 0.37 | 6.7323 | 0.36645
| ACV0.00012020 | ACV | 20 | 100 | 0.0001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36642 | 6.7324 | 0.36645
| ACV0.0012020 | ACV | 20 | 100 | 0.001 | V | mV | 20 | 50 | kHz | 6.8 | 0.36631 | 6.7495 | 0.36632
| ACV0.012020 | ACV | 20 | 100 | 0.01 | V | mV | 20 | 50 | kHz | 8.3578 | 0.35409 | 8.3518 | 0.35408
| ACV0.12020 | ACV | 20 | 100 | 0.1 | V | mV | 20 | 50 | kHz | 57.163 | 0.15069 | 57.159 | 0.15056
| ACV12020 | ACV | 20 | 100 | 1 | V | mV | 20 | 50 | kHz | 577.28 | 0.017106 | 575.94 | 0.01704
| ACV0.000012050 | ACV | 20 | 100 | 1e-05 | V | mV | 50 | 100 | kHz | 22 | 1.2 | 17.722 | 1.2358
| ACV0.00012050 | ACV | 20 | 100 | 0.0001 | V | mV | 50 | 100 | kHz | 22 | 1.2034 | 17.723 | 1.2358
| ACV0.0012050 | ACV | 20 | 100 | 0.001 | V | mV | 50 | 100 | kHz | 22 | 1.2034 | 17.724 | 1.2358
| ACV0.012050 | ACV | 20 | 100 | 0.01 | V | mV | 50 | 100 | kHz | 22 | 1.2047 | 18.315 | 1.2311
| ACV0.12050 | ACV | 20 | 100 | 0.1 | V | mV | 50 | 100 | kHz | 55.391 | 0.9727 | 55.373 | 0.9727
| ACV12050 | ACV | 20 | 100 | 1 | V | mV | 50 | 100 | kHz | 576.34 | 0.18075 | 568.41 | 0.18063
| ACV0.00011000.001 | ACV | 100 | 330 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 48 | 0.46 | 46.118 | 0.46325
| ACV0.0011000.001 | ACV | 100 | 330 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 48 | 0.46086 | 46.12 | 0.46325
| ACV0.011000.001 | ACV | 100 | 330 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 48 | 0.46129 | 46.341 | 0.46283
| ACV0.11000.001 | ACV | 100 | 330 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 66.332 | 0.42712 | 66.294 | 0.42698
| ACV11000.001 | ACV | 100 | 330 | 1 | V | mV | 0.001 | 0.04 | kHz | 573.42 | 0.11327 | 573.42 | 0.11294
| ACV0.00011000.04 | ACV | 100 | 330 | 0.0001 | V | mV | 0.04 | 1 | kHz | 25 | 0.46 | 23.034 | 0.46337
| ACV0.0011000.04 | ACV | 100 | 330 | 0.001 | V | mV | 0.04 | 1 | kHz | 25 | 0.46067 | 23.037 | 0.46337
| ACV0.011000.04 | ACV | 100 | 330 | 0.01 | V | mV | 0.04 | 1 | kHz | 25 | 0.46116 | 23.338 | 0.46274
| ACV0.11000.04 | ACV | 100 | 330 | 0.1 | V | mV | 0.04 | 1 | kHz | 49.032 | 0.41285 | 48.983 | 0.41271
| ACV11000.04 | ACV | 100 | 330 | 1 | V | mV | 0.04 | 1 | kHz | 571.92 | 0.096252 | 571.92 | 0.095921
| ACV0.00011001 | ACV | 100 | 330 | 0.0001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.059 | 0.69384
| ACV0.0011001 | ACV | 100 | 330 | 0.001 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.061 | 0.69384
| ACV0.011001 | ACV | 100 | 330 | 0.01 | V | mV | 1 | 20 | kHz | 23 | 0.7 | 23.291 | 0.69334
| ACV0.11001 | ACV | 100 | 330 | 0.1 | V | mV | 1 | 20 | kHz | 44.083 | 0.65012 | 43.963 | 0.65028
| ACV11001 | ACV | 100 | 330 | 1 | V | mV | 1 | 20 | kHz | 565.1 | 0.19691 | 565.03 | 0.19677
| ACV0.000110020 | ACV | 100 | 330 | 0.0001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.053 | 1.3879
| ACV0.00110020 | ACV | 100 | 330 | 0.001 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.054 | 1.3879
| ACV0.0110020 | ACV | 100 | 330 | 0.01 | V | mV | 20 | 50 | kHz | 23 | 1.4 | 23.188 | 1.3876
| ACV0.110020 | ACV | 100 | 330 | 0.1 | V | mV | 20 | 50 | kHz | 35.929 | 1.3608 | 35.878 | 1.3595
| ACV110020 | ACV | 100 | 330 | 1 | V | mV | 20 | 50 | kHz | 533.77 | 0.66047 | 533.57 | 0.66042
| ACV0.000110050 | ACV | 100 | 330 | 0.0001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.11 | 3.4656
| ACV0.00110050 | ACV | 100 | 330 | 0.001 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.111 | 3.4656
| ACV0.0110050 | ACV | 100 | 330 | 0.01 | V | mV | 50 | 100 | kHz | 23 | 3.5 | 23.164 | 3.4655
| ACV0.110050 | ACV | 100 | 330 | 0.1 | V | mV | 50 | 100 | kHz | 24.914 | 3.4942 | 28.927 | 3.4523
| ACV110050 | ACV | 100 | 330 | 1 | V | mV | 50 | 100 | kHz | 418.09 | 2.6787 | 417.63 | 2.6793
| ACV0.00013300.001 | ACV | 330 | 1000 | 0.0001 | V | mV | 0.001 | 0.04 | kHz | 50 | 0.46 | 46.107 | 0.46338
| ACV0.0013300.001 | ACV | 330 | 1000 | 0.001 | V | mV | 0.001 | 0.04 | kHz | 50 | 0.46054 | 46.108 | 0.46338
| ACV0.013300.001 | ACV | 330 | 1000 | 0.01 | V | mV | 0.001 | 0.04 | kHz | 50 | 0.46059 | 46.301 | 0.46322
| ACV0.13300.001 | ACV | 330 | 1000 | 0.1 | V | mV | 0.001 | 0.04 | kHz | 64.184 | 0.4486 | 64.129 | 0.44862
| ACV13300.001 | ACV | 330 | 1000 | 1 | V | mV | 0.001 | 0.04 | kHz | 564.15 | 0.20595 | 564.12 | 0.20589
| ACV0.00013300.04 | ACV | 330 | 1000 | 0.0001 | V | mV | 0.04 | 1 | kHz | 27 | 0.46 | 23.031 | 0.46343
| ACV0.0013300.04 | ACV | 330 | 1000 | 0.001 | V | mV | 0.04 | 1 | kHz | 27 | 0.46051 | 23.032 | 0.46343
| ACV0.013300.04 | ACV | 330 | 1000 | 0.01 | V | mV | 0.04 | 1 | kHz | 27 | 0.46056 | 23.29 | 0.46321
| ACV0.13300.04 | ACV | 330 | 1000 | 0.1 | V | mV | 0.04 | 1 | kHz | 45.921 | 0.44399 | 45.855 | 0.44402
| ACV13300.04 | ACV | 330 | 1000 | 1 | V | mV | 0.04 | 1 | kHz | 562.27 | 0.1928 | 562.23 | 0.19274
| ACV0.00013301 | ACV | 330 | 1000 | 0.0001 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.056 | 0.69387
| ACV0.0013301 | ACV | 330 | 1000 | 0.001 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.063 | 0.69386
| ACV0.013301 | ACV | 330 | 1000 | 0.01 | V | mV | 1 | 10 | kHz | 23 | 0.7 | 23.256 | 0.69369
| ACV0.13301 | ACV | 330 | 1000 | 0.1 | V | mV | 1 | 10 | kHz | 40.883 | 0.68212 | 41.185 | 0.67806
| ACV13301 | ACV | 330 | 1000 | 1 | V | mV | 1 | 10 | kHz | 547.47 | 0.37317 | 547.4 | 0.3731
| ACA0.000000000100.001 | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0.001 | 0.02 | kHz | 0.077 | 4200 | 0.076632 | 4198.9
| ACA0.00000000100.001 | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0.001 | 0.02 | kHz | 0.077 | 4200 | 0.076633 | 4198.9
| ACA0.0000000100.001 | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0.001 | 0.02 | kHz | 0.077 | 4200 | 0.076839 | 4197.2
| ACA0.000000100.001 | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 0.095326 | 4045.9 | 0.095288 | 4045.8
| ACA0.00000100.001 | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 0.58121 | 1804 | 0.58119 | 1803.1
| ACA0.0000100.001 | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 5.7739 | 210.05 | 5.7723 | 209.55
| ACA0.000100.001 | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0.001 | 0.02 | kHz | 57.736 | 21.197 | 57.734 | 20.994
| ACA0.00100.001 | ACA | 0 | 0.0001 | 0.001 | A | µA | 0.001 | 0.02 | kHz | 577.36 | 2.1805 | 577.35 | 2.0995
| ACA0.0100.001 | ACA | 0 | 0.0001 | 0.01 | A | µA | 0.001 | 0.02 | kHz | 5773.6 | 0.24237 | 5773.6 | 0.20995
| ACA0.100.001 | ACA | 0 | 0.0001 | 0.1 | A | µA | 0.001 | 0.02 | kHz | 57736 | 0.033964 | 57736 | 0.020995
| ACA100.001 | ACA | 0 | 0.0001 | 1 | A | µA | 0.001 | 0.02 | kHz | 577360 | 0.0072871 | 577360 | 0.002099
| ACA0.000000000100.02 | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0.02 | 0.045 | kHz | 0.038 | 1700 | 0.034644 | 1732.1
| ACA0.00000000100.02 | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0.02 | 0.045 | kHz | 0.038 | 1700 | 0.034648 | 1732.1
| ACA0.0000000100.02 | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 0.038 | 1700 | 0.034998 | 1729.4
| ACA0.000000100.02 | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 0.062399 | 1534.5 | 0.062349 | 1534
| ACA0.00000100.02 | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 0.57594 | 378.74 | 0.57592 | 377.44
| ACA0.0000100.02 | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.7734 | 39.656 | 5.77 | 39.099
| ACA0.000100.02 | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 57.736 | 4.1759 | 57.732 | 3.953
| ACA0.00100.02 | ACA | 0 | 0.0001 | 0.001 | A | µA | 0.02 | 0.045 | kHz | 577.36 | 0.50118 | 577.35 | 0.41202
| ACA0.0100.02 | ACA | 0 | 0.0001 | 0.01 | A | µA | 0.02 | 0.045 | kHz | 5773.6 | 0.083554 | 5773.5 | 0.047889
| ACA0.100.02 | ACA | 0 | 0.0001 | 0.1 | A | µA | 0.02 | 0.045 | kHz | 57736 | 0.02173 | 57736 | 0.0074638
| ACA100.02 | ACA | 0 | 0.0001 | 1 | A | µA | 0.02 | 0.045 | kHz | 577360 | 0.007523 | 577360 | 0.0018174
| ACA0.000000000100.045 | ACA | 0 | 0.0001 | 1e-10 | A | µA | 0.045 | 5 | kHz | 0.037 | 700 | 0.035478 | 705.46
| ACA0.00000000100.045 | ACA | 0 | 0.0001 | 1e-09 | A | µA | 0.045 | 5 | kHz | 0.037 | 701.57 | 0.035484 | 705.45
| ACA0.0000000100.045 | ACA | 0 | 0.0001 | 1e-08 | A | µA | 0.045 | 5 | kHz | 0.037 | 703.27 | 0.035899 | 702.95
| ACA0.000000100.045 | ACA | 0 | 0.0001 | 1e-07 | A | µA | 0.045 | 5 | kHz | 0.066315 | 546.37 | 0.066275 | 544.96
| ACA0.00000100.045 | ACA | 0 | 0.0001 | 1e-06 | A | µA | 0.045 | 5 | kHz | 0.57804 | 91.593 | 0.57802 | 90.228
| ACA0.0000100.045 | ACA | 0 | 0.0001 | 1e-05 | A | µA | 0.045 | 5 | kHz | 5.7736 | 9.7715 | 5.7728 | 9.2143
| ACA0.000100.045 | ACA | 0 | 0.0001 | 0.0001 | A | µA | 0.045 | 5 | kHz | 57.736 | 1.1862 | 57.735 | 0.96332
| ACA0.00100.045 | ACA | 0 | 0.0001 | 0.001 | A | µA | 0.045 | 5 | kHz | 577.36 | 0.20221 | 577.35 | 0.11305
| ACA0.0100.045 | ACA | 0 | 0.0001 | 0.01 | A | µA | 0.045 | 5 | kHz | 5773.6 | 0.053657 | 5773.6 | 0.017992
| ACA0.100.045 | ACA | 0 | 0.0001 | 0.1 | A | µA | 0.045 | 5 | kHz | 57736 | 0.01874 | 57736 | 0.0044741
| ACA100.045 | ACA | 0 | 0.0001 | 1 | A | µA | 0.045 | 5 | kHz | 577360 | 0.0072242 | 577360 | 0.0015173
| ACA0.00000000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0.001 | 0.02 | kHz | 0.25 | 4600 | 0.23095 | 4618.8
| ACA0.0000000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.001 | 0.02 | kHz | 0.25 | 4600 | 0.23098 | 4618.8
| ACA0.000000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.001 | 0.02 | kHz | 0.25 | 4600.3 | 0.23363 | 4616.5
| ACA0.00000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 0.4599 | 4424.7 | 0.45952 | 4424.6
| ACA0.0000010.00010.001 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 5.6235 | 1918.3 | 5.6233 | 1917.2
| ACA0.00000000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0.02 | 0.045 | kHz | 0.27 | 1700 | 0.23094 | 1732.1
| ACA0.0000000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.02 | 0.045 | kHz | 0.27 | 1700 | 0.231 | 1732.1
| ACA0.000000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 0.27 | 1700 | 0.23546 | 1728.5
| ACA0.00000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 0.5563 | 1491.3 | 0.55586 | 1490.6
| ACA0.0000010.00010.02 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 5.7533 | 346.75 | 5.7532 | 345.32
| ACA0.00000000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0.045 | 0.1 | kHz | 0.24 | 690 | 0.23094 | 693.02
| ACA0.0000000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.045 | 0.1 | kHz | 0.24 | 690 | 0.23102 | 692.99
| ACA0.000000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.045 | 0.1 | kHz | 0.24 | 690 | 0.23691 | 689
| ACA0.00000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.045 | 0.1 | kHz | 0.60249 | 489.7 | 0.6021 | 487.94
| ACA0.0000010.00010.045 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.045 | 0.1 | kHz | 5.7742 | 74.832 | 5.7741 | 73.333
| ACA0.00000000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 0.1 | 5 | kHz | 0.25 | 350 | 0.23001 | 364.09
| ACA0.0000000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 0.1 | 5 | kHz | 0.25 | 354.91 | 0.23009 | 364.09
| ACA0.000000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 0.1 | 5 | kHz | 0.25 | 359.01 | 0.23663 | 360.5
| ACA0.00000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 0.1 | 5 | kHz | 0.61486 | 217.08 | 0.61454 | 214.58
| ACA0.0000010.00010.1 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 0.1 | 5 | kHz | 5.7771 | 28.936 | 5.777 | 27.424
| ACA0.00000000010.00015 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 5 | 20 | kHz | 0.25 | 710 | 0.23471 | 715.83
| ACA0.0000000010.00015 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 5 | 20 | kHz | 0.25 | 711.93 | 0.23479 | 715.8
| ACA0.000000010.00015 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 5 | 20 | kHz | 0.25 | 712.22 | 0.24058 | 711.87
| ACA0.00000010.00015 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 5 | 20 | kHz | 0.6031 | 511.67 | 0.60272 | 509.95
| ACA0.0000010.00015 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 5 | 20 | kHz | 5.7741 | 79.21 | 5.774 | 77.713
| ACA0.00000000010.000120 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 20 | 50 | kHz | 0.47 | 4800 | 0.46627 | 4761.7
| ACA0.0000000010.000120 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 20 | 50 | kHz | 0.47 | 4761.6 | 0.46636 | 4761.6
| ACA0.000000010.000120 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 20 | 50 | kHz | 0.47 | 4759.9 | 0.46837 | 4760
| ACA0.00000010.000120 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 20 | 50 | kHz | 0.64627 | 4616 | 0.64414 | 4616.1
| ACA0.0000010.000120 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 20 | 50 | kHz | 5.636 | 2159.7 | 5.6349 | 2155.3
| ACA0.00000000010.000150 | ACA | 0.0001 | 0.001 | 1e-10 | A | µA | 50 | 100 | kHz | 1.9 | 6500 | 1.7592 | 6558
| ACA0.0000000010.000150 | ACA | 0.0001 | 0.001 | 1e-09 | A | µA | 50 | 100 | kHz | 1.9 | 6531.9 | 1.7594 | 6558.1
| ACA0.000000010.000150 | ACA | 0.0001 | 0.001 | 1e-08 | A | µA | 50 | 100 | kHz | 1.9 | 6537.4 | 1.7607 | 6558
| ACA0.00000010.000150 | ACA | 0.0001 | 0.001 | 1e-07 | A | µA | 50 | 100 | kHz | 1.9 | 6516.3 | 1.8354 | 6506.9
| ACA0.0000010.000150 | ACA | 0.0001 | 0.001 | 1e-06 | A | µA | 50 | 100 | kHz | 5.8456 | 4362.3 | 5.8321 | 4309.3
| ACA0.000000010.0010.001 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0.001 | 0.02 | kHz | 2.5 | 4600 | 2.3096 | 4618.8
| ACA0.00000010.0010.001 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 2.5 | 4600 | 2.3101 | 4618.8
| ACA0.0000010.0010.001 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 2.5 | 4600.7 | 2.3373 | 4616.5
| ACA0.000010.0010.001 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 4.6177 | 4424.3 | 4.5991 | 4424.5
| ACA0.00010.0010.001 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.001 | 0.02 | kHz | 56.243 | 1921.2 | 56.235 | 1917.8
| ACA0.000000010.0010.02 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 2.7 | 1700 | 2.3095 | 1732.1
| ACA0.00000010.0010.02 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 2.7 | 1700 | 2.311 | 1732
| ACA0.0000010.0010.02 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 2.7 | 1700 | 2.3566 | 1728.5
| ACA0.000010.0010.02 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.585 | 1492.6 | 5.5626 | 1490.9
| ACA0.00010.0010.02 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 57.541 | 350.7 | 57.534 | 346.11
| ACA0.000000010.0010.045 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0.045 | 0.1 | kHz | 2.4 | 690 | 2.3095 | 693.04
| ACA0.00000010.0010.045 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.045 | 0.1 | kHz | 2.4 | 690 | 2.3107 | 693.04
| ACA0.0000010.0010.045 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.045 | 0.1 | kHz | 2.4 | 690 | 2.3717 | 689.14
| ACA0.000010.0010.045 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.045 | 0.1 | kHz | 6.0458 | 494.17 | 6.0252 | 488.84
| ACA0.00010.0010.045 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.045 | 0.1 | kHz | 57.749 | 79.009 | 57.742 | 74.17
| ACA0.000000010.0010.1 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 0.1 | 5 | kHz | 2.5 | 350 | 2.2952 | 364.62
| ACA0.00000010.0010.1 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 0.1 | 5 | kHz | 2.5 | 355.75 | 2.2965 | 364.68
| ACA0.0000010.0010.1 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 0.1 | 5 | kHz | 2.5 | 362.33 | 2.3641 | 361.33
| ACA0.000010.0010.1 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 0.1 | 5 | kHz | 6.1659 | 223.94 | 6.1472 | 216.14
| ACA0.00010.0010.1 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 0.1 | 5 | kHz | 57.778 | 33.176 | 57.771 | 28.291
| ACA0.000000010.0015 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 5 | 20 | kHz | 2.5 | 710 | 2.2891 | 721.65
| ACA0.00000010.0015 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 5 | 20 | kHz | 2.5 | 712.46 | 2.2904 | 721.65
| ACA0.0000010.0015 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 5 | 20 | kHz | 2.5 | 716.59 | 2.3512 | 717.7
| ACA0.000010.0015 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 5 | 20 | kHz | 6.0251 | 518.68 | 6.0044 | 513.52
| ACA0.00010.0015 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 5 | 20 | kHz | 57.745 | 83.689 | 57.738 | 78.853
| ACA0.000000010.00120 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 20 | 50 | kHz | 5.1 | 4700 | 4.5719 | 4747.9
| ACA0.00000010.00120 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 20 | 50 | kHz | 5.1 | 4739.1 | 4.5722 | 4747.9
| ACA0.0000010.00120 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 20 | 50 | kHz | 5.1 | 4739.9 | 4.5921 | 4746.3
| ACA0.000010.00120 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 20 | 50 | kHz | 6.3814 | 4600.8 | 6.366 | 4600.8
| ACA0.00010.00120 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 20 | 50 | kHz | 56.354 | 2143.3 | 56.346 | 2139.9
| ACA0.000000010.00150 | ACA | 0.001 | 0.01 | 1e-08 | A | µA | 50 | 100 | kHz | 19 | 6500 | 17.151 | 6602.1
| ACA0.00000010.00150 | ACA | 0.001 | 0.01 | 1e-07 | A | µA | 50 | 100 | kHz | 19 | 6531.1 | 17.153 | 6602.1
| ACA0.0000010.00150 | ACA | 0.001 | 0.01 | 1e-06 | A | µA | 50 | 100 | kHz | 19 | 6534.1 | 17.164 | 6601.7
| ACA0.000010.00150 | ACA | 0.001 | 0.01 | 1e-05 | A | µA | 50 | 100 | kHz | 19 | 6550 | 17.919 | 6548
| ACA0.00010.00150 | ACA | 0.001 | 0.01 | 0.0001 | A | µA | 50 | 100 | kHz | 58.261 | 4342.2 | 58.146 | 4318.8
| ACA0.00000010.010.001 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0.001 | 0.02 | kHz | 25 | 4600 | 23.095 | 4618.8
| ACA0.0000010.010.001 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.001 | 0.02 | kHz | 25 | 4600 | 23.099 | 4618.8
| ACA0.000010.010.001 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.001 | 0.02 | kHz | 25 | 4600.4 | 23.366 | 4616.5
| ACA0.00010.010.001 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.001 | 0.02 | kHz | 46.067 | 4424.4 | 45.964 | 4424.6
| ACA0.0010.010.001 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.001 | 0.02 | kHz | 562.38 | 1919.1 | 562.34 | 1917.3
| ACA0.00000010.010.02 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 27 | 1700 | 23.095 | 1732.1
| ACA0.0000010.010.02 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 27 | 1700 | 23.102 | 1732.1
| ACA0.000010.010.02 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 27 | 1700 | 23.554 | 1728.5
| ACA0.00010.010.02 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 55.716 | 1491.6 | 55.6 | 1490.7
| ACA0.0010.010.02 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.02 | 0.045 | kHz | 575.37 | 347.9 | 575.33 | 345.55
| ACA0.00000010.010.045 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0.045 | 0.1 | kHz | 24 | 690 | 23.094 | 693.05
| ACA0.0000010.010.045 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.045 | 0.1 | kHz | 24 | 690 | 23.103 | 693.03
| ACA0.000010.010.045 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.045 | 0.1 | kHz | 24 | 690 | 23.701 | 689.05
| ACA0.00010.010.045 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.045 | 0.1 | kHz | 60.336 | 490.93 | 60.227 | 488.21
| ACA0.0010.010.045 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.045 | 0.1 | kHz | 577.45 | 76.063 | 577.41 | 73.582
| ACA0.00000010.010.1 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 0.1 | 5 | kHz | 25 | 350 | 22.95 | 364.61
| ACA0.0000010.010.1 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 0.1 | 5 | kHz | 25 | 355.17 | 22.961 | 364.62
| ACA0.000010.010.1 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 0.1 | 5 | kHz | 25 | 360.06 | 23.625 | 361.08
| ACA0.00010.010.1 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 0.1 | 5 | kHz | 61.548 | 219.19 | 61.449 | 215.18
| ACA0.0010.010.1 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 0.1 | 5 | kHz | 577.74 | 30.199 | 577.71 | 27.695
| ACA0.00000010.015 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 5 | 20 | kHz | 25 | 710 | 22.891 | 721.64
| ACA0.0000010.015 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 5 | 20 | kHz | 25 | 712.1 | 22.9 | 721.62
| ACA0.000010.015 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 5 | 20 | kHz | 25 | 715.16 | 23.497 | 717.59
| ACA0.00010.015 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 5 | 20 | kHz | 60.128 | 515.53 | 60.018 | 512.89
| ACA0.0010.015 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 5 | 20 | kHz | 577.41 | 80.741 | 577.37 | 78.263
| ACA0.00000010.0120 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 20 | 50 | kHz | 51 | 4700 | 45.602 | 4746.9
| ACA0.0000010.0120 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 20 | 50 | kHz | 51 | 4736.8 | 45.606 | 4746.9
| ACA0.000010.0120 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 20 | 50 | kHz | 51 | 4738 | 45.823 | 4745.2
| ACA0.00010.0120 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 20 | 50 | kHz | 64.008 | 4598.6 | 63.621 | 4599.4
| ACA0.0010.0120 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 20 | 50 | kHz | 563.71 | 2145.2 | 563.49 | 2139.2
| ACA0.00000010.0150 | ACA | 0.01 | 0.1 | 1e-07 | A | µA | 50 | 100 | kHz | 190 | 6500 | 171.51 | 6602.2
| ACA0.0000010.0150 | ACA | 0.01 | 0.1 | 1e-06 | A | µA | 50 | 100 | kHz | 190 | 6538.1 | 171.63 | 6602.5
| ACA0.000010.0150 | ACA | 0.01 | 0.1 | 1e-05 | A | µA | 50 | 100 | kHz | 190 | 6565.8 | 179.15 | 6551.2
| ACA0.00010.0150 | ACA | 0.01 | 0.1 | 0.0001 | A | µA | 50 | 100 | kHz | 582.09 | 4394.2 | 581.36 | 4329.2
| ACA0.0010.0150 | ACA | 0.01 | 0.1 | 0.001 | A | µA | 50 | 100 | kHz | 5772.5 | 669.65 | 5772.4 | 620.3
| ACA0.0000010.10.001 | ACA | 0.1 | 1 | 1e-06 | A | mA | 0.001 | 0.02 | kHz | 0.25 | 4.6 | 0.23096 | 4.6188
| ACA0.000010.10.001 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.001 | 0.02 | kHz | 0.25 | 4.6 | 0.231 | 4.6188
| ACA0.00010.10.001 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.001 | 0.02 | kHz | 0.25 | 4.6003 | 0.23368 | 4.6165
| ACA0.0010.10.001 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.001 | 0.02 | kHz | 0.46103 | 4.4237 | 0.45975 | 4.4244
| ACA0.010.10.001 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.001 | 0.02 | kHz | 5.6242 | 1.9179 | 5.6235 | 1.9171
| ACA0.0000010.10.02 | ACA | 0.1 | 1 | 1e-06 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.23096 | 1.8476
| ACA0.000010.10.02 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.23102 | 1.8476
| ACA0.00010.10.02 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.02 | 0.045 | kHz | 0.28 | 1.8 | 0.23545 | 1.844
| ACA0.0010.10.02 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.02 | 0.045 | kHz | 0.55279 | 1.6061 | 0.55118 | 1.6064
| ACA0.010.10.02 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.02 | 0.045 | kHz | 5.7507 | 0.38797 | 5.75 | 0.38676
| ACA0.0000010.10.045 | ACA | 0.1 | 1 | 1e-06 | A | mA | 0.045 | 0.1 | kHz | 0.24 | 0.92 | 0.23091 | 0.92432
| ACA0.000010.10.045 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.045 | 0.1 | kHz | 0.24 | 0.92 | 0.23101 | 0.92427
| ACA0.00010.10.045 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.045 | 0.1 | kHz | 0.24 | 0.92 | 0.23664 | 0.92018
| ACA0.0010.10.045 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 0.59376 | 0.70044 | 0.59215 | 0.69986
| ACA0.010.10.045 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 5.7717 | 0.11872 | 5.771 | 0.11743
| ACA0.0000010.10.1 | ACA | 0.1 | 1 | 1e-06 | A | mA | 0.1 | 5 | kHz | 0.24 | 1.2 | 0.22816 | 1.1855
| ACA0.000010.10.1 | ACA | 0.1 | 1 | 1e-05 | A | mA | 0.1 | 5 | kHz | 0.24 | 1.2 | 0.2283 | 1.1854
| ACA0.00010.10.1 | ACA | 0.1 | 1 | 0.0001 | A | mA | 0.1 | 5 | kHz | 0.24 | 1.2 | 0.23356 | 1.1814
| ACA0.0010.10.1 | ACA | 0.1 | 1 | 0.001 | A | mA | 0.1 | 5 | kHz | 0.58062 | 0.9487 | 0.57901 | 0.94846
| ACA0.010.10.1 | ACA | 0.1 | 1 | 0.01 | A | mA | 0.1 | 5 | kHz | 5.767 | 0.17953 | 5.7663 | 0.17826
| ACA0.0000010.15 | ACA | 0.1 | 1 | 1e-06 | A | mA | 5 | 20 | kHz | 0.27 | 3.5 | 0.22313 | 3.5466
| ACA0.000010.15 | ACA | 0.1 | 1 | 1e-05 | A | mA | 5 | 20 | kHz | 0.27 | 3.5331 | 0.22319 | 3.5466
| ACA0.00010.15 | ACA | 0.1 | 1 | 0.0001 | A | mA | 5 | 20 | kHz | 0.27 | 3.5338 | 0.22642 | 3.5438
| ACA0.0010.15 | ACA | 0.1 | 1 | 0.001 | A | mA | 5 | 20 | kHz | 0.48562 | 3.329 | 0.48418 | 3.3297
| ACA0.010.15 | ACA | 0.1 | 1 | 0.01 | A | mA | 5 | 20 | kHz | 5.6818 | 1.2155 | 5.6811 | 1.2146
| ACA0.0000010.120 | ACA | 0.1 | 1 | 1e-06 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44152 | 11.792
| ACA0.000010.120 | ACA | 0.1 | 1 | 1e-05 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44162 | 11.792
| ACA0.00010.120 | ACA | 0.1 | 1 | 0.0001 | A | mA | 20 | 50 | kHz | 0.47 | 12 | 0.44312 | 11.791
| ACA0.0010.120 | ACA | 0.1 | 1 | 0.001 | A | mA | 20 | 50 | kHz | 0.53612 | 11.934 | 0.5524 | 11.697
| ACA0.010.120 | ACA | 0.1 | 1 | 0.01 | A | mA | 20 | 50 | kHz | 5.1678 | 8.3949 | 5.1615 | 8.3731
| ACA0.0000110.001 | ACA | 1 | 10 | 1e-05 | A | mA | 0.001 | 0.05 | kHz | 0.55 | 0.12 | 0.51999 | 0.093087
| ACA0.000110.001 | ACA | 1 | 10 | 0.0001 | A | mA | 0.001 | 0.05 | kHz | 0.55 | 0.12 | 0.52375 | 0.092861
| ACA0.00110.001 | ACA | 1 | 10 | 0.001 | A | mA | 0.001 | 0.05 | kHz | 0.75909 | 0.099091 | 0.76542 | 0.079775
| ACA0.0110.001 | ACA | 1 | 10 | 0.01 | A | mA | 0.001 | 0.05 | kHz | 5.7992 | 0.016313 | 5.7916 | 0.016333
| ACA0.110.001 | ACA | 1 | 10 | 0.1 | A | mA | 0.001 | 0.05 | kHz | 57.741 | 0.0016634 | 57.738 | 0.0016635
| ACA0.0000110.05 | ACA | 1 | 10 | 1e-05 | A | mA | 0.05 | 1 | kHz | 0.46 | 0.17 | 0.49897 | 0.12063
| ACA0.000110.05 | ACA | 1 | 10 | 0.0001 | A | mA | 0.05 | 1 | kHz | 0.46 | 0.16806 | 0.50269 | 0.12039
| ACA0.00110.05 | ACA | 1 | 10 | 0.001 | A | mA | 0.05 | 1 | kHz | 0.72094 | 0.14191 | 0.74419 | 0.10576
| ACA0.0110.05 | ACA | 1 | 10 | 0.01 | A | mA | 0.05 | 1 | kHz | 5.7925 | 0.023676 | 5.7849 | 0.023706
| ACA0.110.05 | ACA | 1 | 10 | 0.1 | A | mA | 0.05 | 1 | kHz | 57.74 | 0.002428 | 57.737 | 0.0024282
| ACA0.0000111 | ACA | 1 | 10 | 1e-05 | A | mA | 1 | 5 | kHz | 1.4 | 0.66 | 1.326 | 0.66218
| ACA0.000111 | ACA | 1 | 10 | 0.0001 | A | mA | 1 | 5 | kHz | 1.4 | 0.66168 | 1.3279 | 0.66196
| ACA0.00111 | ACA | 1 | 10 | 0.001 | A | mA | 1 | 5 | kHz | 1.4676 | 0.64572 | 1.4594 | 0.64665
| ACA0.0111 | ACA | 1 | 10 | 0.01 | A | mA | 1 | 5 | kHz | 5.82 | 0.34986 | 5.8127 | 0.35015
| ACA0.111 | ACA | 1 | 10 | 0.1 | A | mA | 1 | 5 | kHz | 57.725 | 0.044886 | 57.722 | 0.044889
| ACA0.00001100.001 | ACA | 10 | 20 | 1e-05 | A | mA | 0.001 | 0.05 | kHz | 0.92 | 0.49 | 4.6374 | 0.081764
| ACA0.0001100.001 | ACA | 10 | 20 | 0.0001 | A | mA | 0.001 | 0.05 | kHz | 0.92 | 0.49 | 4.6379 | 0.081759
| ACA0.001100.001 | ACA | 10 | 20 | 0.001 | A | mA | 0.001 | 0.05 | kHz | 0.92 | 0.49 | 4.6724 | 0.081363
| ACA0.01100.001 | ACA | 10 | 20 | 0.01 | A | mA | 0.001 | 0.05 | kHz | 5.1799 | 0.27701 | 7.3622 | 0.058218
| ACA0.1100.001 | ACA | 10 | 20 | 0.1 | A | mA | 0.001 | 0.05 | kHz | 57.914 | 0.0082611 | 57.911 | 0.0082616
| ACA0.00001100.05 | ACA | 10 | 20 | 1e-05 | A | mA | 0.05 | 1 | kHz | 0 | 0.5 | 2.1832 | 0.11356
| ACA0.0001100.05 | ACA | 10 | 20 | 0.0001 | A | mA | 0.05 | 1 | kHz | 0 | 0.5 | 2.1841 | 0.11354
| ACA0.001100.05 | ACA | 10 | 20 | 0.001 | A | mA | 0.05 | 1 | kHz | 0 | 0.5 | 2.2467 | 0.11228
| ACA0.01100.05 | ACA | 10 | 20 | 0.01 | A | mA | 0.05 | 1 | kHz | 3.3354 | 0.33323 | 6.0285 | 0.063256
| ACA0.1100.05 | ACA | 10 | 20 | 0.1 | A | mA | 0.05 | 1 | kHz | 57.758 | 0.0076264 | 57.755 | 0.0076268
| ACA0.00001101 | ACA | 10 | 20 | 1e-05 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.0702 | 0.65987
| ACA0.0001101 | ACA | 10 | 20 | 0.0001 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.0706 | 0.65986
| ACA0.001101 | ACA | 10 | 20 | 0.001 | A | mA | 1 | 5 | kHz | 0 | 0.95 | 1.1025 | 0.65885
| ACA0.01101 | ACA | 10 | 20 | 0.01 | A | mA | 1 | 5 | kHz | 0.20999 | 0.9395 | 3.8087 | 0.57918
| ACA0.1101 | ACA | 10 | 20 | 0.1 | A | mA | 1 | 5 | kHz | 57.017 | 0.12296 | 57.014 | 0.12297
| Frequency0.0000000010 | Frequency | 0 | 0.04 | 1e-09 | kHz | Hz |  |  |  | 0 | 6e-07 | 1.0954e-11 | 5.7735e-07
| Frequency0.000000010 | Frequency | 0 | 0.04 | 1e-08 | kHz | Hz |  |  |  | 0 | 6e-07 | 2.9716e-11 | 5.7735e-07
| Frequency0.00000010 | Frequency | 0 | 0.04 | 1e-07 | kHz | Hz |  |  |  | 0 | 6e-07 | 2.9505e-09 | 5.7728e-07
| Frequency0.0000010 | Frequency | 0 | 0.04 | 1e-06 | kHz | Hz |  |  |  | 2.2344e-07 | 5.9441e-07 | 2.4523e-07 | 5.714e-07
| Frequency0.000010 | Frequency | 0 | 0.04 | 1e-05 | kHz | Hz |  |  |  | 5.3365e-06 | 4.6659e-07 | 5.3408e-06 | 4.616e-07
| Frequency0.00010 | Frequency | 0 | 0.04 | 0.0001 | kHz | Hz |  |  |  | 5.7625e-05 | 1.1396e-07 | 5.7277e-05 | 1.1396e-07
| Frequency0.0010 | Frequency | 0 | 0.04 | 0.001 | kHz | Hz |  |  |  | 0.00057734 | 1.1831e-08 | 0.00057689 | 1.1831e-08
| Frequency0.010 | Frequency | 0 | 0.04 | 0.01 | kHz | Hz |  |  |  | 0.0057736 | 1.1836e-09 | 0.0057731 | 1.1836e-09
| Frequency0.10 | Frequency | 0 | 0.04 | 0.1 | kHz | Hz |  |  |  | 0.057736 | 1.1836e-10 | 0.057735 | 1.1836e-10
| Frequency10 | Frequency | 0 | 0.04 | 1 | kHz | Hz |  |  |  | 0.57736 | 1.1836e-11 | 0.57735 | 1.1836e-11
| Frequency100 | Frequency | 0 | 0.04 | 10 | kHz | Hz |  |  |  | 5.7736 | 1.1836e-12 | 5.7736 | 1.1836e-12
| Frequency0.00000010.04 | Frequency | 0.04 | 1 | 1e-07 | kHz | Hz |  |  |  | 2.9e-07 | 0.00013 | 2.5345e-07 | 0.00012897
| Frequency0.0000010.04 | Frequency | 0.04 | 1 | 1e-06 | kHz | Hz |  |  |  | 6.5701e-07 | 0.00012894 | 5.9818e-07 | 0.00012907
| Frequency0.000010.04 | Frequency | 0.04 | 1 | 1e-05 | kHz | Hz |  |  |  | 2.609e-05 | 0.00013885 | 2.5942e-05 | 0.00013918
| Frequency0.00010.04 | Frequency | 0.04 | 1 | 0.0001 | kHz | Hz |  |  |  | 0.00039689 | 0.0005557 | 0.00039681 | 0.00055588
| Frequency0.0010.04 | Frequency | 0.04 | 1 | 0.001 | kHz | Hz |  |  |  | 0.0039872 | 0.0054141 | 0.0039872 | 0.0054141
| Frequency0.010.04 | Frequency | 0.04 | 1 | 0.01 | kHz | Hz |  |  |  | 0.055949 | 0.054127 | 0.055949 | 0.054127
| Frequency0.0000011 | Frequency | 1 | 10 | 1e-06 | kHz | Hz |  |  |  | 4.4e-09 | 0.00013 | 7.4137e-10 | 0.00012922
| Frequency0.000011 | Frequency | 1 | 10 | 1e-05 | kHz | Hz |  |  |  | 1.9518e-07 | 0.0001292 | 1.3837e-07 | 0.00012921
| Frequency0.00011 | Frequency | 1 | 10 | 0.0001 | kHz | Hz |  |  |  | 1.3574e-05 | 0.00012799 | 1.3571e-05 | 0.00012799
| Frequency0.0011 | Frequency | 1 | 10 | 0.001 | kHz | Hz |  |  |  | 0.00068566 | 9.1517e-05 | 0.00068565 | 9.1518e-05
| Frequency0.011 | Frequency | 1 | 10 | 0.01 | kHz | Hz |  |  |  | 0.0057844 | 1.571e-05 | 0.0057844 | 1.571e-05
| Frequency0.11 | Frequency | 1 | 10 | 0.1 | kHz | Hz |  |  |  | 0.057734 | 1.5904e-06 | 0.05772 | 1.5904e-06
| Frequency0.0000110 | Frequency | 10 | 100 | 1e-05 | kHz | Hz |  |  |  | 4.4e-08 | 7.4137e-09 | 7.4137e-09 | 0.00012922
| Frequency0.000110 | Frequency | 10 | 100 | 0.0001 | kHz | Hz |  |  |  | 1.9518e-06 | 1.3837e-06 | 1.3837e-06 | 0.00012921
| Frequency0.00110 | Frequency | 10 | 100 | 0.001 | kHz | Hz |  |  |  | 0.00013574 | 0.00013571 | 0.00013571 | 0.00012799
| Frequency0.0110 | Frequency | 10 | 100 | 0.01 | kHz | Hz |  |  |  | 0.0068566 | 0.0068565 | 0.0068565 | 9.1518e-05
| Frequency0.110 | Frequency | 10 | 100 | 0.1 | kHz | Hz |  |  |  | 0.057844 | 0.057844 | 0.057844 | 1.571e-05
| Frequency110 | Frequency | 10 | 100 | 1 | kHz | Hz |  |  |  | 0.57734 | 0.5772 | 0.5772 | 1.5904e-06
| Frequency1010 | Frequency | 10 | 100 | 10 | kHz | Hz |  |  |  | 5.7736 | 5.7734 | 5.7734 | 1.5906e-07
| Frequency0.0001100 | Frequency | 100 | 1000 | 0.0001 | kHz | Hz |  |  |  | 4.4e-07 | 0.00013 | 7.4137e-08 | 0.00012922
| Frequency0.001100 | Frequency | 100 | 1000 | 0.001 | kHz | Hz |  |  |  | 1.9518e-05 | 0.0001292 | 1.3837e-05 | 0.00012921
| Frequency0.01100 | Frequency | 100 | 1000 | 0.01 | kHz | Hz |  |  |  | 0.0013574 | 0.00012799 | 0.0013571 | 0.00012799
| Frequency0.1100 | Frequency | 100 | 1000 | 0.1 | kHz | Hz |  |  |  | 0.068566 | 9.1517e-05 | 0.068565 | 9.1518e-05
| Frequency1100 | Frequency | 100 | 1000 | 1 | kHz | Hz |  |  |  | 0.57844 | 1.571e-05 | 0.57844 | 1.571e-05
| Frequency10100 | Frequency | 100 | 1000 | 10 | kHz | Hz |  |  |  | 5.7734 | 1.5904e-06 | 5.772 | 1.5904e-06
| Frequency0.0011000 | Frequency | 1000 | 10000 | 0.001 | kHz | Hz |  |  |  | 4.4e-06 | 0.00013 | 7.4137e-07 | 0.00012922
| Frequency0.011000 | Frequency | 1000 | 10000 | 0.01 | kHz | Hz |  |  |  | 0.00019518 | 0.0001292 | 0.00013837 | 0.00012921
| Frequency0.11000 | Frequency | 1000 | 10000 | 0.1 | kHz | Hz |  |  |  | 0.013574 | 0.00012799 | 0.013571 | 0.00012799
| Frequency11000 | Frequency | 1000 | 10000 | 1 | kHz | Hz |  |  |  | 0.68566 | 9.1517e-05 | 0.68565 | 9.1518e-05
| Frequency101000 | Frequency | 1000 | 10000 | 10 | kHz | Hz |  |  |  | 5.7844 | 1.571e-05 | 5.7844 | 1.571e-05
| Frequency1001000 | Frequency | 1000 | 10000 | 100 | kHz | Hz |  |  |  | 57.734 | 1.5904e-06 | 57.72 | 1.5904e-06
| Frequency0.04 | Frequency | 0.04 | 10000 |  | kHz | Hz |  |  |  | 0 | 0.00012 | 0 | 0.00011561
| CDE0.1-270 | CDE | -270 | 149 | 0.1 | °C | °C |  |  |  | 0.080416 | 0 | 0.067383 | 0
| CDE0.1149.001 | CDE | 149.001 | 593 | 0.1 | °C | °C |  |  |  | 0.069131 | 0 | 0.067566 | 0
| CDE0.1593.001 | CDE | 593.001 | 1000 | 0.1 | °C | °C |  |  |  | 0.071536 | 0 | 0.070024 | 0
| CDE1-270 | CDE | -270 | 149 | 1 | °C | °C |  |  |  | 0.58502 | 0 | 0.57839 | 0
| CDE1149.001 | CDE | 149.001 | 593 | 1 | °C | °C |  |  |  | 0.57916 | 0 | 0.57842 | 0
| CDE1593.001 | CDE | 593.001 | 1000 | 1 | °C | °C |  |  |  | 0.57945 | 0 | 0.57871 | 0
| CDJ0.1-210 | CDJ | -210 | 259.999 | 0.1 | °C | °C |  |  |  | 0.0917 | 0 | 0.070601 | 0
| CDJ0.1260 | CDJ | 260 | 732 | 0.1 | °C | °C |  |  |  | 0.073715 | 0 | 0.068215 | 0
| CDJ0.1732.001 | CDJ | 732.001 | 1200 | 0.1 | °C | °C |  |  |  | 0.076447 | 0 | 0.071753 | 0
| CDJ1-210 | CDJ | -210 | 259.999 | 1 | °C | °C |  |  |  | 0.59049 | 0 | 0.57878 | 0
| CDJ1260 | CDJ | 260 | 732 | 1 | °C | °C |  |  |  | 0.58166 | 0 | 0.57849 | 0
| CDJ1732.001 | CDJ | 732.001 | 1200 | 1 | °C | °C |  |  |  | 0.58132 | 0 | 0.57892 | 0
| CDK0.1-270 | CDK | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.45663 | 0 | 0.45079 | 0
| CDK0.1-240 | CDK | -240 | 259.999 | 0.1 | °C | °C |  |  |  | 0.11392 | 0 | 0.091809 | 0
| CDK0.1260 | CDK | 260 | 816 | 0.1 | °C | °C |  |  |  | 0.08098 | 0 | 0.069993 | 0
| CDK0.1816.001 | CDK | 816.001 | 1371 | 0.1 | °C | °C |  |  |  | 0.090471 | 0 | 0.076709 | 0
| CDK1-270 | CDK | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.74458 | 0 | 0.73021 | 0
| CDK1-240 | CDK | -240 | 259.999 | 1 | °C | °C |  |  |  | 0.59718 | 0 | 0.58175 | 0
| CDK1260 | CDK | 260 | 816 | 1 | °C | °C |  |  |  | 0.58441 | 0 | 0.5787 | 0
| CDK1816.001 | CDK | 816.001 | 1371 | 1 | °C | °C |  |  |  | 0.58744 | 0 | 0.57956 | 0
| CDN0.1-270 | CDN | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.8317 | 0 | 0.82877 | 0
| CDN0.1-240 | CDN | -240 | -101.001 | 0.1 | °C | °C |  |  |  | 0.14224 | 0 | 0.12394 | 0
| CDN0.1-101 | CDN | -101 | 371 | 0.1 | °C | °C |  |  |  | 0.090241 | 0 | 0.068695 | 0
| CDN0.1371.001 | CDN | 371.001 | 816 | 0.1 | °C | °C |  |  |  | 0.081513 | 0 | 0.069464 | 0
| CDN0.1816.001 | CDN | 816.001 | 1300 | 0.1 | °C | °C |  |  |  | 0.086274 | 0 | 0.073915 | 0
| CDN1-270 | CDN | -270 | -240.001 | 1 | °C | °C |  |  |  | 1.018 | 0 | 1.0084 | 0
| CDN1-240 | CDN | -240 | -101.001 | 1 | °C | °C |  |  |  | 0.60402 | 0 | 0.58768 | 0
| CDN1-101 | CDN | -101 | 371 | 1 | °C | °C |  |  |  | 0.59027 | 0 | 0.57855 | 0
| CDN1371.001 | CDN | 371.001 | 816 | 1 | °C | °C |  |  |  | 0.58489 | 0 | 0.57864 | 0
| CDN1816.001 | CDN | 816.001 | 1300 | 1 | °C | °C |  |  |  | 0.58599 | 0 | 0.57919 | 0
| CDR0.1-50 | CDR | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.15085 | 0 | 0.13394 | 0
| CDR0.10 | CDR | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.11485 | 0 | 0.091518 | 0
| CDR0.1593.001 | CDR | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.10063 | 0 | 0.077849 | 0
| CDR0.11204.001 | CDR | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.10872 | 0 | 0.088053 | 0
| CDR1-50 | CDR | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60597 | 0 | 0.58986 | 0
| CDR10 | CDR | 0 | 593 | 1 | °C | °C |  |  |  | 0.59803 | 0 | 0.5817 | 0
| CDR1593.001 | CDR | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59357 | 0 | 0.57971 | 0
| CDR11204.001 | CDR | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.595 | 0 | 0.58117 | 0
| CDS0.1-50 | CDS | -50 | -0.001 | 0.1 | °C | °C |  |  |  | 0.13149 | 0 | 0.11169 | 0
| CDS0.10 | CDS | 0 | 593 | 0.1 | °C | °C |  |  |  | 0.10976 | 0 | 0.085663 | 0
| CDS0.1593.001 | CDS | 593.001 | 1204 | 0.1 | °C | °C |  |  |  | 0.10063 | 0 | 0.079868 | 0
| CDS0.11204.001 | CDS | 1204.001 | 1768 | 0.1 | °C | °C |  |  |  | 0.11141 | 0 | 0.091359 | 0
| CDS1-50 | CDS | -50 | -0.001 | 1 | °C | °C |  |  |  | 0.60145 | 0 | 0.58521 | 0
| CDS10 | CDS | 0 | 593 | 1 | °C | °C |  |  |  | 0.5968 | 0 | 0.58081 | 0
| CDS1593.001 | CDS | 593.001 | 1204 | 1 | °C | °C |  |  |  | 0.59384 | 0 | 0.57998 | 0
| CDS11204.001 | CDS | 1204.001 | 1768 | 1 | °C | °C |  |  |  | 0.59549 | 0 | 0.58168 | 0
| CDT0.1-270 | CDT | -270 | -240.001 | 0.1 | °C | °C |  |  |  | 0.2459 | 0 | 0.23488 | 0
| CDT0.1-240 | CDT | -240 | 92.999 | 0.1 | °C | °C |  |  |  | 0.10841 | 0 | 0.08582 | 0
| CDT0.193 | CDT | 93 | 400 | 0.1 | °C | °C |  |  |  | 0.076722 | 0 | 0.066645 | 0
| CDT1-270 | CDT | -270 | -240.001 | 1 | °C | °C |  |  |  | 0.63746 | 0 | 0.62062 | 0
| CDT1-240 | CDT | -240 | 92.999 | 1 | °C | °C |  |  |  | 0.59575 | 0 | 0.58083 | 0
| CDT193 | CDT | 93 | 400 | 1 | °C | °C |  |  |  | 0.58328 | 0 | 0.57831 | 0
| CDPT1000.01-195.6 | CDPT100 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.008485 | 0 | 0.0076317 | 0
| CDPT1000.010.001 | CDPT100 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.010087 | 0 | 0.0093806 | 0
| CDPT1000.01204.445 | CDPT100 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.012329 | 0 | 0.011526 | 0
| CDPT1000.01426.668 | CDPT100 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015887 | 0 | 0.015272 | 0
| CDPT1000.01648.89 | CDPT100 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019744 | 0 | 0.01888 | 0
| CDPT1000.1-195.6 | CDPT100 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058423 | 0 | 0.05795 | 0
| CDPT1000.10.001 | CDPT100 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058677 | 0 | 0.058206 | 0
| CDPT1000.1204.445 | CDPT100 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.059241 | 0 | 0.058591 | 0
| CDPT1000.1426.668 | CDPT100 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060082 | 0 | 0.059441 | 0
| CDPT1000.1648.89 | CDPT100 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.061563 | 0 | 0.060469 | 0
| CDPT1001-195.6 | CDPT100 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57756 | 0 | 0.57737 | 0
| CDPT10010.001 | CDPT100 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57759 | 0 | 0.5774 | 0
| CDPT1001204.445 | CDPT100 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.5777 | 0 | 0.57744 | 0
| CDPT1001426.668 | CDPT100 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57752 | 0
| CDPT1001648.89 | CDPT100 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57809 | 0 | 0.57763 | 0
| CDPT10000.01-195.6 | CDPT1000 | -195.6 | 0 | 0.01 | °C | °C |  |  |  | 0.0069515 | 0 | 0.0065655 | 0
| CDPT10000.010.001 | CDPT1000 | 0.001 | 204.444 | 0.01 | °C | °C |  |  |  | 0.0097193 | 0 | 0.0093614 | 0
| CDPT10000.01204.445 | CDPT1000 | 204.445 | 426.667 | 0.01 | °C | °C |  |  |  | 0.011768 | 0 | 0.011474 | 0
| CDPT10000.01426.668 | CDPT1000 | 426.668 | 648.889 | 0.01 | °C | °C |  |  |  | 0.015792 | 0 | 0.015173 | 0
| CDPT10000.01648.89 | CDPT1000 | 648.89 | 850 | 0.01 | °C | °C |  |  |  | 0.019042 | 0 | 0.018733 | 0
| CDPT10000.1-195.6 | CDPT1000 | -195.6 | 0 | 0.1 | °C | °C |  |  |  | 0.058 | 0 | 0.05782 | 0
| CDPT10000.10.001 | CDPT1000 | 0.001 | 204.444 | 0.1 | °C | °C |  |  |  | 0.058438 | 0 | 0.058203 | 0
| CDPT10000.1204.445 | CDPT1000 | 204.445 | 426.667 | 0.1 | °C | °C |  |  |  | 0.058813 | 0 | 0.05858 | 0
| CDPT10000.1426.668 | CDPT1000 | 426.668 | 648.889 | 0.1 | °C | °C |  |  |  | 0.060057 | 0 | 0.059416 | 0
| CDPT10000.1648.89 | CDPT1000 | 648.89 | 850 | 0.1 | °C | °C |  |  |  | 0.060808 | 0 | 0.060423 | 0
| CDPT10001-195.6 | CDPT1000 | -195.6 | 0 | 1 | °C | °C |  |  |  | 0.57743 | 0 | 0.57736 | 0
| CDPT100010.001 | CDPT1000 | 0.001 | 204.444 | 1 | °C | °C |  |  |  | 0.57749 | 0 | 0.5774 | 0
| CDPT10001204.445 | CDPT1000 | 204.445 | 426.667 | 1 | °C | °C |  |  |  | 0.57753 | 0 | 0.57744 | 0
| CDPT10001426.668 | CDPT1000 | 426.668 | 648.889 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57752 | 0
| CDPT10001648.89 | CDPT1000 | 648.89 | 850 | 1 | °C | °C |  |  |  | 0.57779 | 0 | 0.57763 | 0
| FDE0.1-454 | FDE | -454 | 300 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.08462 | 0
| FDE0.1300.001 | FDE | 300.001 | 1100 | 0.1 | °F | °F |  |  |  | 0.09 | 0 | 0.084258 | 0
| FDE0.11100.001 | FDE | 1100.001 | 1832 | 0.1 | °F | °F |  |  |  | 0.091751 | 0 | 0.08789 | 0
| FDE1-454 | FDE | -454 | 300 | 1 | °F | °F |  |  |  | 0.60177 | 0 | 0.58066 | 0
| FDE1300.001 | FDE | 300.001 | 1100 | 1 | °F | °F |  |  |  | 0.58299 | 0 | 0.5806 | 0
| FDE11100.001 | FDE | 1100.001 | 1832 | 1 | °F | °F |  |  |  | 0.58352 | 0 | 0.58114 | 0
| FDJ0.1-346 | FDJ | -346 | 499.999 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.091814 | 0
| FDJ0.1500 | FDJ | 500 | 1350 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.085419 | 0
| FDJ0.11350.001 | FDJ | 1350.001 | 2192 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.090692 | 0
| FDJ1-346 | FDJ | -346 | 499.999 | 1 | °F | °F |  |  |  | 0.61872 | 0 | 0.58175 | 0
| FDJ1500 | FDJ | 500 | 1350 | 1 | °F | °F |  |  |  | 0.59113 | 0 | 0.58077 | 0
| FDJ11350.001 | FDJ | 1350.001 | 2192 | 1 | °F | °F |  |  |  | 0.58927 | 0 | 0.58157 | 0
| FDK0.1-454 | FDK | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.77443 | 0 | 0.76327 | 0
| FDK0.1-400 | FDK | -400 | 499.999 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13588 | 0
| FDK0.1500 | FDK | 500 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.088716 | 0
| FDK0.11500.001 | FDK | 1500.001 | 2500 | 0.1 | °F | °F |  |  |  | 0.13135 | 0 | 0.098982 | 0
| FDK1-454 | FDK | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.99057 | 0 | 0.95529 | 0
| FDK1-400 | FDK | -400 | 499.999 | 1 | °F | °F |  |  |  | 0.63828 | 0 | 0.59031 | 0
| FDK1500 | FDK | 500 | 1500 | 1 | °F | °F |  |  |  | 0.59947 | 0 | 0.58127 | 0
| FDK11500.001 | FDK | 1500.001 | 2500 | 1 | °F | °F |  |  |  | 0.60796 | 0 | 0.58292 | 0
| FDN0.1-454 | FDN | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 1.5 | 0 | 1.4288 | 0
| FDN0.1-400 | FDN | -400 | -150.001 | 0.1 | °F | °F |  |  |  | 0.24 | 0 | 0.19873 | 0
| FDN0.1-150 | FDN | -150 | 700 | 0.1 | °F | °F |  |  |  | 0.14 | 0 | 0.088083 | 0
| FDN0.1700.001 | FDN | 700.001 | 1500 | 0.1 | °F | °F |  |  |  | 0.12 | 0 | 0.087954 | 0
| FDN0.11500.001 | FDN | 1500.001 | 2372 | 0.1 | °F | °F |  |  |  | 0.13 | 0 | 0.094806 | 0
| FDN1-454 | FDN | -454 | -400.001 | 1 | °F | °F |  |  |  | 1.5603 | 0 | 1.54 | 0
| FDN1-400 | FDN | -400 | -150.001 | 1 | °F | °F |  |  |  | 0.65772 | 0 | 0.60786 | 0
| FDN1-150 | FDN | -150 | 700 | 1 | °F | °F |  |  |  | 0.61817 | 0 | 0.58117 | 0
| FDN1700.001 | FDN | 700.001 | 1500 | 1 | °F | °F |  |  |  | 0.60109 | 0 | 0.58115 | 0
| FDN11500.001 | FDN | 1500.001 | 2372 | 1 | °F | °F |  |  |  | 0.60386 | 0 | 0.58223 | 0
| FDR0.1-58 | FDR | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.26 | 0 | 0.22215 | 0
| FDR0.132 | FDR | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.19 | 0 | 0.13898 | 0
| FDR0.11100.001 | FDR | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.16 | 0 | 0.10571 | 0
| FDR0.12200.001 | FDR | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.12242 | 0
| FDR1-58 | FDR | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.66465 | 0 | 0.61591 | 0
| FDR132 | FDR | 32 | 1100 | 1 | °F | °F |  |  |  | 0.64166 | 0 | 0.59103 | 0
| FDR11100.001 | FDR | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.6276 | 0 | 0.5841 | 0
| FDR12200.001 | FDR | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.63063 | 0 | 0.58736 | 0
| FDS0.1-58 | FDS | -58 | 31.999 | 0.1 | °F | °F |  |  |  | 0.22 | 0 | 0.17936 | 0
| FDS0.132 | FDS | 32 | 1100 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12667 | 0
| FDS0.11100.001 | FDS | 1100.001 | 2200 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.10571 | 0
| FDS0.12200.001 | FDS | 2200.001 | 3214 | 0.1 | °F | °F |  |  |  | 0.17 | 0 | 0.12242 | 0
| FDS1-58 | FDS | -58 | 31.999 | 1 | °F | °F |  |  |  | 0.6516 | 0 | 0.6018 | 0
| FDS132 | FDS | 32 | 1100 | 1 | °F | °F |  |  |  | 0.63802 | 0 | 0.58826 | 0
| FDS11100.001 | FDS | 1100.001 | 2200 | 1 | °F | °F |  |  |  | 0.62833 | 0 | 0.5849 | 0
| FDS12200.001 | FDS | 2200.001 | 3214 | 1 | °F | °F |  |  |  | 0.63063 | 0 | 0.58878 | 0
| FDT0.1-454 | FDT | -454 | -400.001 | 0.1 | °F | °F |  |  |  | 0.41398 | 0 | 0.39271 | 0
| FDT0.1-400 | FDT | -400 | 199.999 | 0.1 | °F | °F |  |  |  | 0.18 | 0 | 0.12412 | 0
| FDT0.1200 | FDT | 200 | 752 | 0.1 | °F | °F |  |  |  | 0.11 | 0 | 0.083094 | 0
| FDT1-454 | FDT | -454 | -400.001 | 1 | °F | °F |  |  |  | 0.74355 | 0 | 0.69586 | 0
| FDT1-400 | FDT | -400 | 199.999 | 1 | °F | °F |  |  |  | 0.63425 | 0 | 0.58771 | 0
| FDT1200 | FDT | 200 | 752 | 1 | °F | °F |  |  |  | 0.59635 | 0 | 0.58043 | 0
| FDPT1000.01-320 | FDPT100 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.013 | 0 | 0.010679 | 0
| FDPT1000.0132.001 | FDPT100 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014506 | 0
| FDPT1000.01400.001 | FDPT100 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.021 | 0 | 0.018862 | 0
| FDPT1000.01800.001 | FDPT100 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.027261 | 0 | 0.026097 | 0
| FDPT1000.011200.001 | FDPT100 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.034473 | 0 | 0.032867 | 0
| FDPT1000.1-320 | FDPT100 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.059936 | 0 | 0.05843 | 0
| FDPT1000.132.001 | FDPT100 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.060734 | 0 | 0.059249 | 0
| FDPT1000.1400.001 | FDPT100 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.062482 | 0 | 0.060463 | 0
| FDPT1000.1800.001 | FDPT100 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.065034 | 0 | 0.063096 | 0
| FDPT1000.11200.001 | FDPT100 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.069375 | 0 | 0.066183 | 0
| FDPT1001-320 | FDPT100 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57804 | 0 | 0.57742 | 0
| FDPT100132.001 | FDPT100 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57812 | 0 | 0.5775 | 0
| FDPT1001400.001 | FDPT100 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57849 | 0 | 0.57763 | 0
| FDPT1001800.001 | FDPT100 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57877 | 0 | 0.57791 | 0
| FDPT10011200.001 | FDPT100 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57975 | 0 | 0.57826 | 0
| FDPT10000.01-320 | FDPT1000 | -320 | 32 | 0.01 | °F | °F |  |  |  | 0.0090498 | 0 | 0.0080621 | 0
| FDPT10000.0132.001 | FDPT1000 | 32.001 | 400 | 0.01 | °F | °F |  |  |  | 0.016 | 0 | 0.014466 | 0
| FDPT10000.01400.001 | FDPT1000 | 400.001 | 800 | 0.01 | °F | °F |  |  |  | 0.02 | 0 | 0.018759 | 0
| FDPT10000.01800.001 | FDPT1000 | 800.001 | 1200 | 0.01 | °F | °F |  |  |  | 0.02708 | 0 | 0.025909 | 0
| FDPT10000.011200.001 | FDPT1000 | 1200.001 | 1562 | 0.01 | °F | °F |  |  |  | 0.033169 | 0 | 0.032594 | 0
| FDPT10000.1-320 | FDPT1000 | -320 | 32 | 0.1 | °F | °F |  |  |  | 0.058588 | 0 | 0.058009 | 0
| FDPT10000.132.001 | FDPT1000 | 32.001 | 400 | 0.1 | °F | °F |  |  |  | 0.059981 | 0 | 0.059239 | 0
| FDPT10000.1400.001 | FDPT1000 | 400.001 | 800 | 0.1 | °F | °F |  |  |  | 0.061159 | 0 | 0.060431 | 0
| FDPT10000.1800.001 | FDPT1000 | 800.001 | 1200 | 0.1 | °F | °F |  |  |  | 0.064958 | 0 | 0.063018 | 0
| FDPT10000.11200.001 | FDPT1000 | 1200.001 | 1562 | 0.1 | °F | °F |  |  |  | 0.067183 | 0 | 0.066048 | 0
| FDPT10001-320 | FDPT1000 | -320 | 32 | 1 | °F | °F |  |  |  | 0.57761 | 0 | 0.57738 | 0
| FDPT1000132.001 | FDPT1000 | 32.001 | 400 | 1 | °F | °F |  |  |  | 0.57781 | 0 | 0.5775 | 0
| FDPT10001400.001 | FDPT1000 | 400.001 | 800 | 1 | °F | °F |  |  |  | 0.57793 | 0 | 0.57763 | 0
| FDPT10001800.001 | FDPT1000 | 800.001 | 1200 | 1 | °F | °F |  |  |  | 0.57876 | 0 | 0.5779 | 0
| FDPT100011200.001 | FDPT1000 | 1200.001 | 1562 | 1 | °F | °F |  |  |  | 0.57876 | 0 | 0.57824 | 0

## Standard3
Updated: | 2021-02-23 | 34420A
| Main Name | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCVmeas0.000000010 | DCVmeas | 0 | 0.001 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.027845 | 71.95
| DCVmeas0.00000010 | DCVmeas | 0 | 0.001 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.062207 | 53.073
| DCVmeas0.0000010 | DCVmeas | 0 | 0.001 | 1e-06 | V | µV |  |  |  | 0.62 | 4.5 | 0.5802 | 8.2952
| DCVmeas0.000010 | DCVmeas | 0 | 0.001 | 1e-05 | V | µV |  |  |  | 5.8 | 0.89276 | 5.8 | 0.83626
| DCVmeas0.00010 | DCVmeas | 0 | 0.001 | 0.0001 | V | µV |  |  |  | 58 | 0.10624 | 58 | 0.083633
| DCVmeas0.0010 | DCVmeas | 0 | 0.001 | 0.001 | V | µV |  |  |  | 580 | 0.017406 | 580 | 0.0083633
| DCVmeas0.000000010.001 | DCVmeas | 0.001 | 0.01 | 1e-08 | V | µV |  |  |  | 0.62 | 4.5 | 0.060672 | 48.506
| DCVmeas0.00000010.001 | DCVmeas | 0.001 | 0.01 | 1e-07 | V | µV |  |  |  | 0.62 | 4.5 | 0.076236 | 47.254
| DCVmeas0.0000010.001 | DCVmeas | 0.001 | 0.01 | 1e-06 | V | µV |  |  |  | 0.62 | 19.155 | 0.56725 | 22.912
| DCVmeas0.000010.001 | DCVmeas | 0.001 | 0.01 | 1e-05 | V | µV |  |  |  | 5.7983 | 2.757 | 5.7983 | 2.7323
| DCVmeas0.00010.001 | DCVmeas | 0.001 | 0.01 | 0.0001 | V | µV |  |  |  | 58 | 0.2838 | 58 | 0.27385
| DCVmeas0.0010.001 | DCVmeas | 0.001 | 0.01 | 0.001 | V | µV |  |  |  | 580 | 0.031364 | 580 | 0.027386
| DCVmeas0.00000010.01 | DCVmeas | 0.01 | 0.1 | 1e-07 | V | µV |  |  |  | 0.48 | 37 | 0.47363 | 36.624
| DCVmeas0.0000010.01 | DCVmeas | 0.01 | 0.1 | 1e-06 | V | µV |  |  |  | 0.6696 | 35.075 | 0.66823 | 35.079
| DCVmeas0.000010.01 | DCVmeas | 0.01 | 0.1 | 1e-05 | V | µV |  |  |  | 5.7207 | 14.051 | 5.7199 | 14.036
| DCVmeas0.00010.01 | DCVmeas | 0.01 | 0.1 | 0.0001 | V | µV |  |  |  | 57.991 | 1.5774 | 57.99 | 1.5689
| DCVmeas0.0010.01 | DCVmeas | 0.01 | 0.1 | 0.001 | V | µV |  |  |  | 580 | 0.16049 | 580 | 0.1571
| DCVmeas0.010.01 | DCVmeas | 0.01 | 0.1 | 0.01 | V | µV |  |  |  | 5800 | 0.017067 | 5800 | 0.01571
| DCVmeas0.0000010.1 | DCVmeas | 0.1 | 1 | 1e-06 | V | µV |  |  |  | 4.7 | 41 | 4.6518 | 40.782
| DCVmeas0.000010.1 | DCVmeas | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 6.5523 | 39.255 | 6.5389 | 39.26
| DCVmeas0.00010.1 | DCVmeas | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 56.989 | 16.705 | 56.981 | 16.693
| DCVmeas0.0010.1 | DCVmeas | 0.1 | 1 | 0.001 | V | µV |  |  |  | 579.88 | 1.9084 | 579.88 | 1.9012
| DCVmeas0.010.1 | DCVmeas | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5800 | 0.19329 | 5800 | 0.19042
| DCVmeas0.10.1 | DCVmeas | 0.1 | 1 | 0.1 | V | µV |  |  |  | 58000 | 0.020188 | 58000 | 0.019042
| DCVmeas0.000011 | DCVmeas | 1 | 10 | 1e-05 | V | µV |  |  |  | 47 | 35 | 46.376 | 34.813
| DCVmeas0.00011 | DCVmeas | 1 | 10 | 0.0001 | V | µV |  |  |  | 66.458 | 33.232 | 66.382 | 33.232
| DCVmeas0.0011 | DCVmeas | 1 | 10 | 0.001 | V | µV |  |  |  | 572.81 | 12.879 | 572.76 | 12.867
| DCVmeas0.011 | DCVmeas | 1 | 10 | 0.01 | V | µV |  |  |  | 5799.2 | 1.4325 | 5799.1 | 1.4259
| DCVmeas0.11 | DCVmeas | 1 | 10 | 0.1 | V | µV |  |  |  | 58000 | 0.1454 | 58000 | 0.14276
| DCVmeas11 | DCVmeas | 1 | 10 | 1 | V | µV |  |  |  | 580000 | 0.015331 | 580000 | 0.014276
| DCVmeas0.000110 | DCVmeas | 10 | 100 | 0.0001 | V | mV |  |  |  | 0.58 | 0.041 | 0.57895 | 0.040741
| DCVmeas0.00110 | DCVmeas | 10 | 100 | 0.001 | V | mV |  |  |  | 0.7495 | 0.039399 | 0.74879 | 0.039399
| DCVmeas0.0110 | DCVmeas | 10 | 100 | 0.01 | V | mV |  |  |  | 5.7109 | 0.017264 | 5.7105 | 0.017251
| DCVmeas0.110 | DCVmeas | 10 | 100 | 0.1 | V | mV |  |  |  | 57.989 | 0.0019846 | 57.989 | 0.0019773
| DCVmeas110 | DCVmeas | 10 | 100 | 1 | V | mV |  |  |  | 580 | 0.000201 | 580 | 0.00019806
| Resistance0.0000000010 | Resistance | 0 | 1 | 1e-09 | kOhms | µΩ |  |  |  | 3 | 120 | 2.3273 | 80.8131
| Resistance0.000000010 | Resistance | 0 | 1 | 1e-08 | kOhms | µΩ |  |  |  | 3 | 120 | 3.9637 | 79.3768
| Resistance0.00000010 | Resistance | 0 | 1 | 1e-07 | kOhms | µΩ |  |  |  | 63.6928 | 62.6594 | 54.2074 | 47.1632
| Resistance0.0000010 | Resistance | 0 | 1 | 1e-06 | kOhms | µΩ |  |  |  | 600.322 | 96.4808 | 579.4447 | 6.4836
| Resistance0.000010 | Resistance | 0 | 1 | 1e-05 | kOhms | mΩ |  |  |  | 5.8513 | 0.24805 | 5.7999 | 0.00065169 | 0.0058513 | 0.00024805 | 0.0057999 | 6.5169e-07
| Resistance0.000000011 | Resistance | 1 | 10 | 1e-08 | kOhms | µΩ |  |  |  | 24.0351 | 69.2135 | 23.2955 | 69.2642
| Resistance0.00000011 | Resistance | 1 | 10 | 1e-07 | kOhms | µΩ |  |  |  | 66.5586 | 66.0557 | 41.3894 | 67.687
| Resistance0.0000011 | Resistance | 1 | 10 | 1e-06 | kOhms | µΩ |  |  |  | 643.2147 | 54.7509 | 550.1922 | 37.1183
| Resistance0.000011 | Resistance | 1 | 10 | 1e-05 | kOhms | mΩ |  |  |  | 6.0044 | 0.095076 | 5.7959 | 0.004809 | 0.0060044 | 9.5076e-05 | 0.0057959 | 4.809e-06
| Resistance0.00011 | Resistance | 1 | 10 | 0.0001 | kOhms | mΩ |  |  |  | 58.5135 | 0.24789 | 57.9996 | 0.00048274 | 0.058513 | 0.00024789 | 0.058 | 4.8274e-07
| Resistance0.000000110 | Resistance | 10 | 100 | 1e-07 | kOhms | mΩ |  |  |  | 0.24035 | 0.069214 | 0.23295 | 0.069264
| Resistance0.00000110 | Resistance | 10 | 100 | 1e-06 | kOhms | mΩ |  |  |  | 0.66559 | 0.066056 | 0.41389 | 0.067687
| Resistance0.0000110 | Resistance | 10 | 100 | 1e-05 | kOhms | mΩ |  |  |  | 6.4321 | 0.054751 | 5.5019 | 0.037118
| Resistance0.000110 | Resistance | 10 | 100 | 0.0001 | kOhms | mΩ |  |  |  | 60.0444 | 0.095076 | 57.9593 | 0.004809
| Resistance0.00110 | Resistance | 10 | 100 | 0.001 | kOhms | mΩ |  |  |  | 585.1347 | 0.24789 | 579.9959 | 0.000483
| Resistance0.001100 | Resistance | 100 | 1 | 0.001 | kOhm | mΩ |  |  |  | 2.4035 | 0.069214 | 2.3295 | 0.069264
| Resistance0.01100 | Resistance | 100 | 1 | 0.01 | kOhm | mΩ |  |  |  | 6.6559 | 0.066056 | 4.1389 | 0.067687
| Resistance0.1100 | Resistance | 100 | 1 | 0.1 | kOhm | mΩ |  |  |  | 64.3215 | 0.054751 | 55.0192 | 0.037118
| Resistance1100 | Resistance | 100 | 1 | 1 | kOhm | mΩ |  |  |  | 600.444 | 0.095076 | 579.5927 | 0.004809
| Resistance10100 | Resistance | 100 | 1 | 10 | kOhm | Ω |  |  |  | 5.8513 | 0.00024789 | 5.8 | 4.83e-07 | 0.0058513 | 2.4789e-07 | 0.0058 | 4.83e-10
| Resistance0.011 | Resistance | 1 | 10 | 0.01 | kOhm | mΩ |  |  |  | 24.0351 | 0.069214 | 23.2955 | 0.069264
| Resistance0.11 | Resistance | 1 | 10 | 0.1 | kOhm | mΩ |  |  |  | 66.5586 | 0.066056 | 41.3894 | 0.067688
| Resistance11 | Resistance | 1 | 10 | 1 | kOhm | mΩ |  |  |  | 643.2147 | 0.054751 | 550.192 | 0.037118
| Resistance101 | Resistance | 1 | 10 | 10 | kOhm | Ω |  |  |  | 6.0044 | 9.5076e-05 | 5.7959 | 4.809e-06 | 0.0060044 | 9.5076e-08 | 0.0057959 | 4.809e-09
| Resistance1001 | Resistance | 1 | 10 | 100 | kOhm | Ω |  |  |  | 58.5135 | 0.00024789 | 57.9996 | 4.83e-07 | 0.058513 | 2.4789e-07 | 0.058 | 4.83e-10
| Resistance0.110 | Resistance | 10 | 100 | 0.1 | kOhm | Ω |  |  |  | 0.46937 | 6.92e-05 | 0.46349 | 6.93e-05
| Resistance110 | Resistance | 10 | 100 | 1 | kOhm | Ω |  |  |  | 0.82928 | 6.67e-05 | 0.61213 | 6.81e-05
| Resistance1010 | Resistance | 10 | 100 | 10 | kOhm | Ω |  |  |  | 6.4547 | 5.59e-05 | 5.5271 | 3.87e-05
| Resistance10010 | Resistance | 10 | 100 | 100 | kOhm | Ω |  |  |  | 60.0461 | 9.53e-05 | 57.9607 | 5.09e-06
| Resistance100010 | Resistance | 10 | 100 | 1000 | kOhm | Ω |  |  |  | 585.1349 | 0.000248 | 579.996 | 5.1e-07
| Resistance1000100 | Resistance | 100 | 1 | 1000 | kOhm | Ω |  |  |  | 5 | 8.05e-05 | 4.6334 | 8.08e-05
| Resistance10000100 | Resistance | 100 | 1 | 10000 | kOhm | Ω |  |  |  | 8.0349 | 7.84e-05 | 5.9988 | 7.96e-05
| Resistance100000100 | Resistance | 100 | 1 | 100000 | kOhm | Ω |  |  |  | 63.945 | 6.39e-05 | 54.4969 | 4.88e-05
| Resistance1000000100 | Resistance | 100 | 1 | 1000000 | kOhm | Ω |  |  |  | 600.3396 | 9.67e-05 | 579.4589 | 6.8e-06
| Resistance10000000100 | Resistance | 100 | 1 | 10000000 | kOhm | kΩ |  |  |  | 5.8513 | 2.481e-07 | 5.7999 | 6.9e-10 | 0.0058513 | 2.481e-10 | 0.0057999 | 6.9e-13

## Standard4
Updated: | 2021-02-23 | Transmille
| Main Name | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.0000010 | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 12 | 92 | 11.624 | 92.068
| DCV0.000010 | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 12.963 | 86.429 | 12.956 | 86.458
| DCV0.00010 | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.889 | 24.899 | 58.883 | 24.902
| DCV0.0010 | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 577.47 | 2.5928 | 577.47 | 2.5928
| DCV0.010 | DCV | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5773.6 | 0.25939 | 5773.6 | 0.25939
| DCV0.000010.1 | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 36 | 92 | 35.885 | 91.648
| DCV0.00010.1 | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 65.685 | 74.286 | 65.574 | 74.301
| DCV0.0010.1 | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 577.78 | 13.555 | 577.73 | 13.515
| DCV0.010.1 | DCV | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5773.6 | 1.3866 | 5773.6 | 1.3696
| DCV0.10.1 | DCV | 0.1 | 1 | 0.1 | V | µV |  |  |  | 57736 | 0.14379 | 57736 | 0.13698
| DCV0.00011 | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 360 | 92 | 354.54 | 91.829
| DCV0.0011 | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 653.72 | 74.371 | 653.17 | 74.329
| DCV0.011 | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5777.2 | 13.559 | 5777 | 13.491
| DCV0.11 | DCV | 1 | 10 | 0.1 | V | µV |  |  |  | 57736 | 1.3954 | 57736 | 13.496
| DCV11 | DCV | 1 | 10 | 1 | V | µV |  |  |  | 577360 | 0.14808 | 577360 | 13.497
| DCV0.00110 | DCV | 10 | 100 | 0.001 | V | mV |  |  |  | 6.8 | 0.082 | 6.7338 | 0.082465
| DCV0.0110 | DCV | 10 | 100 | 0.01 | V | mV |  |  |  | 8.774 | 0.072816 | 8.766 | 0.072781
| DCV0.110 | DCV | 10 | 100 | 0.1 | V | mV |  |  |  | 58.072 | 0.015845 | 58.068 | 0.015769
| DCV110 | DCV | 10 | 100 | 1 | V | mV |  |  |  | 577.39 | 0.0016415 | 577.39 | 0.0016093
| DCV1010 | DCV | 10 | 100 | 10 | V | mV |  |  |  | 5773.6 | 0.00017383 | 5773.6 | 0.00016097
| DCV0.01100 | DCV | 100 | 1020 | 0.01 | V | mV |  |  |  | 36 | 0.092 | 35.843 | 0.091772
| DCV0.1100 | DCV | 100 | 1020 | 0.1 | V | mV |  |  |  | 65.632 | 0.07463 | 65.525 | 0.074608
| DCV1100 | DCV | 100 | 1020 | 1 | V | mV |  |  |  | 577.76 | 0.013752 | 577.71 | 0.013677
| DCV10100 | DCV | 100 | 1020 | 10 | V | mV |  |  |  | 5773.6 | 0.0014179 | 5773.6 | 0.0013864
| DCV100100 | DCV | 100 | 1020 | 100 | V | mV |  |  |  | 57736 | 0.00015125 | 57736 | 0.00013866
| DCA0.0000000010 | DCA | 0 | 0.0001 | 1e-09 | A | nA |  |  |  | 35 | 350000 | 34.933 | 349170
| DCA0.000000010 | DCA | 0 | 0.0001 | 1e-08 | A | nA |  |  |  | 35.368 | 346810 | 35.366 | 346810
| DCA0.00000010 | DCA | 0 | 0.0001 | 1e-07 | A | nA |  |  |  | 67.34 | 230750 | 67.336 | 230720
| DCA0.0000010 | DCA | 0 | 0.0001 | 1e-06 | A | nA |  |  |  | 578.4 | 31405 | 578.39 | 31376
| DCA0.000010 | DCA | 0 | 0.0001 | 1e-05 | A | nA |  |  |  | 5773.7 | 3163.2 | 5773.7 | 3151.6
| DCA0.000000010.0001 | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 120 | 350000 | 115.86 | 347000
| DCA0.00000010.0001 | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 127.29 | 339170 | 127.24 | 339180
| DCA0.0000010.0001 | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 580.76 | 159300 | 580.72 | 159250
| DCA0.000010.0001 | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.7 | 18430 | 5773.7 | 18401
| DCA0.00010.0001 | DCA | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 57736 | 1854.7 | 57736 | 1843.4
| DCA0.00000010.001 | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 1.2 | 350 | 1.1597 | 346.89
| DCA0.0000010.001 | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 1.2739 | 339.07 | 1.2734 | 339.08
| DCA0.000010.001 | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.8079 | 159.28 | 5.8074 | 159.23
| DCA0.00010.001 | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.737 | 18.431 | 57.737 | 18.399
| DCA0.0010.001 | DCA | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 1.856 | 577.36 | 1.8431
| DCA0.0000010.01 | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 12 | 350 | 11.589 | 348.45
| DCA0.000010.01 | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 12.73 | 340.62 | 12.725 | 340.63
| DCA0.00010.01 | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 58.071 | 160.28 | 58.066 | 160.23
| DCA0.0010.01 | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.37 | 18.56 | 577.37 | 18.528
| DCA0.010.01 | DCA | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 1.8689 | 5773.6 | 1.856
| DCA0.000010.1 | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 180 | 370 | 172.98 | 365.23
| DCA0.00010.1 | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 181.28 | 360.03 | 181.24 | 360.03
| DCA0.0010.1 | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 594.75 | 194.63 | 594.71 | 194.58
| DCA0.010.1 | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5775 | 23.623 | 5775 | 23.591
| DCA0.10.1 | DCA | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 2.3778 | 57736 | 2.3649
| DCA0.00011 | DCA | 1 | 10.2 | 0.0001 | A | µA |  |  |  | 2.4 | 0.59 | 2.3095 | 0.57975
| DCA0.0011 | DCA | 1 | 10.2 | 0.001 | A | µA |  |  |  | 2.4 | 0.63081 | 2.3701 | 0.57578
| DCA0.011 | DCA | 1 | 10.2 | 0.01 | A | µA |  |  |  | 5.8935 | 0.56293 | 6.0655 | 0.39037
| DCA0.11 | DCA | 1 | 10.2 | 0.1 | A | µA |  |  |  | 57.625 | 0.18253 | 57.752 | 0.055477
| DCA11 | DCA | 1 | 10.2 | 1 | A | µA |  |  |  | 577.31 | 0.057402 | 577.36 | 0.0055789
| DCA 10 turn0.000110 | DCA 10 turn | 10 | 102 | 0.0001 | A | mA |  |  |  | 16 | 2 | 15.052 | 2.0088
| DCA 10 turn0.00110 | DCA 10 turn | 10 | 102 | 0.001 | A | mA |  |  |  | 16 | 2.0045 | 15.566 | 2.0045
| DCA 10 turn0.0110 | DCA 10 turn | 10 | 102 | 0.01 | A | mA |  |  |  | 50.243 | 1.7371 | 50.216 | 1.7371
| DCA 10 turn0.110 | DCA 10 turn | 10 | 102 | 0.1 | A | mA |  |  |  | 574.15 | 0.42863 | 574.14 | 0.42836
| DCA 10 turn110 | DCA 10 turn | 10 | 102 | 1 | A | mA |  |  |  | 5773.2 | 0.044479 | 5773.2 | 0.044361
| DCA 10 turn1010 | DCA 10 turn | 10 | 102 | 10 | A | mA |  |  |  | 57735 | 0.0044851 | 57735 | 0.0044377
| DCA 20 turn0.00120 | DCA 20 turn | 20 | 204 | 0.001 | A | mA |  |  |  | 32 | 2 | 30.087 | 2.0089
| DCA 20 turn0.0120 | DCA 20 turn | 20 | 204 | 0.01 | A | mA |  |  |  | 32 | 2.0077 | 30.344 | 2.0078
| DCA 20 turn0.120 | DCA 20 turn | 20 | 204 | 0.1 | A | mA |  |  |  | 52.695 | 1.9168 | 52.601 | 1.917
| DCA 20 turn120 | DCA 20 turn | 20 | 204 | 1 | A | mA |  |  |  | 565.98 | 0.78427 | 565.94 | 0.78385
| DCA 20 turn1020 | DCA 20 turn | 20 | 204 | 10 | A | mA |  |  |  | 5772.2 | 0.088858 | 5772.2 | 0.088623
| DCA 50 turn0.0015 | DCA 50 turn | 5 | 51 | 0.001 | A | mA |  |  |  | 4.7 | 2 | 4.5993 | 1.98
| DCA 50 turn0.015 | DCA 50 turn | 5 | 51 | 0.01 | A | mA |  |  |  | 5.9279 | 1.9581 | 5.8001 | 1.9595
| DCA 50 turn0.15 | DCA 50 turn | 5 | 51 | 0.1 | A | mA |  |  |  | 53.039 | 1.3233 | 52.916 | 1.3219
| DCA 50 turn15 | DCA 50 turn | 5 | 51 | 1 | A | mA |  |  |  | 576.56 | 0.20649 | 576.52 | 0.20417
| DCA 50 turn105 | DCA 50 turn | 5 | 51 | 10 | A | mA |  |  |  | 5773.5 | 0.021536 | 5773.5 | 0.020588
| DCA 50 turn0.00150 | DCA 50 turn | 50 | 510 | 0.001 | A | mA |  |  |  | 80 | 2 | 75.208 | 2.0089
| DCA 50 turn0.0150 | DCA 50 turn | 50 | 510 | 0.01 | A | mA |  |  |  | 80 | 2.0086 | 75.318 | 2.0087
| DCA 50 turn0.150 | DCA 50 turn | 50 | 510 | 0.1 | A | mA |  |  |  | 85.569 | 1.9918 | 85.292 | 1.9921
| DCA 50 turn150 | DCA 50 turn | 50 | 510 | 1 | A | mA |  |  |  | 534.37 | 1.3886 | 534.07 | 1.3883
| DCA 50 turn1050 | DCA 50 turn | 50 | 510 | 10 | A | mA |  |  |  | 5765.3 | 0.22044 | 5765.2 | 0.21986
| ACV0.00000100.01 | ACV | 0 | 0.1 | 1e-06 | V | µV | 0.01 | 2 | KHz | 35 | 940 | 34.74 | 937.4
| ACV0.0000100.01 | ACV | 0 | 0.1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 35.162 | 934.08 | 35.157 | 934.05
| ACV0.000100.01 | ACV | 0 | 0.1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 67.116 | 737.26 | 67.105 | 737.07
| ACV0.00100.01 | ACV | 0 | 0.1 | 0.001 | V | µV | 0.01 | 2 | KHz | 578.33 | 131.69 | 578.32 | 131.46
| ACV0.0100.01 | ACV | 0 | 0.1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5773.7 | 13.415 | 5773.6 | 13.317
| ACV0.00000102 | ACV | 0 | 0.1 | 1e-06 | V | µV | 2 | 20 | KHz | 36 | 930 | 35.471 | 926.49
| ACV0.0000102 | ACV | 0 | 0.1 | 1e-05 | V | µV | 2 | 20 | KHz | 36 | 923.18 | 35.865 | 923.15
| ACV0.000102 | ACV | 0 | 0.1 | 0.0001 | V | µV | 2 | 20 | KHz | 67.126 | 727.32 | 67.116 | 727.12
| ACV0.00102 | ACV | 0 | 0.1 | 0.001 | V | µV | 2 | 20 | KHz | 578.33 | 129.33 | 578.32 | 129.1
| ACV0.0102 | ACV | 0 | 0.1 | 0.01 | V | µV | 2 | 20 | KHz | 5773.7 | 13.173 | 5773.6 | 13.075
| ACV0.000010.10.01 | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 350 | 940 | 347.15 | 937.18
| ACV0.00010.10.01 | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 351.19 | 934.51 | 351.16 | 934.46
| ACV0.0010.10.01 | ACV | 0.1 | 1 | 0.001 | V | µV | 0.01 | 2 | KHz | 650.74 | 757.68 | 650.66 | 757.46
| ACV0.010.10.01 | ACV | 0.1 | 1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5776.6 | 138.39 | 5776.5 | 138.13
| ACV0.10.10.01 | ACV | 0.1 | 1 | 0.1 | V | µV | 0.01 | 2 | KHz | 57736 | 14.108 | 57736 | 14
| ACV0.000010.12 | ACV | 0.1 | 1 | 1e-05 | V | µV | 2 | 20 | KHz | 850 | 1700 | 808.55 | 1738.9
| ACV0.00010.12 | ACV | 0.1 | 1 | 0.0001 | V | µV | 2 | 20 | KHz | 850 | 1737.8 | 810.33 | 1737.8
| ACV0.0010.12 | ACV | 0.1 | 1 | 0.001 | V | µV | 2 | 20 | KHz | 976.25 | 1636.3 | 975.89 | 1636.2
| ACV0.010.12 | ACV | 0.1 | 1 | 0.01 | V | µV | 2 | 20 | KHz | 5806.3 | 505.05 | 5806.1 | 504.49
| ACV0.10.12 | ACV | 0.1 | 1 | 0.1 | V | µV | 2 | 20 | KHz | 57739 | 53.381 | 57739 | 53.128
| ACV0.000110.01 | ACV | 1 | 10 | 0.0001 | V | mV | 0.01 | 2 | KHz | 3.5 | 0.94 | 3.4755 | 0.93718
| ACV0.00110.01 | ACV | 1 | 10 | 0.001 | V | mV | 0.01 | 2 | KHz | 3.5127 | 0.93465 | 3.5116 | 0.93446
| ACV0.0110.01 | ACV | 1 | 10 | 0.01 | V | mV | 0.01 | 2 | KHz | 6.5091 | 0.7583 | 6.5066 | 0.75746
| ACV0.110.01 | ACV | 1 | 10 | 0.1 | V | mV | 0.01 | 2 | KHz | 57.766 | 0.1391 | 57.765 | 0.13813
| ACV110.01 | ACV | 1 | 10 | 1 | V | mV | 0.01 | 2 | KHz | 577.36 | 0.014397 | 577.36 | 0.014
| ACV0.000112 | ACV | 1 | 10 | 0.0001 | V | mV | 2 | 20 | KHz | 8.2 | 1.8 | 8.1347 | 1.7537
| ACV0.00112 | ACV | 1 | 10 | 0.001 | V | mV | 2 | 20 | KHz | 8.2 | 1.7525 | 8.1525 | 1.7526
| ACV0.0112 | ACV | 1 | 10 | 0.01 | V | mV | 2 | 20 | KHz | 9.8126 | 1.6513 | 9.799 | 1.6514
| ACV0.112 | ACV | 1 | 10 | 0.1 | V | mV | 2 | 20 | KHz | 58.073 | 0.51337 | 58.064 | 0.51216
| ACV112 | ACV | 1 | 10 | 1 | V | mV | 2 | 20 | KHz | 577.39 | 0.054542 | 577.39 | 0.053976
| ACV0.001100.04 | ACV | 10 | 100 | 0.001 | V | mV | 0.04 | 1 | KHz | 35 | 0.95 | 34.943 | 0.94799
| ACV0.01100.04 | ACV | 10 | 100 | 0.01 | V | mV | 0.04 | 1 | KHz | 35.349 | 0.94557 | 35.282 | 0.94527
| ACV0.1100.04 | ACV | 10 | 100 | 0.1 | V | mV | 0.04 | 1 | KHz | 65.076 | 0.76909 | 64.997 | 0.7675
| ACV1100.04 | ACV | 10 | 100 | 1 | V | mV | 0.04 | 1 | KHz | 577.66 | 0.14257 | 577.63 | 0.14057
| ACV10100.04 | ACV | 10 | 100 | 10 | V | mV | 0.04 | 1 | KHz | 5773.6 | 0.015076 | 5773.6 | 0.01425
| ACV0.011000.04 | ACV | 100 | 1020 | 0.01 | V | mV | 0.04 | 1 | KHz | 360 | 1.1 | 337.15 | 1.1167
| ACV0.11000.04 | ACV | 100 | 1020 | 0.1 | V | mV | 0.04 | 1 | KHz | 360 | 1.1141 | 341.11 | 1.1139
| ACV11000.04 | ACV | 100 | 1020 | 1 | V | mV | 0.04 | 1 | KHz | 639.27 | 0.92932 | 638.46 | 0.92803
| ACV101000.04 | ACV | 100 | 1020 | 10 | V | mV | 0.04 | 1 | KHz | 5773 | 0.18489 | 5772.7 | 0.18294
| ACV1001000.04 | ACV | 100 | 1020 | 100 | V | mV | 0.04 | 1 | KHz | 57736 | 0.019421 | 57735 | 0.018613
| ACA0.0000000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-09 | A | µA | 0.01 | 2 | KHz | 0.47 | 1200 | 0.46304 | 1233.9
| ACA0.000000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.47 | 1234 | 0.46308 | 1233.9
| ACA0.00000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.47 | 1227.5 | 0.4666 | 1227
| ACA0.0000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 0.73981 | 857.29 | 0.73979 | 855.77
| ACA0.000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8088 | 394.38 | 5.8088 | 393.5
| ACA0.000000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.97 | 1200 | 0.9653 | 1192.1
| ACA0.00000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.97 | 1191.5 | 0.96691 | 1191.3
| ACA0.0000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 1.1172 | 1119.7 | 1.1168 | 1118.6
| ACA0.000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8401 | 353.65 | 5.8398 | 351.6
| ACA0.00010.00010.01 | ACA | 0.0001 | 0.00104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 57.912 | 316.08 | 57.912 | 315.2
| ACA0.00000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 9.7 | 1200 | 9.653 | 1192.1
| ACA0.0000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 9.7 | 1191.5 | 9.6691 | 1191.3
| ACA0.000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 11.172 | 1119.7 | 11.168 | 1118.6
| ACA0.00010.0010.01 | ACA | 0.001 | 0.0104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 58.401 | 353.65 | 58.398 | 351.6
| ACA0.0010.0010.01 | ACA | 0.001 | 0.0104 | 0.001 | A | µA | 0.01 | 2 | KHz | 579.12 | 316.08 | 579.12 | 315.2
| ACA0.0000010.010.01 | ACA | 0.01 | 0.104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 97 | 1200 | 96.53 | 1192.1
| ACA0.000010.010.01 | ACA | 0.01 | 0.104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 97 | 1191.5 | 96.691 | 1191.3
| ACA0.00010.010.01 | ACA | 0.01 | 0.104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 111.72 | 1119.7 | 111.68 | 1118.6
| ACA0.0010.010.01 | ACA | 0.01 | 0.104 | 0.001 | A | µA | 0.01 | 2 | KHz | 583.96 | 353.65 | 583.98 | 351.6
| ACA0.010.010.01 | ACA | 0.01 | 0.104 | 0.01 | A | µA | 0.01 | 2 | KHz | 5791.2 | 316.08 | 5791.2 | 315.2
| ACA0.000010.10.01 | ACA | 0.1 | 1.04 | 1e-05 | A | mA | 0.01 | 2 | KHz | 1 | 1.5 | 0.99589 | 1.489
| ACA0.00010.10.01 | ACA | 0.1 | 1.04 | 0.0001 | A | mA | 0.01 | 2 | KHz | 1 | 1.4884 | 0.99738 | 1.4882
| ACA0.0010.10.01 | ACA | 0.1 | 1.04 | 0.001 | A | mA | 0.01 | 2 | KHz | 1.1829 | 1.4133 | 1.1409 | 1.4124
| ACA0.010.10.01 | ACA | 0.1 | 1.04 | 0.01 | A | mA | 0.01 | 2 | KHz | 5.9116 | 0.47821 | 5.8383 | 0.47621
| ACA0.10.10.01 | ACA | 0.1 | 1.04 | 0.1 | A | mA | 0.01 | 2 | KHz | 57.947 | 0.32961 | 57.91 | 0.32873
| ACA0.000110.01 | ACA | 1 | 10.4 | 0.0001 | A | mA | 0.01 | 2 | KHz | 18 | 1.4 | 17.128 | 1.4001
| ACA0.00110.01 | ACA | 1 | 10.4 | 0.001 | A | mA | 0.01 | 2 | KHz | 18 | 1.3999 | 17.138 | 1.3997
| ACA0.0110.01 | ACA | 1 | 10.4 | 0.01 | A | mA | 0.01 | 2 | KHz | 18.047 | 1.3635 | 18.045 | 1.3626
| ACA0.110.01 | ACA | 1 | 10.4 | 0.1 | A | mA | 0.01 | 2 | KHz | 60.057 | 0.58243 | 60.055 | 0.58052
| ACA110.01 | ACA | 1 | 10.4 | 1 | A | mA | 0.01 | 2 | KHz | 579.27 | 0.34275 | 579.27 | 0.34186
| Capacitance0.0000000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-09 | mF | nF |  |  |  | 0.036 | 6600 | 0.035476 | 6559.3
| Capacitance0.000000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-08 | mF | nF |  |  |  | 0.036 | 6536.7 | 0.035893 | 6534.2
| Capacitance0.00000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-07 | mF | nF |  |  |  | 0.066425 | 5023.2 | 0.066402 | 5012.6
| Capacitance0.0000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-06 | mF | nF |  |  |  | 0.57863 | 1104.8 | 0.57863 | 1095.3
| Capacitance0.000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-05 | mF | nF |  |  |  | 5.7903 | 3024.4 | 5.7903 | 3020.5
| Capacitance0.000000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-08 | mF | nF |  |  |  | 0.5 | 5800 | 0.48989 | 5801.6
| Capacitance0.00000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-07 | mF | nF |  |  |  | 0.5 | 5788.2 | 0.49306 | 5785.5
| Capacitance0.0000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-06 | mF | nF |  |  |  | 0.74933 | 4690.3 | 0.74914 | 4679.8
| Capacitance0.000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-05 | mF | nF |  |  |  | 5.797 | 1102.8 | 5.797 | 1093.3
| Capacitance0.00010.00001 | Capacitance | 1e-05 | 0.0001 | 0.0001 | mF | nF |  |  |  | 57.904 | 3024.3 | 57.904 | 3020.4
| Capacitance0.00000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-07 | mF | nF |  |  |  | 6.2 | 5200 | 6.1439 | 5209.5
| Capacitance0.0000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-06 | mF | nF |  |  |  | 6.2 | 5201.1 | 6.1698 | 5198.4
| Capacitance0.000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-05 | mF | nF |  |  |  | 8.3825 | 4378.3 | 8.3808 | 4368.2
| Capacitance0.00010.0001 | Capacitance | 0.0001 | 0.001 | 0.0001 | mF | nF |  |  |  | 58.094 | 1100.5 | 58.094 | 1091
| Capacitance0.0010.0001 | Capacitance | 0.0001 | 0.001 | 0.001 | mF | nF |  |  |  | 579.05 | 3024.2 | 579.05 | 3020.3
| Frequency0.0010.001 | Frequency | 0.001 | 100 | 0.001 | kHz | Hz |  |  |  | 0.87 | 0.016 | 0.81638 | 0.016364
| Frequency0.010.001 | Frequency | 0.001 | 100 | 0.01 | kHz | Hz |  |  |  | 5.8379 | 0.0048237 | 5.8243 | 0.0046826
| Frequency0.10.001 | Frequency | 0.001 | 100 | 0.1 | kHz | Hz |  |  |  | 58.005 | 0.0031699 | 58 | 0.0031065
| Frequency10.001 | Frequency | 0.001 | 100 | 1 | kHz | Hz |  |  |  | 579.98 | 0.026531 | 579.97 | 0.026505
| Frequency100.001 | Frequency | 0.001 | 100 | 10 | kHz | Hz |  |  |  | 5799.7 | 0.26461 | 5799.7 | 0.2646
| Insulation DCV11 | Insulation DCV | 1 | 100 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.4054 | 0.010453
| Insulation DCV101 | Insulation DCV | 1 | 100 | 10 | V | V |  |  |  | 5.9149 | 0.0038256 | 5.9069 | 0.0036668
| Insulation DCV1001 | Insulation DCV | 1 | 100 | 100 | V | V |  |  |  | 57.918 | 0.0033656 | 57.915 | 0.0032952
| Insulation DCV1100 | Insulation DCV | 100 | 250 | 1 | V | V |  |  |  | 1.4 | 0.011 | 1.3333 | 0.011162
| Insulation DCV10100 | Insulation DCV | 100 | 250 | 10 | V | V |  |  |  | 5.6899 | 0.0057678 | 5.6798 | 0.0056527
| Insulation DCV100100 | Insulation DCV | 100 | 250 | 100 | V | V |  |  |  | 57.719 | 0.002458 | 57.716 | 0.0023936
| Insulation DCV1250 | Insulation DCV | 250 | 500 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.2696 | 0.011414
| Insulation DCV10250 | Insulation DCV | 250 | 500 | 10 | V | V |  |  |  | 5.1032 | 0.0080276 | 5.0881 | 0.007932
| Insulation DCV100250 | Insulation DCV | 250 | 500 | 100 | V | V |  |  |  | 57.497 | 0.0022225 | 57.502 | 0.0021425
| Insulation DCV 1500 | Insulation DCV | 500 | 1000 | 1 | V | V |  |  |  | 1.8 | 0.011 | 1.2242 | 0.011504
| Insulation DCV 10500 | Insulation DCV | 500 | 1000 | 10 | V | V |  |  |  | 4.1693 | 0.0098749 | 4.0987 | 0.0098767
| Insulation DCV 100500 | Insulation DCV | 500 | 1000 | 100 | V | V |  |  |  | 57.066 | 0.0024926 | 57.041 | 0.0024497
| Insulation Res10250 | Insulation Res | 250 | 100000 | 10 | kΩ | kΩ |  |  |  | 18 | 0.0092 | 13.773 | 0.0092419
| Insulation Res100250 | Insulation Res | 250 | 100000 | 100 | kΩ | kΩ |  |  |  | 57.523 | 0.0088223 | 57.453 | 0.0088228
| Insulation Res1000250 | Insulation Res | 250 | 100000 | 1000 | kΩ | kΩ |  |  |  | 576.26 | 0.0052659 | 576.24 | 0.0052655
| Insulation Res10000250 | Insulation Res | 250 | 100000 | 10000 | kΩ | kΩ |  |  |  | 5775 | 0.0010208 | 5775 | 0.0010204
| Insulation Res10100000 | Insulation Res | 100000 | 250000 | 10 | kΩ | kΩ |  |  |  | 14 | 0.011 | 13.511 | 0.010788
| Insulation Res100100000 | Insulation Res | 100000 | 250000 | 100 | kΩ | kΩ |  |  |  | 57.138 | 0.010616 | 57.125 | 0.010616
| Insulation Res1000100000 | Insulation Res | 100000 | 250000 | 1000 | kΩ | kΩ |  |  |  | 575.4 | 0.0087861 | 575.35 | 0.0087861
| Insulation Res10000100000 | Insulation Res | 100000 | 250000 | 10000 | kΩ | kΩ |  |  |  | 5773 | 0.002517 | 5772.9 | 0.0025167
| Insulation Res100000100000 | Insulation Res | 100000 | 250000 | 100000 | kΩ | kΩ |  |  |  | 57896 | 0.0013129 | 57896 | 0.0013128
| Insulation Res10250000 | Insulation Res | 250000 | 500000 | 10 | kΩ | kΩ |  |  |  | 18 | 0.011 | 12.764 | 0.011009
| Insulation Res100250000 | Insulation Res | 250000 | 500000 | 100 | kΩ | kΩ |  |  |  | 55.079 | 0.010925 | 54.77 | 0.010926
| Insulation Res1000250000 | Insulation Res | 250000 | 500000 | 1000 | kΩ | kΩ |  |  |  | 572.67 | 0.0099501 | 572.64 | 0.0099501
| Insulation Res10000250000 | Insulation Res | 250000 | 500000 | 10000 | kΩ | kΩ |  |  |  | 5771.4 | 0.0044676 | 5771.3 | 0.0044674
| Insulation Res10500000 | Insulation Res | 500000 | 1000000 | 10 | kΩ | kΩ |  |  |  | 25 | 0.011 | -76.51 | 0.011211
| Insulation Res100500000 | Insulation Res | 500000 | 1000000 | 100 | kΩ | kΩ |  |  |  | 25 | 0.011209 | -75.391 | 0.011209
| Insulation Res1000500000 | Insulation Res | 500000 | 1000000 | 1000 | kΩ | kΩ |  |  |  | 25 | 0.011045 | 24.396 | 0.011045
| Insulation Res10000500000 | Insulation Res | 500000 | 1000000 | 10000 | kΩ | kΩ |  |  |  | 5114.5 | 0.007343 | 5114.5 | 0.0073429
| Insulation Res100000500000 | Insulation Res | 500000 | 1000000 | 100000 | kΩ | kΩ |  |  |  | 57638 | 0.0014395 | 57638 | 0.0014395
| TcindicatorJ0.1-346 | TcindicatorJ | -346 | -148 | 0.1 | °F | °F |  |  |  | 2.4168 | 0 | 2.4145 | 0
| TcindicatorJ0.1-148 | TcindicatorJ | -148 | 302 | 0.1 | °F | °F |  |  |  | 1.9165 | 0 | 1.9145 | 0
| TcindicatorJ0.1302 | TcindicatorJ | 302 | 1400 | 0.1 | °F | °F |  |  |  | 2.0393 | 0 | 2.0387 | 0
| TcindicatorJ0.11400 | TcindicatorJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 2.2469 | 0 | 2.2464 | 0
| TcindicatorJ1-346 | TcindicatorJ | -346 | -148 | 1 | °F | °F |  |  |  | 2.4908 | 0 | 2.4819 | 0
| TcindicatorJ1-148 | TcindicatorJ | -148 | 302 | 1 | °F | °F |  |  |  | 2.0067 | 0 | 1.9988 | 0
| TcindicatorJ1302 | TcindicatorJ | 302 | 1400 | 1 | °F | °F |  |  |  | 2.1202 | 0 | 2.1181 | 0
| TcindicatorJ11400 | TcindicatorJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 2.3207 | 0 | 2.3187 | 0
| TcindicatorJ0.1-210 | TcindicatorJ | -210 | -100 | 0.1 | °C | °C |  |  |  | 1.3461 | 0 | 1.3448 | 0
| TcindicatorJ0.1-100 | TcindicatorJ | -100 | 150 | 0.1 | °C | °C |  |  |  | 1.0665 | 0 | 1.0653 | 0
| TcindicatorJ0.1150 | TcindicatorJ | 150 | 760 | 0.1 | °C | °C |  |  |  | 1.1342 | 0 | 1.1339 | 0
| TcindicatorJ0.1760 | TcindicatorJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 1.2495 | 0 | 1.2492 | 0
| TcindicatorJ1-210 | TcindicatorJ | -210 | -100 | 1 | °C | °C |  |  |  | 1.467 | 0 | 1.4624 | 0
| TcindicatorJ1-100 | TcindicatorJ | -100 | 150 | 1 | °C | °C |  |  |  | 1.2144 | 0 | 1.2103 | 0
| TcindicatorJ1150 | TcindicatorJ | 150 | 760 | 1 | °C | °C |  |  |  | 1.2722 | 0 | 1.2711 | 0
| TcindicatorJ1760 | TcindicatorJ | 760 | 1200 | 1 | °C | °C |  |  |  | 1.376 | 0 | 1.375 | 0
| TcindicatorK0.1-328 | TcindicatorK | -328 | -148 | 0.1 | °F | °F |  |  |  | 2.5841 | 0 | 2.5813 | 0
| TcindicatorK0.1-148 | TcindicatorK | -148 | 248 | 0.1 | °F | °F |  |  |  | 2.0844 | 0 | 2.0809 | 0
| TcindicatorK0.1248 | TcindicatorK | 248 | 2500 | 0.1 | °F | °F |  |  |  | 2.5391 | 0 | 2.5376 | 0
| TcindicatorK1-328 | TcindicatorK | -328 | -148 | 1 | °F | °F |  |  |  | 2.6555 | 0 | 2.6444 | 0
| TcindicatorK1-148 | TcindicatorK | -148 | 248 | 1 | °F | °F |  |  |  | 2.1723 | 0 | 2.1587 | 0
| TcindicatorK1248 | TcindicatorK | 248 | 2500 | 1 | °F | °F |  |  |  | 2.6075 | 0 | 2.6018 | 0
| TcindicatorK0.1-200 | TcindicatorK | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.4398 | 0 | 1.4382 | 0
| TcindicatorK0.1-100 | TcindicatorK | -100 | 120 | 0.1 | °C | °C |  |  |  | 1.16 | 0 | 1.1581 | 0
| TcindicatorK0.1120 | TcindicatorK | 120 | 1370 | 0.1 | °C | °C |  |  |  | 1.4121 | 0 | 1.4113 | 0
| TcindicatorK1-200 | TcindicatorK | -200 | -100 | 1 | °C | °C |  |  |  | 1.5545 | 0 | 1.5487 | 0
| TcindicatorK1-100 | TcindicatorK | -100 | 120 | 1 | °C | °C |  |  |  | 1.2998 | 0 | 1.2927 | 0
| TcindicatorK1120 | TcindicatorK | 120 | 1370 | 1 | °C | °C |  |  |  | 1.5267 | 0 | 1.5237 | 0
| TcindicatorN0.1-328 | TcindicatorN | -328 | -148 | 0.1 | °F | °F |  |  |  | 3.205 | 0 | 3.2025 | 0
| TcindicatorN0.1-148 | TcindicatorN | -148 | 770 | 0.1 | °F | °F |  |  |  | 2.2918 | 0 | 2.2894 | 0
| TcindicatorN0.1770 | TcindicatorN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 2.4558 | 0 | 2.4545 | 0
| TcindicatorN1-328 | TcindicatorN | -328 | -148 | 1 | °F | °F |  |  |  | 3.2633 | 0 | 3.2536 | 0
| TcindicatorN1-148 | TcindicatorN | -148 | 770 | 1 | °F | °F |  |  |  | 2.3698 | 0 | 2.3604 | 0
| TcindicatorN1770 | TcindicatorN | 770 | 2372 | 1 | °F | °F |  |  |  | 2.5259 | 0 | 2.5208 | 0
| TcindicatorN0.1-200 | TcindicatorN | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.782 | 0 | 1.7806 | 0
| TcindicatorN0.1-100 | TcindicatorN | -100 | 410 | 0.1 | °C | °C |  |  |  | 1.2762 | 0 | 1.2749 | 0
| TcindicatorN0.1410 | TcindicatorN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 1.3658 | 0 | 1.3651 | 0
| TcindicatorN1-200 | TcindicatorN | -200 | -100 | 1 | °C | °C |  |  |  | 1.8762 | 0 | 1.871 | 0
| TcindicatorN1-100 | TcindicatorN | -100 | 410 | 1 | °C | °C |  |  |  | 1.4032 | 0 | 1.3983 | 0
| TcindicatorN1410 | TcindicatorN | 410 | 1300 | 1 | °C | °C |  |  |  | 1.4837 | 0 | 1.481 | 0
| TcindicatorR0.132 | TcindicatorR | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
| TcindicatorR0.1482 | TcindicatorR | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5823 | 0 | 3.5805 | 0
| TcindicatorR132 | TcindicatorR | 32 | 482 | 1 | °F | °F |  |  |  | 4.8335 | 0 | 4.8271 | 0
| TcindicatorR1482 | TcindicatorR | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6336 | 0 | 3.6263 | 0
| TcindicatorR0.10 | TcindicatorR | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6784 | 0 | 2.6775 | 0
| TcindicatorR0.1250 | TcindicatorR | 250 | 1760 | 0.1 | °C | °C |  |  |  | 1.9965 | 0 | 1.9955 | 0
| TcindicatorR10 | TcindicatorR | 0 | 250 | 1 | °C | °C |  |  |  | 2.7419 | 0 | 2.7384 | 0
| TcindicatorR1250 | TcindicatorR | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0804 | 0 | 2.0765 | 0
| TcindicatorS0.132 | TcindicatorS | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
| TcindicatorS0.1482 | TcindicatorS | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5846 | 0 | 3.5828 | 0
| TcindicatorS132 | TcindicatorS | 32 | 482 | 1 | °F | °F |  |  |  | 4.8334 | 0 | 4.8271 | 0
| TcindicatorS1482 | TcindicatorS | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6358 | 0 | 3.6285 | 0
| TcindicatorS0.10 | TcindicatorS | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6783 | 0 | 2.6775 | 0
| TcindicatorS0.1250 | TcindicatorS | 250 | 1760 | 0.1 | °C | °C |  |  |  | 2.0006 | 0 | 1.9996 | 0
| TcindicatorS10 | TcindicatorS | 0 | 250 | 1 | °C | °C |  |  |  | 2.7418 | 0 | 2.7384 | 0
| TcindicatorS1250 | TcindicatorS | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0844 | 0 | 2.0805 | 0
| TcindicatorT0.1-418 | TcindicatorT | -418 | -238 | 0.1 | °F | °F |  |  |  | 3.9624 | 0 | 3.9602 | 0
| TcindicatorT0.1-238 | TcindicatorT | -238 | 752 | 0.1 | °F | °F |  |  |  | 1.9193 | 0 | 1.9156 | 0
| TcindicatorT1-418 | TcindicatorT | -418 | -238 | 1 | °F | °F |  |  |  | 4.0103 | 0 | 4.0017 | 0
| TcindicatorT1-238 | TcindicatorT | -238 | 752 | 1 | °F | °F |  |  |  | 2.0141 | 0 | 1.9999 | 0
| TcindicatorT0.1-250 | TcindicatorT | -250 | -150 | 0.1 | °C | °C |  |  |  | 2.2146 | 0 | 2.2134 | 0
| TcindicatorT0.1-150 | TcindicatorT | -150 | 400 | 0.1 | °C | °C |  |  |  | 1.0694 | 0 | 1.0674 | 0
| TcindicatorT1-250 | TcindicatorT | -250 | -150 | 1 | °C | °C |  |  |  | 2.2914 | 0 | 2.2868 | 0
| TcindicatorT1-150 | TcindicatorT | -150 | 400 | 1 | °C | °C |  |  |  | 1.2194 | 0 | 1.2121 | 0

## Software Validation
Date: | Action
2015-12-02 | Validated all functions and calculations, repeat readings, uncertainty derivations from table
2017-02-17 | Changed calculations so that any reading with no uncertainty is marked with an asterisk
2018-04-08 | Reordered standard 1 values to fix uncertainty calcualtor errors
2019-05-15 | Added Standard 2, HP 3458A uncertainties
2019-05-30 | Fixed lookup functions so that ACA works properly
2019-06-04 | Fixed error in lookup function
2019-06-04 | Changed 3458A uncertainties to Ω instead of Ohms
2019-07-03 | Added Standards 3 and 4, HP 34420A and Transmille uncertainties
2019-08-23 | Updated 34420A uncertainties
2019-09-13 | Updated 3458A uncertainties

## conversion_factors
Prefix | Name | Conversion Factor
G | Giga | 1000000000
M | Mega | 1000000
k | Kilo | 1000
h | Hecto | 100
da | Deca | 10
d | Deci | 0.1
c | Centi | 0.01
m | Milli | 0.001
µ | Micro | 1e-06
n | Nano | 1e-09
p | Pico | 1e-12