## Worksheet
Johnson Gage & Inspection | [17 empty] | INSPECTOR NOTE: PRINT PAGE 1 ONLY | [6 empty] | µF | µΩ | °C
| Supplementary Data Sheet to Report # |  |  |  |  | Page 2 of 2 | [11 empty] | This section is for Creation of Datasheet Only | [6 empty] | µV | @ | °F
| Fluke 179 Unit # |  |  | Date: | [10 empty] | Secondary Unit for Range field |  |  |  |  | Note: You can now use "O" instead of Ω, "DGC" instead |  |  |  |  | µA |  |  | RESULT | Special Character Filter
[10 empty] | REPEAT READINGS | [9 empty] | of "°C", and "DGF" instead of "°F" if you want to. | [7 empty] | CONV.
Function | Range | Signal Applied | Result | Min | Max | Uncertainty @k=2 | Coil Used | Meter Used | Notes: | 1 | 2 | 3 | 4 | 5 |  | Resolution | Range | Primary Unit | DC/AC | Nominal | tolerance | max | min | (leave blank) | Frequency value | Secondary Unit |  | TO NUMBER
Volts AC | 600 mV | 300 mVAC @ 45Hz |  | 296.7mVAC | 303.3mVAC |  |  | 5502A | [7 empty] | 0.0 | 600 | mV | AC | 300 | 3.3 | 303.3 | 296.7 |  | 45 | Hz | @ 45Hz | 0 | mV | 0 | 303.3 | 296.7 | 0 | 0 | 1
| 6 V | 5 VAC @ 500Hz |  | 4.947VAC | 5.053VAC |  |  | 5502A | [7 empty] | 0.000 | 6 | V | AC | 5 | 0.053 | 5.053 | 4.947 |  | 500 | Hz | @ 500Hz | 0 | V | 0 | 5.053 | 4.947 | 0 | 0 | 1
| 6 V | 5 VAC @ 1kHz |  | 4.897VAC | 5.103VAC |  |  | 5502A | [7 empty] | 0.000 | 6 | V | AC | 5 | 0.103 | 5.103 | 4.897 |  | 1 | kHz | @ 1kHz | 0 | V | 0 | 5.103 | 4.897 | 0 | 0 | 1
| 60 V | 50 VAC @ 45Hz |  | 49.47VAC | 50.53VAC |  |  | 5502A | [7 empty] | 0.00 | 60 | V | AC | 50 | 0.53 | 50.53 | 49.47 |  | 45 | Hz | @ 45Hz | 0 | V | 0 | 50.53 | 49.47 | 0 | 0 | 1
| 60 V | 50 VAC @ 1kHz |  | 48.97VAC | 51.03VAC |  |  | 5502A | [7 empty] | 0.00 | 60 | V | AC | 50 | 1.03 | 51.03 | 48.97 |  | 1 | kHz | @ 1kHz | 0 | V | 0 | 51.03 | 48.97 | 0 | 0 | 1
| 600 V | 300 VAC @ 45Hz |  | 296.7VAC | 303.3VAC |  |  | 5502A | [7 empty] | 0.0 | 600 | V | AC | 300 | 3.3 | 303.3 | 296.7 |  | 45 | Hz | @ 45Hz | 0 | V | 0 | 303.3 | 296.7 | 0 | 0 | 1
| 600 V | 500 VAC @ 500Hz |  | 494.7VAC | 505.3VAC |  |  | 5502A | [7 empty] | 0.0 | 600 | V | AC | 500 | 5.3 | 505.3 | 494.7 |  | 500 | Hz | @ 500Hz | 0 | V | 0 | 505.3 | 494.7 | 0 | 0 | 1
| 600 V | 500 VAC @ 1kHz |  | 489.7VAC | 510.3VAC |  |  | 5502A | [7 empty] | 0.0 | 600 | V | AC | 500 | 10.3 | 510.3 | 489.7 |  | 1 | kHz | @ 1kHz | 0 | V | 0 | 510.3 | 489.7 | 0 | 0 | 1
| 1000 V | 1000 VAC @ 45Hz |  | 987VAC | 1013VAC |  |  | 5502A | [7 empty] | 0 | 1000 | V | AC | 1000 | 13 | 1013 | 987 |  | 45 | Hz | @ 45Hz | 0 | V | 0 | 1013 | 987 | 0 | 0 | 1
Volts AC | 99.99 Hz | 1V @ 45Hz |  | 44.94Hz | 45.06Hz |  |  | 5502A | [7 empty] | 0.00 | 99.99 | Hz |  | 45 | 0.06 | 45.06 | 44.94 |  | 1 | V | 1V | 0 | Hz | 0 | 45.06 | 44.94 | 0 | 0 | 1
Frequency | 99.99 kHz | 5V @ 50kHz |  | 49.94kHz | 50.06kHz |  |  | 5502A | [7 empty] | 0.00 | 99.99 | kHz |  | 50 | 0.06 | 50.06 | 49.94 |  | 5 | V | 5V | 0 | kHz | 0 | 50.06 | 49.94 | 0 | 0 | 1
Volts DC | 6 V | 5 VDC |  | 4.993VDC | 5.007VDC |  |  | 5502A | [7 empty] | 0.000 | 6 | V | DC | 5 | 0.007 | 5.007 | 4.993 |  |  |  |  | 0 | V | 0 | 5.007 | 4.993 | 0 | 0 | 1
| 600 V | 300 VDC |  | 299.5VDC | 300.5VDC |  |  | 5502A | [7 empty] | 0.0 | 600 | V | DC | 300 | 0.5 | 300.5 | 299.5 |  |  |  |  | 0 | V | 0 | 300.5 | 299.5 | 0 | 0 | 1
| 1000 V | 1000 VDC |  | 997VDC | 1003VDC |  |  | 5502A | [7 empty] | 0 | 1000 | V | DC | 1000 | 3 | 1003 | 997 |  |  |  |  | 0 | V | 0 | 1003 | 997 | 0 | 0 | 1
| 1000 V | -1000 VDC |  | -1003VDC | -997VDC |  |  | 5502A | [7 empty] | 0 | 1000 | V | DC | -1000 | 3 | -997 | -1003 |  |  |  |  | 0 | V | 0 | -997 | -1003 | 0 | 1 | 0
Volts DC | 99.99 Hz | 3V @ 45Hz |  | 44.94Hz | 45.06Hz |  |  | 5502A | [7 empty] | 0.00 | 99.99 | Hz |  | 45 | 0.06 | 45.06 | 44.94 |  | 3 | V | 3V | 0 | Hz | 0 | 45.06 | 44.94 | 0 | 0 | 1
Frequency | 99.99 kHz | 30V @ 50kHz |  | 49.94kHz | 50.06kHz |  |  | 5502A | [7 empty] | 0.00 | 99.99 | kHz |  | 50 | 0.06 | 50.06 | 49.94 |  | 30 | V | 30V | 0 | kHz | 0 | 50.06 | 49.94 | 0 | 0 | 1
Millivolts DC | 600 mV | 30 mVDC |  | 29.8mVDC | 30.2mVDC |  |  | 5502A | [7 empty] | 0.0 | 600 | mV | DC | 30 | 0.2 | 30.2 | 29.8 |  |  |  |  | 0 | mV | 0 | 30.2 | 29.8 | 0 | 0 | 1
| 600 mV | -300 mVDC |  | -300.5mVDC | -299.5mVDC |  |  | 5502A | [7 empty] | 0.0 | 600 | mV | DC | -300 | 0.5 | -299.5 | -300.5 |  |  |  |  | 0 | mV | 0 | -299.5 | -300.5 | 0 | 1 | 0
| 600 mV | 600 mVDC |  | 599.3mVDC | 600.7mVDC |  |  | 5502A | [7 empty] | 0.0 | 600 | mV | DC | 600 | 0.7 | 600.7 | 599.3 |  |  |  |  | 0 | mV | 0 | 600.7 | 599.3 | 0 | 0 | 1
Resistance | 600 Ω | 500 Ω |  | 495.3Ω | 504.7Ω |  |  | 5502A | [7 empty] | 0.0 | 600 | Ω |  | 500 | 4.7 | 504.7 | 495.3 |  |  |  |  | 0 | Ω | 0 | 504.7 | 495.3 | 0 | 0 | 1
| 50 MΩ | 19 MΩ |  | 18.68MΩ | 19.32MΩ |  |  | 5502A | [7 empty] | 0.00 | 50 | MΩ |  | 19 | 0.32 | 19.32 | 18.68 |  |  |  |  | 0 | MΩ | 0 | 19.32 | 18.68 | 0 | 0 | 1
Capacitance | 1000 nF | 900 nF |  | 887nF | 913nF |  |  | 5502A | [7 empty] | 0 | 1000 | nF |  | 900 | 13 | 913 | 887 |  |  |  |  | 0 | nF | 0 | 913 | 887 | 0 | 0 | 1
Continuity | 600 Ω | 25 Ω |  | Tone ON |  |  |  | 5502A | [7 empty] | 0.0 | 600 | Ω |  | 25 |  | 25.0 | 25.0 |  |  |  |  | 0 | Ω | 0 | 25 | 25 | 0 | 0 | 1
| 600 Ω | 250 Ω |  | Tone OFF |  |  |  | 5502A | [7 empty] | 0.0 | 600 | Ω |  | 250 |  | 250.0 | 250.0 |  |  |  |  | 0 | Ω | 0 | 250 | 250 | 0 | 0 | 1
Diode | 2.400 V | 2.0 VDC |  | 1.978VDC | 2.022VDC |  |  | 5502A | [7 empty] | 0.000 | 2.400 | V | DC | 2.0 | 0.022 | 2.022 | 1.978 |  |  |  |  | 0 | V | 0 | 2.022 | 1.978 | 0 | 0 | 1
Milliamps AC | 60 mA | 3 mAAC @ 45Hz |  | 2.92mAAC | 3.08mAAC |  |  | 5502A | [7 empty] | 0.00 | 60 | mA | AC | 3 | 0.08 | 3.08 | 2.92 |  | 45 | Hz | @ 45Hz | 0 | mA | 0 | 3.08 | 2.92 | 0 | 0 | 1
| 60 mA | 50 mAAC @ 1kHz |  | 49.22mAAC | 50.78mAAC |  |  | 5502A | [7 empty] | 0.00 | 60 | mA | AC | 50 | 0.78 | 50.78 | 49.22 |  | 1 | kHz | @ 1kHz | 0 | mA | 0 | 50.78 | 49.22 | 0 | 0 | 1
| 400 mA | 400 mAAC @ 1kHz |  | 393.7mAAC | 406.3mAAC |  |  | 5502A | [7 empty] | 0.0 | 400 | mA | AC | 400 | 6.3 | 406.3 | 393.7 |  | 1 | kHz | @ 1kHz | 0 | mA | 0 | 406.3 | 393.7 | 0 | 0 | 1
Amps AC | 6 A | 4 AAC @ 45Hz |  | 3.937AAC | 4.063AAC |  |  | 5502A | [7 empty] | 0.000 | 6 | A | AC | 4 | 0.063 | 4.063 | 3.937 |  | 45 | Hz | @ 45Hz | 0 | A | 0 | 4.063 | 3.937 | 0 | 0 | 1
| 10 A | 9 AAC @ 1kHz |  | 8.83AAC | 9.17AAC |  |  | 5502A | [7 empty] | 0.00 | 10 | A | AC | 9 | 0.17 | 9.17 | 8.83 |  | 1 | kHz | @ 1kHz | 0 | A | 0 | 9.17 | 8.83 | 0 | 0 | 1
Milliamps DC | 60 mA | 3 mADC |  | 2.94mADC | 3.06mADC |  |  | 5502A | [7 empty] | 0.00 | 60 | mA | DC | 3 | 0.06 | 3.06 | 2.94 |  |  |  |  | 0 | mA | 0 | 3.06 | 2.94 | 0 | 0 | 1
| 60 mA | 50 mADC |  | 49.47mADC | 50.53mADC |  |  | 5502A | [7 empty] | 0.00 | 60 | mA | DC | 50 | 0.53 | 50.53 | 49.47 |  |  |  |  | 0 | mA | 0 | 50.53 | 49.47 | 0 | 0 | 1
| 400 mA | -400 mADC |  | -404.3mADC | -395.7mADC |  |  | 5502A | [7 empty] | 0.0 | 400 | mA | DC | -400 | 4.3 | -395.7 | -404.3 |  |  |  |  | 0 | mA | 0 | -395.7 | -404.3 | 0 | 1 | 0
Amps DC | 6 A | 4 ADC |  | 3.957ADC | 4.043ADC |  |  | 5502A | [7 empty] | 0.000 | 6 | A | DC | 4 | 0.043 | 4.043 | 3.957 |  |  |  |  | 0 | A | 0 | 4.043 | 3.957 | 0 | 0 | 1
| 10 A | -9 ADC |  | -9.12ADC | -8.88ADC |  |  | 5502A | [7 empty] | 0.00 | 10 | A | DC | -9 | 0.12 | -8.88 | -9.12 |  |  |  |  | 0 | A | 0 | -8.88 | -9.12 | 0 | 1 | 0
Temperature | 400 °C | 0 °C |  | -1°C | 1°C |  |  | 5502A | [7 empty] | 0 | 400 | °C |  | 0 | 1 | 1 | -1 |  |  |  |  | 0 | °C | 0 | 1 | -1 | 1 | 1 | 1
| 400 °C | -40 °C |  | -41.4°C | -38.6°C |  |  | 5502A | [7 empty] | 0.0 | 400 | °C |  | -40 | 1.4 | -38.6 | -41.4 |  |  |  |  | 0 | °C | 0 | -38.6 | -41.4 | 0 | 1 | 0
| 400 °C | 400 °C |  | 395°C | 405°C |  |  | 5502A | [7 empty] | 0 | 400 | °C |  | 400 | 5 | 405 | 395 |  |  |  |  | 0 | °C | 0 | 405 | 395 | 0 | 0 | 1
Temperature |  | Temperature | 70°F |  |  | * | [15 empty] | 0 | 0 |  |  |  |  | #VALUE! | 0 | #VALUE! | 0 | 0 | #VALUE! | #VALUE! | #VALUE!
Condtion |  | Humidity | 40% |  |  | * | [15 empty] | 0 | 0 |  |  |  |  | 0.4 | 0 | 0.4 | 0 | 0 | 0 | 1 | 0
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1
|  | 0 |  | 00 | 00 | [16 empty] | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1

## Unc. Calculator
RED=Key Calculator Component |  | 5502A | Standard1
GREEN=Formula Uses a Table Lookup |  | Transmille | Standard2 |  | Troubleshooter Lookup | [11 empty] | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | [38 empty] | ACV Freq. Rng Low | ACI Freq. Rng Low | ACA10T. Rng low | ACA20T. Rng low | ACA50T. Rng Low
ORANGE= END USER RAW INPUTS |  | 3458A | Standard3 |  | Rdg | L.up ID | Standard | Base | Mult | 2nd B | 2nd Mult | Rep. Mult | Unc. Units | Result |  | Standard6 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [13 empty] | 1000 | [23 empty] | Standard1 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A
BLUE= PART OF MAIN LOOKUP ID |  | 34420A | Standard4 |  | 0.003 | DCA0.000010 | Standard1 | 5.8001 | 1.6501 | 5.8001 | 1.5718 | 0 | µA | 5.8050503µA |  | Standard5 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [37 empty] | Standard2 | $I$80:$I$119 | $I$120:$I$149 | #N/A | #N/A | #N/A
|  |  | Standard5 | [12 empty] | Standard4 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [37 empty] | Standard3 | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  | Standard6 | [12 empty] | Standard3 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | [37 empty] | Standard4 | #N/A | #N/A | #N/A | #N/A | #N/A | Col. Index # for frequency Lookup | Input Unit | Unc units
|  |  | Standard7 | [12 empty] | Standard2 | $D$4:$D$28 | $D$80:$D$119 | $D$29:$D$58 | $D$120:$D$149 | #N/A | $D$150:$D$164 | $D$165:$D$169 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Unit Multiplier | [13 empty] | 2ndary Unit Multiplier | [22 empty] | Standard5 | #N/A | #N/A | #N/A | #N/A | #N/A |  | V11 | CN11
[16 empty] | Standard1 | $D$4:$D$32 | $D$189:$D$348 | $D$33:$D$66 | $D$373:$D$538 | $D$121:$D$188 | $D$744:$D$785 | $D$349:$D$809 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | Base unit | [13 empty] | Base Unit | [18 empty] | RESOLUTION CONVERTER |  |  |  | Standard6 | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  | USER INPUTS |  |  |  |  | Function Selector | [8 empty] | MAIN INPUT FUNCTION LOOKUP | [12 empty] | V | V | Amps | Amps | kOhms | mF | kHz | Amps | Amps | Amps | Amps | Amps | Amps |  |  |  |  | LOOKED UP UNC. COMPONENTS |  |  |  |  | REPEAT USER INPUTS |  |  |  |  | CONVERTED REPEAT ENTRIES | [5 empty] | "0.000 to 0.001" | [10 empty] | Repeat Multiplier Ratio
Main Lookup ID | Standard Used | Unc. | Res. | In. Value | Units | Range 2 | Units | L.upbldr | L.upbldr | L.upbldr | L.upbldr | L.upbldr | L.upbldr | Func.Sel. | Main L.up | Converted Rdg | DCV | ACV | DCA | ACA | Resistance | Capacitance | Frequency | DCA10TURN | DCA20TURN | DCA50TURN | ACA10TURN | ACA20TURN | ACA50TURN | DCV | ACV | DCA | ACA | Resistance | Capacitance | Frequency | DCA10TURN | DCA20TURN | DCA50TURN | ACA10TURN | ACA20TURN | ACA50TURN | Primary Conv. Factor | kHz | Converted Rdg | 2ndary L. Up | Base | Mult | 2nd B. | 2nd Mult. | Rept Mult | R1 | R2 | R3 | R4 | R5 | R1(Mult) | R2(Mult) | R3(Mult) | R4(Mult) | R5(Mult) | Unc. Units | Input | Conv. Txt | Result |  | Unc. Table Address for L. Ups | ACV | ACA | ACA10TURN | ACA20TURN | ACA50TURN |  | Input | Uncert. | Multiplier
ACV0.00010.0330.045 | 5502A | 140 µV | 0.0001 | 300 | mV | 45 | Hz | m | V | AC | ACV |  |  | ACV | 0.033 | 0.3 | 0 | 0.033 | 0.033 | 0.033 | 0.11 | 0.11 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.001 | 0.045 | 0.045 | 58.586 | 276.03 | 58.483 | 276.21 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 2 | 0.001 | 1e-06
ACV0.0013.30.045 | 5502A | 2.5 mV | 0.001 | 5 | V | 500 | Hz |  | V | AC | ACV |  |  | ACV | 3.3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.5 | 0.045 | 0.78957 | 0.34543 | 0.78891 | 0.34543 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 3 | 1 | 0.001
ACV0.0013.30.045 | 5502A | 2.5 mV | 0.001 | 5 | V | 1 | kHz |  | V | AC | ACV |  |  | ACV | 3.3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.045 | 0.78957 | 0.34543 | 0.78891 | 0.34543 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 4 | 1 | 0.001
ACV0.01330.045 | 5502A | 58 mV | 0.01 | 50 | V | 45 | Hz |  | V | AC | ACV |  |  | ACV | 33 | 50 | 33 | 33 | 11 | 11 | 33 | 33 | 30 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 12 | 0.92331 | 10.834 | 0.9234 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.00 | 0.01 | 0.01 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 5 | 1 | 0.001
ACV0.01330.045 | 5502A | 58 mV | 0.01 | 50 | V | 1 | kHz |  | V | AC | ACV |  |  | ACV | 33 | 50 | 33 | 33 | 11 | 11 | 33 | 33 | 30 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.045 | 12 | 0.92331 | 10.834 | 0.9234 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.00 | 0.01 | 0.01 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 6 | 1 | 0.001
ACV0.1330.045 | 5502A | 300 mV | 0.1 | 300 | V | 45 | Hz |  | V | AC | ACV |  |  | ACV | 33 | 300 | 33 | 33 | 11 | 11 | 110 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 43.309 | 0.841 | 43.21 | 0.84115 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.0 | 0.1 | 0.1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 7 | 1 | 0.001
ACV0.13300.045 | 5502A | 670 mV | 0.1 | 500 | V | 500 | Hz |  | V | AC | ACV |  |  | ACV | 330 | 500 | 330 | 330 | 11 | 11 | 330 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.5 | 0.045 | 510.18 | 0.32785 | 510.26 | 0.32724 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.0 | 0.1 | 0.1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 8 | 1 | 0.001
ACV0.13300.045 | 5502A | 670 mV | 0.1 | 500 | V | 1 | kHz |  | V | AC | ACV |  |  | ACV | 330 | 500 | 330 | 330 | 11 | 11 | 330 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.045 | 510.18 | 0.32785 | 510.26 | 0.32724 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.0 | 0.1 | 0.1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 9 | 1 | 0.001
ACV13300.045 | 5502A | 5800 mV | 1 | 1000 | V | 45 | Hz |  | V | AC | ACV |  |  | ACV | 330 | 1000 | 330 | 330 | 11 | 11 | 330 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 5790.3 | 0.041506 | 5790.4 | 0.041123 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0 | 1 | 1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 10 | 1 | 0.001
Frequency0.000010.033 | 5502A | * | 1e-05 | 45 | Hz | 1 | V | H | z | AC | ACz |  | Frequency | Frequency | 0.033 | 0.045 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.00 | 0.01 | 1e-05 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 11 | 1 | #N/A
Frequency0.0130 | 5502A | * | 0.01 | 50 | kHz | 5 | V | H | z | AC | ACz |  | Frequency | Frequency | 30 | 50 | 33 | 33 | 11 | 11 | 33 | 33 | 30 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 5 |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.00 | 0.01 | 0.01 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 12 | 1000 | #N/A
DCV0.0013 | 5502A | 0.72 mV | 0.001 | 5 | V |  |  |  | V | DC | DCV |  |  | DCV | 3 | 5 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.47784 | 0.047829 | 0.47304 | 0.047859 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 13 | 1 | 0.001
DCV0.133 | 5502A | 62 mV | 0.1 | 300 | V |  |  |  | V | DC | DCV |  |  | DCV | 33 | 300 | 33 | 33 | 11 | 11 | 110 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 57.654 | 0.012974 | 57.637 | 0.012872 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0.0 | 0.1 | 0.1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 14 | 1 | 0.001
DCV1330 | 5502A | 580 mV | 1 | 1000 | V |  |  |  | V | DC | DCV |  |  | DCV | 330 | 1000 | 330 | 330 | 11 | 11 | 330 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0 | 1 | 1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 15 | 1 | 0.001
DCV1330 | 5502A | 580 mV | 1 | 1000 | V |  |  |  | V | DC | DCV |  |  | DCV | 330 | 1000 | 330 | 330 | 11 | 11 | 330 | 33 | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mV | 0 | 1 | 1 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 16 | 1 | 0.001
Frequency0.000010.033 | 5502A | * | 1e-05 | 45 | Hz | 3 | V | H | z | AC | ACz |  | Frequency | Frequency | 0.033 | 0.045 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 | 3 |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.00 | 0.01 | 1e-05 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 17 | 1 | #N/A
Frequency0.0130 | 5502A | * | 0.01 | 50 | kHz | 30 | V | H | z | AC | ACz |  | Frequency | Frequency | 30 | 50 | 33 | 33 | 11 | 11 | 33 | 33 | 30 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 30 |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.00 | 0.01 | 0.01 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 18 | 1000 | #N/A
DCV0.00010 | 5502A | 59 µV | 0.0001 | 30 | mV |  |  | m | V | DC | DCV |  |  | DCV | 0 | 0.03 | 0 | 0.001 | 0.0033 | 0.0033 | 0.011 | 0.011 | 0.0033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 58.11 | 17.037 | 58.091 | 16.984 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 19 | 0.001 | 1e-06
DCV0.00010 | 5502A | 63 µV | 0.0001 | 300 | mV |  |  | m | V | DC | DCV |  |  | DCV | 0 | 0.3 | 0 | 0.033 | 0.033 | 0.033 | 0.11 | 0.11 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 58.11 | 17.037 | 58.091 | 16.984 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 20 | 0.001 | 1e-06
DCV0.00010.33 | 5502A | 76 µV | 0.0001 | 600 | mV |  |  | m | V | DC | DCV |  |  | DCV | 0.33 | 0.6 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.001 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 47.82 | 47.701 | 47.342 | 47.73 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 21 | 0.001 | 1e-06
Resistance0.00010.33 | 5502A | 97 mΩ | 0.0001 | 500 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0.33 | 0.5 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 52.968 | 87.892 | 52.867 | 87.789 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 22 | 1 | 0.001
Resistance1011000 | 5502A | 26 kΩ | 10 | 19 | MΩ |  |  | M | O | DC | O |  | Resistance | Resistance | 11000 | 19000 | 330 | 330 | 11 | 11 | 11000 | 33 | 1200 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1000 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1000 | 1 |  |  | 4.2484 | 0.0011259 | 4.2478 | 0.0011259 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | kΩ | 0.00 | 0.01 | 10 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 23 | 1000000 | 1000
Capacitance0.0000010.00033 | 5502A | * | 1e-06 | 900 | nF |  |  | n | F | DC | DCF |  | Capacitance | Capacitance | 0.00033 | 0.0009 | 0 | #N/A | 0.00033 | 0.00033 | 0 | 0.00033 | 0.00033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1e-06 | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 1 | 1e-06 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 24 | 1e-09 | #N/A
Resistance0.00010.011 | 5502A | 62 mΩ | 0.0001 | 25 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0.011 | 0.025 | 0 | 0.001 | 0.0033 | 0.0033 | 0.011 | 0.011 | 0.0033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 60.487 | 46.304 | 60.481 | 46.04 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 25 | 1 | 0.001
Resistance0.00010.11 | 5502A | 110 mΩ | 0.0001 | 250 | Ω |  |  |  | O | DC | O |  | Resistance | Resistance | 0.11 | 0.25 | 0 | 0.033 | 0.033 | 0.033 | 0.11 | 0.11 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 0.001 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 1 |  |  | 61.466 | 184.35 | 61.46 | 183.05 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mΩ | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 26 | 1 | 0.001
DCV0.0010.33 | 5502A | 600 µV | 0.001 | 2 | V |  |  |  | V | DC | DCV |  |  | DCV | 0.33 | 2 | 0.33 | 0.33 | 0.33 | 1.1 | 1.1 | 1.1 | 1.1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 577.17 | 10.824 | 576.99 | 10.721 | 1000000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µV | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 27 | 1 | 1e-06
ACA0.000010.000330.045 | 5502A | 6.9 µA | 1e-05 | 3 | mA | 45 | Hz | m | A | AC | ACA |  |  | ACA | 0.00033 | 0.003 | 0 | 0.001 | 0.00033 | 0.00033 | 0 | 0.0011 | 0.00033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.045 | 0.045 | 5.6654 | 409.29 | 5.6651 | 409.25 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µA | 0.00 | 0.01 | 1e-05 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 28 | 0.001 | 1e-06
ACA0.000010.0331 | 5502A | 0.12 mA | 1e-05 | 50 | mA | 1 | kHz | m | A | AC | ACA |  |  | ACA | 0.033 | 0.05 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 1 | 0.058 | 1.2 | 0.057825 | 1.1583 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.00 | 0.01 | 1e-05 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 29 | 0.001 | 0.001
ACA0.00010.331 | 5502A | 4 mA | 0.0001 | 400 | mA | 1 | kHz | m | A | AC | ACA |  |  | ACA | 0.33 | 0.4 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 | 1 | 1 | 1.2 | 6.9 | 1.1555 | 6.9278 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 30 | 0.001 | 0.001
ACA0.00130.045 | 5502A | 5.2 mA | 0.001 | 4 | A | 45 | Hz |  | A | AC | ACA |  |  | ACA | 3 | 4 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0.001 | 0.045 | 0.045 | 2.4 | 0.69 | 2.3709 | 0.69112 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 31 | 1 | 0.001
ACA0.0131 | 5502A | 320 mA | 0.01 | 9 | A | 1 | kHz |  | A | AC | ACA |  |  | ACA | 3 | 9 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 2.5 | 35 | 2.5084 | 34.627 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.00 | 0.01 | 0.01 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 32 | 1 | 0.001
DCA0.000010.00033 | 5502A | 5.8 µA | 1e-05 | 3 | mA |  |  | m | A | DC | DCA |  |  | DCA | 0.00033 | 0.003 | 0 | 0.001 | 0.00033 | 0.00033 | 0 | 0.0011 | 0.00033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 |  |  | 5.7989 | 9.0557 | 5.7981 | 8.8116 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µA | 0.00 | 0.01 | 1e-05 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 33 | 0.001 | 1e-06
DCA0.000010.033 | 5502A | 11 µA | 1e-05 | 50 | mA |  |  | m | A | DC | DCA |  |  | DCA | 0.033 | 0.05 | 0 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 |  |  | 5.2853 | 109.48 | 5.266 | 109.51 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | µA | 0.00 | 0.01 | 1e-05 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 34 | 0.001 | 1e-06
DCA0.00010.33 | 5502A | 0.23 mA | 0.0001 | 400 | mA |  |  | m | A | DC | DCA |  |  | DCA | 0.33 | 0.4 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 0.001 | 0.001 | 1 | 1 | 1 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 1 |  |  | 0.060249 | 0.43614 | 0.060204 | 0.43607 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.0 | 0.1 | 0.0001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 35 | 0.001 | 0.001
DCA0.0013 | 5502A | 3.7 mA | 0.001 | 4 | A |  |  |  | A | DC | DCA |  |  | DCA | 3 | 4 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 0.055271 | 0.90973 | 0.052344 | 0.90974 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.000 | 0.001 | 0.001 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 36 | 1 | 0.001
DCA0.013 | 5502A | 10 mA | 0.01 | 9 | A |  |  |  | A | DC | DCA |  |  | DCA | 3 | 9 | 3 | 3.3 | 3 | 3 | 3.3 | 3.3 | 1 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  | 4.4629 | 0.64947 | 4.4597 | 0.64887 | 1000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | mA | 0.00 | 0.01 | 0.01 | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 37 | 1 | 0.001
#N/A | 5502A | * | #N/A | 0 | °C |  |  |  | DGC | DC | DGC |  |  | DGC | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 38 | 1 | #N/A
#N/A | 5502A | * | #N/A | 40 | °C |  |  |  | DGC | DC | DGC |  |  | DGC | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0.0 | 0.1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 39 | 1 | #N/A
#N/A | 5502A | * | #N/A | 400 | °C |  |  |  | DGC | DC | DGC |  |  | DGC | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 40 | 1 | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 41 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 42 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 43 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 44 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 45 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 46 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 47 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 48 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 49 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 50 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 51 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 52 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 53 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 54 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 55 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 56 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 57 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 58 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 59 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 60 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 61 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 62 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 63 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 64 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 65 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 66 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 67 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 68 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 69 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 70 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 71 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 72 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 73 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 74 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 75 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 76 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 77 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 78 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 79 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 80 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 81 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 82 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 83 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 84 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 85 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 86 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 87 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 88 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 89 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 90 |  | #N/A
#N/A | 5502A | * | #N/A | 0 | 0 |  |  |  | 0 | DC | DC0 |  |  | DC0 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | #N/A | 1 |  |  | #N/A | #N/A | #N/A | #N/A | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #N/A | 0 | 1 | #N/A | Standard1 | $B$4:$P$936 | $I$189:$I$348 | $I$373:$I$538 | #N/A | #N/A | #N/A | 91 |  | #N/A
Note on S1: Lookup table thatTranslates Standard Name to Generic name. i.e. "5502A" to "Standard1". See Colum "CQ"
Note on BA7: Columns AU-BA convert the reading entered into the standard unit of the table for each function
Note on BV7: Converts the 2ndary units (usually frequency) to the standard table format (kHz)
Note on T9: User inputs. Resolution is converted from "0.0" format to "0.1" format in Column CA-CE, and Column T references that
Note on Y9: Columns Y-AD select which function (column C) is used based on the inputs
Note on AH9: Columns AG- AM determine in what range the primary unit falls (Ex: 5VDC falls in the 3.3-30V Range, so it returns 3.3, as it is set to show the "Range Low" part as a marker
Note on DB9: This is where the ratio between the Input units and the Uncertainty Units is created to generate the repeat multiplier in CC11
Note on Q10: Constructed input ID in the form of Function,resolution,InputRangeLOW,InputRange2LOW
Note on T10: From Column CE
Note on U10: Returns the Absolute value of the reading (filters out Negative Numbers)
Note on Z10: This is where special characters get filtered out and converted to whatever we designate. (i.e. Ω changes to O)
Note on AE10: Determines which category in Column "C" the reading falls under
Note on AF10: Column AE Retrieves the appropriate "Range Low" result from Columns AG-AM, based on the Function Selector in Column AD. It is an Hlookup formula, desgined to return the Nth row down, where N is the number of rows
Note on AG10: Column AF Converts Column AE to whatever the standard unit is in the table, based on a multiplier in Column BI
Note on BU10: Looks up the correct converted Unit based on the function selector
Note on BV10: 2ndary unit conv. factor
Note on BX10: Returns the appropriate "Range 2 Low" value (usually frequency) from the table. Part of the INPUT ID that looks up the Uncertainty components
Note on CQ10: Final Resolution Input: Part of the INPUT ID used for uncertainty Lookup

## Standard1
Updated: | 2021-02-26 | 5502A
|  | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.00000010 | DCV | 0 | 0.33 | 1e-07 | V | µV |  |  |  | 3.6 | 69 | 3.4727 | 69.257
| DCV0.0000010 | DCV | 0 | 0.33 | 1e-06 | V | µV |  |  |  | 3.6 | 69.095 | 3.5194 | 69.135
| DCV0.000010 | DCV | 0 | 0.33 | 1e-05 | V | µV |  |  |  | 6.7738 | 61.231 | 6.7344 | 61.286
| DCV0.00010 | DCV | 0 | 0.33 | 0.0001 | V | µV |  |  |  | 58.11 | 17.037 | 58.091 | 16.984
| DCV0.0010 | DCV | 0 | 0.33 | 0.001 | V | µV |  |  |  | 580.02 | 1.8079 | 580.01 | 1.7824
| DCV0.010 | DCV | 0 | 0.33 | 0.01 | V | µV |  |  |  | 5800 | 0.18854 | 5800 | 0.17833
| DCV0.0000010.33 | DCV | 0.33 | 3.3 | 1e-06 | V | µV |  |  |  | 5.9 | 58 | 5.7708 | 57.787
| DCV0.000010.33 | DCV | 0.33 | 3.3 | 1e-05 | V | µV |  |  |  | 6.8066 | 57.529 | 6.4957 | 57.593
| DCV0.00010.33 | DCV | 0.33 | 3.3 | 0.0001 | V | µV |  |  |  | 47.82 | 47.701 | 47.342 | 47.73
| DCV0.0010.33 | DCV | 0.33 | 3.3 | 0.001 | V | µV |  |  |  | 577.17 | 10.824 | 576.99 | 10.721
| DCV0.010.33 | DCV | 0.33 | 3.3 | 0.01 | V | µV |  |  |  | 5799.8 | 1.1473 | 5799.7 | 1.1022
| DCV0.10.33 | DCV | 0.33 | 3.3 | 0.1 | V | µV |  |  |  | 58000 | 0.12833 | 5799.8 | 0.11025
| DCV0.000013 | DCV | 3 | 33 | 1e-05 | V | mV |  |  |  | 0.059 | 0.058 | 0.058348 | 0.057919
| DCV0.00013 | DCV | 3 | 33 | 0.0001 | V | mV |  |  |  | 0.067758 | 0.057661 | 0.064652 | 0.057725
| DCV0.0013 | DCV | 3 | 33 | 0.001 | V | mV |  |  |  | 0.47784 | 0.047829 | 0.47304 | 0.047859
| DCV0.013 | DCV | 3 | 33 | 0.01 | V | mV |  |  |  | 5.7715 | 0.010868 | 5.7698 | 0.010764
| DCV0.13 | DCV | 3 | 33 | 0.1 | V | mV |  |  |  | 57.998 | 0.0011519 | 57.997 | 0.0011068
| DCV13 | DCV | 3 | 33 | 1 | V | mV |  |  |  | 580 | 0.00012879 | 580 | 0.00011071
| DCV0.000133 | DCV | 33 | 330 | 0.0001 | V | mV |  |  |  | 0.59 | 0.064 | 0.57834 | 0.063644
| DCV0.00133 | DCV | 33 | 330 | 0.001 | V | mV |  |  |  | 0.67471 | 0.063403 | 0.64578 | 0.063463
| DCV0.0133 | DCV | 33 | 330 | 0.01 | V | mV |  |  |  | 4.6629 | 0.053693 | 4.6153 | 0.053731
| DCV0.133 | DCV | 33 | 330 | 0.1 | V | mV |  |  |  | 57.654 | 0.012974 | 57.637 | 0.012872
| DCV133 | DCV | 33 | 330 | 1 | V | mV |  |  |  | 579.97 | 0.0013757 | 579.96 | 0.0013305
| DCV1033 | DCV | 33 | 330 | 10 | V | mV |  |  |  | 5800 | 0.00015118 | 5800 | 0.0001331
| DCV0.001330 | DCV | 330 | 1000 | 0.001 | V | mV |  |  |  | 1.9 | 0.064 | 1.7549 | 0.063615
| DCV0.01330 | DCV | 330 | 1000 | 0.01 | V | mV |  |  |  | 3.0277 | 0.062894 | 2.7052 | 0.062919
| DCV0.1330 | DCV | 330 | 1000 | 0.1 | V | mV |  |  |  | 50.21 | 0.038072 | 49.941 | 0.037448
| DCV1330 | DCV | 330 | 1000 | 1 | V | mV |  |  |  | 578.9 | 0.0053112 | 578.86 | 0.0048153
| DCV10330 | DCV | 330 | 1000 | 10 | V | mV |  |  |  | 5799.9 | 0.00068362 | 5799.9 | 0.00048322
| DCA0.0000000010 | DCA | 0 | 0.00033 | 1e-09 | A | µA |  |  |  | 0.024 | 180 | 0.023107 | 175.06
| DCA0.000000010 | DCA | 0 | 0.00033 | 1e-08 | A | µA |  |  |  | 0.024 | 173.53 | 0.023814 | 173.54
| DCA0.00000010 | DCA | 0 | 0.00033 | 1e-07 | A | µA |  |  |  | 0.062453 | 112.75 | 0.062384 | 112.54
| DCA0.0000010 | DCA | 0 | 0.00033 | 1e-06 | A | µA |  |  |  | 0.58048 | 15.83 | 0.58045 | 15.637
| DCA0.000010 | DCA | 0 | 0.00033 | 1e-05 | A | µA |  |  |  | 5.8001 | 1.6501 | 5.8001 | 1.5718
| DCA0.00010 | DCA | 0 | 0.00033 | 0.0001 | A | µA |  |  |  | 58 | 0.18852 | 58 | 0.15719
| DCA0.000000010.00033 | DCA | 0.00033 | 0.0033 | 1e-08 | A | µA |  |  |  | 0.068 | 150 | 0.057598 | 152.89
| DCA0.00000010.00033 | DCA | 0.00033 | 0.0033 | 1e-07 | A | µA |  |  |  | 0.076187 | 148.47 | 0.07332 | 149.02
| DCA0.0000010.00033 | DCA | 0.00033 | 0.0033 | 1e-06 | A | µA |  |  |  | 0.56792 | 73.551 | 0.56576 | 73.312
| DCA0.000010.00033 | DCA | 0.00033 | 0.0033 | 1e-05 | A | µA |  |  |  | 5.7989 | 9.0557 | 5.7981 | 8.8116
| DCA0.00010.00033 | DCA | 0.00033 | 0.0033 | 0.0001 | A | µA |  |  |  | 58 | 0.98192 | 58 | 0.88328
| DCA0.00000010.0033 | DCA | 0.0033 | 0.033 | 1e-07 | A | µA |  |  |  | 0.21 | 150 | 0.201 | 145.79
| DCA0.0000010.0033 | DCA | 0.0033 | 0.033 | 1e-06 | A | µA |  |  |  | 0.44201 | 139.58 | 0.43216 | 139.79
| DCA0.000010.0033 | DCA | 0.0033 | 0.033 | 1e-05 | A | µA |  |  |  | 5.6422 | 61.544 | 5.6369 | 61.47
| DCA0.00010.0033 | DCA | 0.0033 | 0.033 | 0.0001 | A | µA |  |  |  | 57.983 | 7.2053 | 57.981 | 7.143
| DCA0.0010.0033 | DCA | 0.0033 | 0.033 | 0.001 | A | µA |  |  |  | 580 | 0.74076 | 580 | 0.71564
| DCA0.0000010.033 | DCA | 0.033 | 0.33 | 1e-06 | A | µA |  |  |  | 3 | 120 | 2.9414 | 115.33
| DCA0.000010.033 | DCA | 0.033 | 0.33 | 1e-05 | A | µA |  |  |  | 5.2853 | 109.48 | 5.266 | 109.51
| DCA0.00010.033 | DCA | 0.033 | 0.33 | 0.0001 | A | µA |  |  |  | 56.995 | 42.585 | 56.984 | 42.551
| DCA0.0010.033 | DCA | 0.033 | 0.33 | 0.001 | A | µA |  |  |  | 579.89 | 4.761 | 579.88 | 4.7411
| DCA0.010.033 | DCA | 0.033 | 0.33 | 0.01 | A | µA |  |  |  | 5800 | 0.4827 | 5800 | 0.47471
| DCA0.000010.33 | DCA | 0.33 | 3 | 1e-05 | A | mA |  |  |  | 0.052 | 0.44 | 0.051 | 0.43873
| DCA0.00010.33 | DCA | 0.33 | 3 | 0.0001 | A | mA |  |  |  | 0.060249 | 0.43614 | 0.060204 | 0.43607
| DCA0.0010.33 | DCA | 0.33 | 3 | 0.001 | A | mA |  |  |  | 0.50421 | 0.32729 | 0.50424 | 0.32696
| DCA0.010.33 | DCA | 0.33 | 3 | 0.01 | A | mA |  |  |  | 5.784 | 0.058643 | 5.7841 | 0.058299
| DCA0.10.33 | DCA | 0.33 | 3 | 0.1 | A | mA |  |  |  | 57.998 | 0.0060523 | 57.998 | 0.0059106
| DCA0.00013 | DCA | 3 | 11 | 0.0001 | A | mA |  |  |  | -0.024 | 0.92 | -0.024619 | 0.91523
| DCA0.0013 | DCA | 3 | 11 | 0.001 | A | mA |  |  |  | 0.055271 | 0.90973 | 0.052344 | 0.90974
| DCA0.013 | DCA | 3 | 11 | 0.01 | A | mA |  |  |  | 4.4629 | 0.64947 | 4.4597 | 0.64887
| DCA0.13 | DCA | 3 | 11 | 0.1 | A | mA |  |  |  | 57.764 | 0.1006 | 57.764 | 0.099909
| DCA0.000111 | DCA | 11 | 20.5 | 0.0001 | A | mA |  |  |  | 2.7 | 1.1 | 2.6473 | 1.1008
| DCA0.00111 | DCA | 11 | 20.5 | 0.001 | A | mA |  |  |  | 2.7 | 1.1006 | 2.6626 | 1.1004
| DCA0.0111 | DCA | 11 | 20.5 | 0.01 | A | mA |  |  |  | 4.2516 | 1.0555 | 4.2551 | 1.0545
| DCA0.111 | DCA | 11 | 20.5 | 0.1 | A | mA |  |  |  | 55.899 | 0.35962 | 55.916 | 0.35746
| DCA 10 turn0.00011 | DCA 10 turn | 1 | 11 | 0.0001 | A | mA |  |  |  | 0.082 | 2 | 0.080875 | 1.9947
| DCA 10 turn0.0011 | DCA 10 turn | 1 | 11 | 0.001 | A | mA |  |  |  | 0.16836 | 1.9875 | 0.16627 | 1.9876
| DCA 10 turn0.011 | DCA 10 turn | 1 | 11 | 0.01 | A | mA |  |  |  | 4.4751 | 1.6631 | 4.4718 | 1.6632
| DCA 10 turn0.11 | DCA 10 turn | 1 | 11 | 0.1 | A | mA |  |  |  | 57.372 | 0.40227 | 57.371 | 0.40202
| DCA 10 turn11 | DCA 10 turn | 1 | 11 | 1 | A | mA |  |  |  | 577.32 | 0.041721 | 577.32 | 0.041613
| DCA 10 turn0.00111 | DCA 10 turn | 11 | 30 | 0.001 | A | mA |  |  |  | 1.6 | 2 | 0.12764 | 2.0471
| DCA 10 turn0.0111 | DCA 10 turn | 11 | 30 | 0.01 | A | mA |  |  |  | 1.6 | 2.0127 | 1.1038 | 2.0235
| DCA 10 turn0.111 | DCA 10 turn | 11 | 30 | 0.1 | A | mA |  |  |  | 49.094 | 1.1796 | 49.066 | 1.1772
| DCA 10 turn111 | DCA 10 turn | 11 | 30 | 1 | A | mA |  |  |  | 576.17 | 0.15058 | 576.16 | 0.14877
| DCA 10 turn1011 | DCA 10 turn | 11 | 30 | 10 | A | mA |  |  |  | 5773.4 | 0.015656 | 5773.4 | 0.014925
| DCA 10 turn10011 | DCA 10 turn | 11 | 30 | 100 | A | mA |  |  |  | 57736 | 0.0017851 | 57735 | 0.0014925
| DCA 10 turn0.0130 | DCA 10 turn | 30 | 110 | 0.01 | A | mA |  |  |  | 4 | 2.1 | 2.509 | 2.1122
| DCA 10 turn0.130 | DCA 10 turn | 30 | 110 | 0.1 | A | mA |  |  |  | 29.551 | 1.93 | 29.511 | 1.9297
| DCA 10 turn130 | DCA 10 turn | 30 | 110 | 1 | A | mA |  |  |  | 565.25 | 0.52854 | 565.25 | 0.52743
| DCA 10 turn1030 | DCA 10 turn | 30 | 110 | 10 | A | mA |  |  |  | 5772.3 | 0.055475 | 5772.3 | 0.054985
| DCA 10 turn10030 | DCA 10 turn | 30 | 110 | 100 | A | mA |  |  |  | 57735 | 0.0056968 | 57735 | 0.005501
| DCA 10 turn0.01110 | DCA 10 turn | 110 | 205 | 0.01 | A | mA |  |  |  | 6.5 | 2.3 | 4.5922 | 2.3086
| DCA 10 turn0.1110 | DCA 10 turn | 110 | 205 | 0.1 | A | mA |  |  |  | 14.301 | 2.2788 | 14.221 | 2.2784
| DCA 10 turn1110 | DCA 10 turn | 110 | 205 | 1 | A | mA |  |  |  | 497.21 | 1.2324 | 497.26 | 1.2301
| DCA 10 turn10110 | DCA 10 turn | 110 | 205 | 10 | A | mA |  |  |  | 5763.1 | 0.14824 | 5763.2 | 0.1469
| DCA 10 turn100110 | DCA 10 turn | 110 | 205 | 100 | A | mA |  |  |  | 57734 | 0.01526 | 57734 | 0.014722
| DCA 20 turn0.00120 | DCA 20 turn | 20 | 60 | 0.001 | A | mA |  |  |  | 3.9 | 2.1 | 0.23928 | 2.0473
| DCA 20 turn0.0120 | DCA 20 turn | 20 | 60 | 0.01 | A | mA |  |  |  | 3.9 | 2.1 | 0.76998 | 2.0407
| DCA 20 turn0.120 | DCA 20 turn | 20 | 60 | 0.1 | A | mA |  |  |  | 38.447 | 1.626 | 38.404 | 1.6257
| DCA 20 turn120 | DCA 20 turn | 20 | 60 | 1 | A | mA |  |  |  | 573.08 | 0.28848 | 573.07 | 0.28764
| DCA 20 turn1020 | DCA 20 turn | 20 | 60 | 10 | A | mA |  |  |  | 5773.1 | 0.029468 | 5773.1 | 0.02912
| DCA 20 turn10020 | DCA 20 turn | 20 | 60 | 100 | A | mA |  |  |  | 57735 | 0.0030514 | 57735 | 0.0029124
| DCA 20 turn0.00160 | DCA 20 turn | 60 | 220 | 0.001 | A | mA |  |  |  | 19 | 2.2 | 4.5221 | 2.114
| DCA 20 turn0.0160 | DCA 20 turn | 60 | 220 | 0.01 | A | mA |  |  |  | 19 | 2.2 | 19.726 | 2.0608
| DCA 20 turn0.160 | DCA 20 turn | 60 | 220 | 0.1 | A | mA |  |  |  | 535.04 | 0.95226 | 535.04 | 0.95071
| DCA 20 turn160 | DCA 20 turn | 60 | 220 | 1 | A | mA |  |  |  | 5768.4 | 0.11068 | 5768.5 | 0.10982
| DCA 20 turn1060 | DCA 20 turn | 60 | 220 | 10 | A | mA |  |  |  | 57735 | 0.011347 | 57735 | 0.011002
| DCA 20 turn10060 | DCA 20 turn | 60 | 220 | 100 | A | mA |  |  |  | 577360 | 0.0012382 | 577350 | 0.0011002
| DCA 20 turn0.01220 | DCA 20 turn | 220 | 410 | 0.01 | A | mA |  |  |  | 13 | 2.3 | 9.0437 | 2.3088
| DCA 20 turn0.1220 | DCA 20 turn | 220 | 410 | 0.1 | A | mA |  |  |  | 13.978 | 2.3014 | 13.925 | 2.3011
| DCA 20 turn1220 | DCA 20 turn | 220 | 410 | 1 | A | mA |  |  |  | 379.54 | 1.7989 | 379.5 | 1.7976
| DCA 20 turn10220 | DCA 20 turn | 220 | 410 | 10 | A | mA |  |  |  | 5732.3 | 0.29332 | 5732.4 | 0.29188
| DCA 20 turn100220 | DCA 20 turn | 220 | 410 | 100 | A | mA |  |  |  | 57731 | 0.03003 | 57731 | 0.029442
| DCA 50 turn0.00155 | DCA 50 turn | 55 | 150 | 0.001 | A | mA |  |  |  | 7.7 | 2 | 0.58265 | 2.0474
| DCA 50 turn0.0155 | DCA 50 turn | 55 | 150 | 0.01 | A | mA |  |  |  | 7.7 | 2.0006 | 0.78384 | 2.0464
| DCA 50 turn0.155 | DCA 50 turn | 55 | 150 | 0.1 | A | mA |  |  |  | 19.451 | 1.9578 | 19.379 | 1.9579
| DCA 50 turn155 | DCA 50 turn | 55 | 150 | 1 | A | mA |  |  |  | 550.23 | 0.69408 | 550.2 | 0.69349
| DCA 50 turn1055 | DCA 50 turn | 55 | 150 | 10 | A | mA |  |  |  | 5770.6 | 0.074862 | 5770.6 | 0.074565
| DCA 50 turn10055 | DCA 50 turn | 55 | 150 | 100 | A | mA |  |  |  | 57735 | 0.0075815 | 57732 | 0.0074625
| DCA 50 turn0.01150 | DCA 50 turn | 150 | 550 | 0.01 | A | mA |  |  |  | 19 | 2.2 | -17.053 | 2.1654
| DCA 50 turn0.1150 | DCA 50 turn | 150 | 550 | 0.1 | A | mA |  |  |  | 19 | 2.2 | -10.261 | 2.1556
| DCA 50 turn1150 | DCA 50 turn | 150 | 550 | 1 | A | mA |  |  |  | 409.49 | 1.6357 | 409.01 | 1.6349
| DCA 50 turn10150 | DCA 50 turn | 150 | 550 | 10 | A | mA |  |  |  | 5740.6 | 0.27618 | 5740.5 | 0.27485
| DCA 50 turn100150 | DCA 50 turn | 150 | 550 | 100 | A | mA |  |  |  | 57732 | 0.028329 | 57732 | 0.027783
| DCA 50 turn0.01550 | DCA 50 turn | 550 | 1025 | 0.01 | A | mA |  |  |  | 32 | 2.3 | 22.487 | 2.3089
| DCA 50 turn0.1550 | DCA 50 turn | 550 | 1025 | 0.1 | A | mA |  |  |  | 32 | 2.3014 | 24.491 | 2.3076
| DCA 50 turn1550 | DCA 50 turn | 550 | 1025 | 1 | A | mA |  |  |  | 209.6 | 2.1949 | 208.5 | 2.1945
| DCA 50 turn10550 | DCA 50 turn | 550 | 1025 | 10 | A | mA |  |  |  | 5532.1 | 0.70111 | 5532.1 | 0.69875
| DCA 50 turn100550 | DCA 50 turn | 550 | 1025 | 100 | A | mA |  |  |  | 57709 | 0.074654 | 57710 | 0.073571
| Resistance0.0000010 | Resistance | 0 | 0.011 | 1e-06 | kOhms | mΩ |  |  |  | 12 | 140 | 11.576 | 138.3
| Resistance0.000010 | Resistance | 0 | 0.011 | 1e-05 | kOhms | mΩ |  |  |  | 12.939 | 125.81 | 12.935 | 125.25
| Resistance0.00010 | Resistance | 0 | 0.011 | 0.0001 | kOhms | mΩ |  |  |  | 59.144 | 29.356 | 59.141 | 28.788
| Resistance0.0010 | Resistance | 0 | 0.011 | 0.001 | kOhms | mΩ |  |  |  | 580.12 | 3.176 | 580.12 | 2.9427
| Resistance0.0000010.011 | Resistance | 0.011 | 0.033 | 1e-06 | kOhms | mΩ |  |  |  | 18 | 140 | 17.339 | 138.85
| Resistance0.000010.011 | Resistance | 0.011 | 0.033 | 1e-05 | kOhms | mΩ |  |  |  | 18.263 | 133.72 | 18.258 | 133.57
| Resistance0.00010.011 | Resistance | 0.011 | 0.033 | 0.0001 | kOhms | mΩ |  |  |  | 60.487 | 46.304 | 60.481 | 46.04
| Resistance0.0010.011 | Resistance | 0.011 | 0.033 | 0.001 | kOhms | mΩ |  |  |  | 580.26 | 4.9938 | 580.25 | 4.8783
| Resistance0.0000010.033 | Resistance | 0.033 | 0.11 | 1e-06 | kOhms | mΩ |  |  |  | 18 | 100 | 17.342 | 104.06
| Resistance0.000010.033 | Resistance | 0.033 | 0.11 | 1e-05 | kOhms | mΩ |  |  |  | 18.224 | 101.51 | 18.221 | 101.28
| Resistance0.00010.033 | Resistance | 0.033 | 0.11 | 0.0001 | kOhms | mΩ |  |  |  | 60.259 | 41.284 | 60.259 | 40.82
| Resistance0.0010.033 | Resistance | 0.033 | 0.11 | 0.001 | kOhms | mΩ |  |  |  | 580.23 | 4.6533 | 580.23 | 4.4422
| Resistance0.0000010.11 | Resistance | 0.11 | 0.33 | 1e-06 | kOhms | mΩ |  |  |  | 25 | 300 | 24.812 | 296.87
| Resistance0.000010.11 | Resistance | 0.11 | 0.33 | 1e-05 | kOhms | mΩ |  |  |  | 25.358 | 294.79 | 25.35 | 294.41
| Resistance0.00010.11 | Resistance | 0.11 | 0.33 | 0.0001 | kOhms | mΩ |  |  |  | 61.466 | 184.35 | 61.46 | 183.05
| Resistance0.0010.11 | Resistance | 0.11 | 0.33 | 0.001 | kOhms | mΩ |  |  |  | 580.25 | 24.237 | 580.25 | 23.449
| Resistance0.000010.33 | Resistance | 0.33 | 1.1 | 1e-05 | kOhms | mΩ |  |  |  | 28 | 100 | 23.818 | 103.66
| Resistance0.00010.33 | Resistance | 0.33 | 1.1 | 0.0001 | kOhms | mΩ |  |  |  | 52.968 | 87.892 | 52.867 | 87.789
| Resistance0.0010.33 | Resistance | 0.33 | 1.1 | 0.001 | kOhms | mΩ |  |  |  | 577.2 | 17.401 | 577.18 | 17.221
| Resistance0.010.33 | Resistance | 0.33 | 1.1 | 0.01 | kOhms | mΩ |  |  |  | 5799.7 | 1.8248 | 5799.7 | 1.75
| Resistance0.000011.1 | Resistance | 1.1 | 3.3 | 1e-05 | kOhms | Ω |  |  |  | 0.25 | 0.1 | 0.23096 | 0.10411
| Resistance0.00011.1 | Resistance | 1.1 | 3.3 | 0.0001 | kOhms | Ω |  |  |  | 0.25 | 0.10329 | 0.23666 | 0.10325
| Resistance0.0011.1 | Resistance | 1.1 | 3.3 | 0.001 | kOhms | Ω |  |  |  | 0.60487 | 0.064429 | 0.60441 | 0.064221
| Resistance0.011.1 | Resistance | 1.1 | 3.3 | 0.01 | kOhms | Ω |  |  |  | 5.8014 | 0.0083761 | 5.8012 | 0.0082294
| Resistance0.00013.3 | Resistance | 3.3 | 11 | 0.0001 | kOhms | Ω |  |  |  | 0.17 | 0.1 | 0.12474 | 0.10353
| Resistance0.0013.3 | Resistance | 3.3 | 11 | 0.001 | kOhms | Ω |  |  |  | 0.46536 | 0.084151 | 0.46464 | 0.084044
| Resistance0.013.3 | Resistance | 3.3 | 11 | 0.01 | kOhms | Ω |  |  |  | 5.7682 | 0.015398 | 5.7682 | 0.015238
| Resistance0.13.3 | Resistance | 3.3 | 11 | 0.1 | kOhms | Ω |  |  |  | 57.997 | 0.0016098 | 57.997 | 0.0015438
| Resistance0.000111 | Resistance | 11 | 33 | 0.0001 | kOhms | Ω |  |  |  | 1.3 | 0.1 | 1.1561 | 0.10412
| Resistance0.00111 | Resistance | 11 | 33 | 0.001 | kOhms | Ω |  |  |  | 1.3 | 0.10253 | 1.2442 | 0.10252
| Resistance0.0111 | Resistance | 11 | 33 | 0.01 | kOhms | Ω |  |  |  | 5.6614 | 0.052695 | 5.6605 | 0.052621
| Resistance0.111 | Resistance | 11 | 33 | 0.1 | kOhms | Ω |  |  |  | 57.978 | 0.0062175 | 57.978 | 0.0061738
| Resistance0.00133 | Resistance | 33 | 110 | 0.001 | kOhms | Ω |  |  |  | 1.3 | 0.13 | 1.2347 | 0.12665
| Resistance0.0133 | Resistance | 33 | 110 | 0.01 | kOhms | Ω |  |  |  | 4.353 | 0.10805 | 4.3501 | 0.10797
| Resistance0.133 | Resistance | 33 | 110 | 0.1 | kOhms | Ω |  |  |  | 57.521 | 0.022173 | 57.522 | 0.022051
| Resistance133 | Resistance | 33 | 110 | 1 | kOhms | Ω |  |  |  | 579.95 | 0.0022971 | 579.95 | 0.0022465
| Resistance0.001110 | Resistance | 110 | 330 | 0.001 | kOhms | Ω |  |  |  | 12 | 0.14 | 11.565 | 0.1393
| Resistance0.01110 | Resistance | 110 | 330 | 0.01 | kOhms | Ω |  |  |  | 12.289 | 0.13784 | 12.284 | 0.13782
| Resistance0.1110 | Resistance | 110 | 330 | 0.1 | kOhms | Ω |  |  |  | 55.026 | 0.080748 | 55.024 | 0.080667
| Resistance1110 | Resistance | 110 | 330 | 1 | kOhms | Ω |  |  |  | 579.51 | 0.010142 | 579.51 | 0.010091
| Resistance0.01330 | Resistance | 330 | 1100 | 0.01 | kOhms | Ω |  |  |  | 16 | 0.17 | 12.341 | 0.17274
| Resistance0.1330 | Resistance | 330 | 1100 | 0.1 | kOhms | Ω |  |  |  | 38.744 | 0.15613 | 38.702 | 0.15611
| Resistance1330 | Resistance | 330 | 1100 | 1 | kOhms | Ω |  |  |  | 571.18 | 0.039241 | 571.18 | 0.039165
| Resistance10330 | Resistance | 330 | 1100 | 10 | kOhms | Ω |  |  |  | 5799.1 | 0.0040776 | 5799.1 | 0.0040446
| Resistance0.011100 | Resistance | 1100 | 3300 | 0.01 | kOhms | kΩ |  |  |  | 0.19 | 0.00017 | 0.17339 | 0.00017318
| Resistance0.11100 | Resistance | 1100 | 3300 | 0.1 | kOhms | kΩ |  |  |  | 0.19 | 0.00017215 | 0.1791 | 0.00017213
| Resistance11100 | Resistance | 1100 | 3300 | 1 | kOhms | kΩ |  |  |  | 0.55516 | 0.00011804 | 0.55501 | 0.00011789
| Resistance101100 | Resistance | 1100 | 3300 | 10 | kOhms | kΩ |  |  |  | 5.7933 | 1.659e-05 | 5.7933 | 1.6469e-05
| Resistance0.13300 | Resistance | 3300 | 11000 | 0.1 | kOhms | kΩ |  |  |  | 0.33 | 0.00069 | 0.29033 | 0.00069272
| Resistance13300 | Resistance | 3300 | 11000 | 1 | kOhms | kΩ |  |  |  | 0.37289 | 0.00068715 | 0.37252 | 0.00068716
| Resistance103300 | Resistance | 3300 | 11000 | 10 | kOhms | kΩ |  |  |  | 4.8626 | 0.00044972 | 4.8623 | 0.00044966
| Resistance1003300 | Resistance | 3300 | 11000 | 100 | kOhms | kΩ |  |  |  | 57.852 | 6.2356e-05 | 57.852 | 6.2304e-05
| Resistance0.111000 | Resistance | 11000 | 33000 | 0.1 | kOhms | kΩ |  |  |  | 2.9 | 0.0012 | 2.8873 | 0.0011548
| Resistance111000 | Resistance | 11000 | 33000 | 1 | kOhms | kΩ |  |  |  | 2.9004 | 0.0011545 | 2.9 | 0.0011545
| Resistance1011000 | Resistance | 11000 | 33000 | 10 | kOhms | kΩ |  |  |  | 4.2484 | 0.0011259 | 4.2478 | 0.0011259
| Resistance10011000 | Resistance | 11000 | 33000 | 100 | kOhms | kΩ |  |  |  | 54.575 | 0.00049852 | 54.575 | 0.00049846
| Resistance133000 | Resistance | 33000 | 110000 | 1 | kOhms | kΩ |  |  |  | 3.5 | 0.0058 | 3.4663 | 0.0057735
| Resistance1033000 | Resistance | 33000 | 110000 | 10 | kOhms | kΩ |  |  |  | 3.5803 | 0.0057727 | 3.5792 | 0.0057727
| Resistance10033000 | Resistance | 33000 | 110000 | 100 | kOhms | kΩ |  |  |  | 14.463 | 0.0056975 | 14.465 | 0.0056974
| Resistance100033000 | Resistance | 33000 | 110000 | 1000 | kOhms | kΩ |  |  |  | 503.99 | 0.0032606 | 503.99 | 0.0032605
| Resistance1110000 | Resistance | 110000 | 330000 | 1 | kOhms | kΩ |  |  |  | 120 | 0.0058 | 115.47 | 0.0057735
| Resistance10110000 | Resistance | 110000 | 330000 | 10 | kOhms | kΩ |  |  |  | 120 | 0.0057734 | 115.52 | 0.0057734
| Resistance100110000 | Resistance | 110000 | 330000 | 100 | kOhms | kΩ |  |  |  | 120 | 0.0057671 | 118.42 | 0.0057671
| Resistance1000110000 | Resistance | 110000 | 330000 | 1000 | kOhms | kΩ |  |  |  | 371.66 | 0.0052445 | 371.67 | 0.0052444
| Resistance10330000 | Resistance | 330000 | 1100000 | 10 | kOhms | kΩ |  |  |  | 930 | 0.017 | 577.19 | 0.017321
| Resistance100330000 | Resistance | 330000 | 1100000 | 100 | kOhms | kΩ |  |  |  | 930 | 0.01732 | 578 | 0.01732
| Resistance1000330000 | Resistance | 330000 | 1100000 | 1000 | kOhms | kΩ |  |  |  | 930 | 0.017297 | 611.79 | 0.017297
| Resistance10000330000 | Resistance | 330000 | 1100000 | 10000 | kOhms | kΩ |  |  |  | 3454.2 | 0.015468 | 3454.1 | 0.015468
| ACV0.0000010.0010.01 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 0.01 | 0.045 | Hz | 25 | 1700 | 23.181 | 1732
| ACV0.000010.0010.01 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 0.01 | 0.045 | Hz | 25 | 1718.2 | 23.854 | 1717.9
| ACV0.00010.0010.01 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 0.01 | 0.045 | Hz | 62.01 | 1125.5 | 61.999 | 1123.8
| ACV0.0010.0010.01 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 0.01 | 0.045 | Hz | 580.38 | 157.54 | 580.38 | 156.33
| ACV0.010.0010.01 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 0.01 | 0.045 | Hz | 5800 | 16.202 | 5800 | 15.714
| ACV0.0000010.0010.045 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 1200 | 23.185 | 1155.4
| ACV0.000010.0010.045 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 1143.4 | 23.87 | 1142.8
| ACV0.00010.0010.045 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 0.045 | 10 | Hz | 62.237 | 673.84 | 62.225 | 671.86
| ACV0.0010.0010.045 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 0.045 | 10 | Hz | 580.43 | 86.247 | 580.43 | 85.036
| ACV0.010.0010.045 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800.1 | 9.0181 | 5800.1 | 8.5307
| ACV0.0000010.00110 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 10 | 20 | kHz | 25 | 1700 | 23.181 | 1732
| ACV0.000010.00110 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 10 | 20 | kHz | 25 | 1718.2 | 23.854 | 1717.9
| ACV0.00010.00110 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 10 | 20 | kHz | 62.01 | 1125.5 | 61.999 | 1123.8
| ACV0.0010.00110 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 10 | 20 | kHz | 580.38 | 157.54 | 580.38 | 156.33
| ACV0.010.00110 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 10 | 20 | kHz | 5800 | 16.202 | 5800 | 15.714
| ACV0.0000010.00120 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 20 | 50 | kHz | 24 | 2300 | 23.565 | 2300.9
| ACV0.000010.00120 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 20 | 50 | kHz | 24.224 | 2286.5 | 24.216 | 2286.3
| ACV0.00010.00120 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 20 | 50 | kHz | 61.904 | 1615.7 | 61.89 | 1614.4
| ACV0.0010.00120 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 20 | 50 | kHz | 580.34 | 247.93 | 580.33 | 246.73
| ACV0.010.00120 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 20 | 50 | kHz | 5800 | 25.351 | 5800 | 24.864
| ACV0.0000010.00150 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 50 | 100 | kHz | 40 | 4000 | 38.937 | 4027
| ACV0.000010.00150 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 50 | 100 | kHz | 40 | 4017.5 | 39.332 | 4018
| ACV0.00010.00150 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 50 | 100 | kHz | 68.858 | 3411.4 | 68.766 | 3411.7
| ACV0.0010.00150 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 50 | 100 | kHz | 580.9 | 730.69 | 580.86 | 729
| ACV0.010.00150 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 50 | 100 | kHz | 5800.1 | 75.275 | 5800.1 | 74.549
| ACV0.0000010.001100 | ACV | 0.001 | 0.033 | 1e-06 | V | µV | 100 | 500 | kHz | 89 | 11000 | 72.86 | 11473
| ACV0.000010.001100 | ACV | 0.001 | 0.033 | 1e-05 | V | µV | 100 | 500 | kHz | 89 | 11468 | 73.063 | 11468
| ACV0.00010.001100 | ACV | 0.001 | 0.033 | 0.0001 | V | µV | 100 | 500 | kHz | 91.464 | 11024 | 91.322 | 11026
| ACV0.0010.001100 | ACV | 0.001 | 0.033 | 0.001 | V | µV | 100 | 500 | kHz | 581.54 | 4655.9 | 581.45 | 4653.2
| ACV0.010.001100 | ACV | 0.001 | 0.033 | 0.01 | V | µV | 100 | 500 | kHz | 5800.1 | 530.77 | 5800.1 | 529.12
| ACV0.0000010.0330.01 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 0.01 | 0.045 | Hz | 33 | 550 | 23.236 | 577.95
| ACV0.000010.0330.01 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 0.01 | 0.045 | Hz | 33 | 550.18 | 23.663 | 576.89
| ACV0.00010.0330.01 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 0.01 | 0.045 | Hz | 55.234 | 504.47 | 55.135 | 504.68
| ACV0.0010.0330.01 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 0.01 | 0.045 | Hz | 577.51 | 123.44 | 577.47 | 123.45
| ACV0.010.0330.01 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 0.01 | 0.045 | Hz | 5799.8 | 12.763 | 577.47 | 12.763
| ACV0.0000010.0330.045 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 0.045 | 10 | Hz | 24 | 350 | 23.254 | 347.52
| ACV0.000010.0330.045 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 0.045 | 10 | Hz | 24 | 346.2 | 23.771 | 346.32
| ACV0.00010.0330.045 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 0.045 | 10 | Hz | 58.586 | 276.03 | 58.483 | 276.21
| ACV0.0010.0330.045 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 0.045 | 10 | Hz | 579.4 | 50.965 | 579.36 | 50.969
| ACV0.010.0330.045 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 0.045 | 10 | Hz | 5800 | 5.1718 | 5799.9 | 5.1718
| ACV0.0000010.03310 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 10 | 20 | kHz | 24 | 810 | 23.218 | 808.67
| ACV0.000010.03310 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 10 | 20 | kHz | 24 | 807.65 | 23.583 | 807.74
| ACV0.00010.03310 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 10 | 20 | kHz | 52.238 | 738.2 | 52.145 | 738.41
| ACV0.0010.03310 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 10 | 20 | kHz | 574.82 | 223.39 | 574.77 | 223.41
| ACV0.010.03310 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 10 | 20 | kHz | 5799.5 | 23.686 | 5799.4 | 23.686
| ACV0.0000010.03320 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 20 | 50 | kHz | 47 | 1200 | 46.389 | 1155.4
| ACV0.000010.03320 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 20 | 50 | kHz | 47 | 1154.8 | 46.618 | 1154.8
| ACV0.00010.03320 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 20 | 50 | kHz | 66.007 | 1107.9 | 65.939 | 1108
| ACV0.0010.03320 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 20 | 50 | kHz | 571.22 | 452.98 | 571.18 | 452.85
| ACV0.010.03320 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 20 | 50 | kHz | 5799 | 51.02 | 5798.9 | 50.942
| ACV0.0000010.03350 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 50 | 100 | kHz | 200 | 2700 | 196.5 | 2656.4
| ACV0.000010.03350 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 50 | 100 | kHz | 200 | 2656.2 | 196.56 | 2656.3
| ACV0.00010.03350 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 50 | 100 | kHz | 202.89 | 2641.9 | 202.83 | 2642
| ACV0.0010.03350 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 50 | 100 | kHz | 582.19 | 1932.6 | 582.1 | 1932.5
| ACV0.010.03350 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 50 | 100 | kHz | 5796.8 | 308.19 | 5796.8 | 308.02
| ACV0.0000010.033100 | ACV | 0.033 | 0.33 | 1e-06 | V | µV | 100 | 500 | kHz | 390 | 5800 | 381.75 | 5779.4
| ACV0.000010.033100 | ACV | 0.033 | 0.33 | 1e-05 | V | µV | 100 | 500 | kHz | 390 | 5779.3 | 381.8 | 5779.3
| ACV0.00010.033100 | ACV | 0.033 | 0.33 | 0.0001 | V | µV | 100 | 500 | kHz | 390 | 5771.9 | 384.93 | 5772
| ACV0.0010.033100 | ACV | 0.033 | 0.33 | 0.001 | V | µV | 100 | 500 | kHz | 643.31 | 5206.4 | 643.12 | 5206.6
| ACV0.010.033100 | ACV | 0.033 | 0.33 | 0.01 | V | µV | 100 | 500 | kHz | 5783 | 1371.2 | 5783 | 1370.9
| ACV0.000010.330.01 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 0.01 | 0.045 | Hz | 0.11 | 0.56 | 0.10391 | 0.56221
| ACV0.00010.330.01 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 0.01 | 0.045 | Hz | 0.11 | 0.56033 | 0.10921 | 0.56037
| ACV0.0010.330.01 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 0.01 | 0.045 | Hz | 0.48274 | 0.46448 | 0.4824 | 0.4645
| ACV0.010.330.01 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 0.01 | 0.045 | Hz | 5.7719 | 0.10338 | 5.7717 | 0.10332
| ACV0.10.330.01 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 0.01 | 0.045 | Hz | 57.997 | 0.010643 | 57.997 | 0.010612
| ACV0.000010.330.045 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 0.045 | 10 | Hz | 0.07 | 0.35 | 0.069477 | 0.34756
| ACV0.00010.330.045 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.079443 | 0.34499 | 0.079136 | 0.34505
| ACV0.0010.330.045 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.52675 | 0.24887 | 0.5264 | 0.24885
| ACV0.010.330.045 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.7894 | 0.041576 | 5.7892 | 0.041503
| ACV0.10.330.045 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.999 | 0.0042263 | 57.999 | 0.004196
| ACV0.000010.3310 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 10 | 20 | kHz | 0.07 | 0.81 | 0.069435 | 0.80878
| ACV0.00010.3310 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 10 | 20 | kHz | 0.075008 | 0.80729 | 0.074829 | 0.80733
| ACV0.0010.3310 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 10 | 20 | kHz | 0.43427 | 0.71669 | 0.43391 | 0.71674
| ACV0.010.3310 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 10 | 20 | kHz | 5.7428 | 0.2035 | 5.7426 | 0.20344
| ACV0.10.3310 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 10 | 20 | kHz | 57.994 | 0.021456 | 57.994 | 0.021426
| ACV0.000010.3320 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 20 | 50 | kHz | 0.07 | 1.2 | 0.069506 | 1.1559
| ACV0.00010.3320 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 20 | 50 | kHz | 0.073711 | 1.1548 | 0.073371 | 1.1549
| ACV0.0010.3320 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 20 | 50 | kHz | 0.38065 | 1.0748 | 0.37987 | 1.0749
| ACV0.010.3320 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 20 | 50 | kHz | 5.6886 | 0.39166 | 5.6883 | 0.39154
| ACV0.10.3320 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 20 | 50 | kHz | 57.988 | 0.043219 | 57.988 | 0.043148
| ACV0.000010.3350 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 50 | 100 | kHz | 0.24 | 2.7 | 0.23108 | 2.6571
| ACV0.00010.3350 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 50 | 100 | kHz | 0.24 | 2.6566 | 0.23256 | 2.6567
| ACV0.0010.3350 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 50 | 100 | kHz | 0.38846 | 2.6152 | 0.38733 | 2.6154
| ACV0.010.3350 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 50 | 100 | kHz | 5.3723 | 1.6168 | 5.3715 | 1.6167
| ACV0.10.3350 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 50 | 100 | kHz | 57.935 | 0.23029 | 57.935 | 0.23012
| ACV0.000010.33100 | ACV | 0.33 | 3.3 | 1e-05 | V | mV | 100 | 500 | kHz | 1.1 | 5.8 | 1.0377 | 5.7942
| ACV0.00010.33100 | ACV | 0.33 | 3.3 | 0.0001 | V | mV | 100 | 500 | kHz | 1.1 | 5.794 | 1.038 | 5.7941
| ACV0.0010.33100 | ACV | 0.33 | 3.3 | 0.001 | V | mV | 100 | 500 | kHz | 1.1 | 5.7778 | 1.0964 | 5.778
| ACV0.010.33100 | ACV | 0.33 | 3.3 | 0.01 | V | mV | 100 | 500 | kHz | 4.9001 | 4.8711 | 4.8981 | 4.8713
| ACV0.10.33100 | ACV | 0.33 | 3.3 | 0.1 | V | mV | 100 | 500 | kHz | 57.706 | 1.1207 | 57.705 | 1.1204
| ACV0.00013.30.01 | ACV | 3.3 | 33 | 0.0001 | V | mV | 0.01 | 0.045 | Hz | 0.93 | 0.58 | 0.92292 | 0.5783
| ACV0.0013.30.01 | ACV | 3.3 | 33 | 0.001 | V | mV | 0.01 | 0.045 | Hz | 0.98726 | 0.57661 | 0.98668 | 0.57662
| ACV0.013.30.01 | ACV | 3.3 | 33 | 0.01 | V | mV | 0.01 | 0.045 | Hz | 4.8574 | 0.48407 | 4.8565 | 0.48406
| ACV0.13.30.01 | ACV | 3.3 | 33 | 0.1 | V | mV | 0.01 | 0.045 | Hz | 57.704 | 0.11064 | 57.704 | 0.11059
| ACV13.30.01 | ACV | 3.3 | 33 | 1 | V | mV | 0.01 | 0.045 | Hz | 579.97 | 0.011401 | 579.97 | 0.011382
| ACV0.00013.30.045 | ACV | 3.3 | 33 | 0.0001 | V | mV | 0.045 | 10 | Hz | 0.7 | 0.35 | 0.69218 | 0.34795
| ACV0.0013.30.045 | ACV | 3.3 | 33 | 0.001 | V | mV | 0.045 | 10 | Hz | 0.78957 | 0.34543 | 0.78891 | 0.34543
| ACV0.013.30.045 | ACV | 3.3 | 33 | 0.01 | V | mV | 0.045 | 10 | Hz | 5.2633 | 0.24921 | 5.2624 | 0.24918
| ACV0.13.30.045 | ACV | 3.3 | 33 | 0.1 | V | mV | 0.045 | 10 | Hz | 57.892 | 0.04162 | 57.892 | 0.041575
| ACV13.30.045 | ACV | 3.3 | 33 | 1 | V | mV | 0.045 | 10 | Hz | 579.99 | 0.0042219 | 579.99 | 0.0042034
| ACV0.00013.310 | ACV | 3.3 | 33 | 0.0001 | V | mV | 10 | 20 | kHz | 0.7 | 0.81 | 0.69304 | 0.80896
| ACV0.0013.310 | ACV | 3.3 | 33 | 0.001 | V | mV | 10 | 20 | kHz | 0.74744 | 0.8075 | 0.74699 | 0.80751
| ACV0.013.310 | ACV | 3.3 | 33 | 0.01 | V | mV | 10 | 20 | kHz | 4.3392 | 0.7169 | 4.3383 | 0.7169
| ACV0.13.310 | ACV | 3.3 | 33 | 0.1 | V | mV | 10 | 20 | kHz | 57.426 | 0.20355 | 57.426 | 0.20351
| ACV13.310 | ACV | 3.3 | 33 | 1 | V | mV | 10 | 20 | kHz | 579.94 | 0.021452 | 579.94 | 0.021433
| ACV0.00013.320 | ACV | 3.3 | 33 | 0.0001 | V | mV | 20 | 50 | kHz | 0.7 | 1.2 | 0.69122 | 1.1565
| ACV0.0013.320 | ACV | 3.3 | 33 | 0.001 | V | mV | 20 | 50 | kHz | 0.73466 | 1.1553 | 0.73165 | 1.1554
| ACV0.013.320 | ACV | 3.3 | 33 | 0.01 | V | mV | 20 | 50 | kHz | 3.8002 | 1.0754 | 3.7971 | 1.0754
| ACV0.13.320 | ACV | 3.3 | 33 | 0.1 | V | mV | 20 | 50 | kHz | 56.883 | 0.3919 | 56.882 | 0.39182
| ACV13.320 | ACV | 3.3 | 33 | 1 | V | mV | 20 | 50 | kHz | 579.88 | 0.043223 | 579.88 | 0.043182
| ACV0.00013.350 | ACV | 3.3 | 33 | 0.0001 | V | mV | 50 | 100 | kHz | 2.3 | 2.7 | 2.3079 | 2.6577
| ACV0.0013.350 | ACV | 3.3 | 33 | 0.001 | V | mV | 50 | 100 | kHz | 2.3265 | 2.6572 | 2.3234 | 2.6573
| ACV0.013.350 | ACV | 3.3 | 33 | 0.01 | V | mV | 50 | 100 | kHz | 3.8797 | 2.6158 | 3.8711 | 2.616
| ACV0.13.350 | ACV | 3.3 | 33 | 0.1 | V | mV | 50 | 100 | kHz | 53.719 | 1.6172 | 53.713 | 1.6172
| ACV13.350 | ACV | 3.3 | 33 | 1 | V | mV | 50 | 100 | kHz | 579.35 | 0.23028 | 579.35 | 0.23021
| ACV0.00013.3100 | ACV | 3.3 | 33 | 0.0001 | V | mV | 100 | 500 | kHz* | 13 | 4.6 | 2.3364 | 2.6995
| ACV0.0013.3100 | ACV | 3.3 | 33 | 0.001 | V | mV | 100 | 500 | kHz* | 13 | 4.6 | 2.3501 | 2.699
| ACV0.013.3100 | ACV | 3.3 | 33 | 0.01 | V | mV | 100 | 500 | kHz* | 13 | 4.6 | 3.831 | 2.6581
| ACV0.13.3100 | ACV | 3.3 | 33 | 0.1 | V | mV | 100 | 500 | kHz* | 47.329 | 3.5597 | 53.608 | 1.6549
| ACV13.3100 | ACV | 3.3 | 33 | 1 | V | mV | 100 | 500 | kHz* | 579.33 | 0.23723 | 579.33 | 0.23717
| ACV0.0013310 | ACV | 33 | 330 | 0.001 | V | mV | 10 | 45 | Hz* | 3.5 | 0.58 | 3.4705 | 0.57868
| ACV0.013310 | ACV | 33 | 330 | 0.01 | V | mV | 10 | 45 | Hz* | 4.2872 | 0.57647 | 4.2606 | 0.57652
| ACV0.13310 | ACV | 33 | 330 | 0.1 | V | mV | 10 | 45 | Hz* | 46.641 | 0.47359 | 46.602 | 0.47361
| ACV13310 | ACV | 33 | 330 | 1 | V | mV | 10 | 45 | Hz* | 576.98 | 0.10542 | 576.96 | 0.10533
| ACV103310 | ACV | 33 | 330 | 10 | V | mV | 10 | 45 | Hz* | 5799.7 | 0.010862 | 5799.7 | 0.010822
| ACV0.001330.045 | ACV | 33 | 330 | 0.001 | V | mV | 0.045 | 10 | Hz | 12 | 0.92 | 10.393 | 0.92459
| ACV0.01330.045 | ACV | 33 | 330 | 0.01 | V | mV | 0.045 | 10 | Hz | 12 | 0.92331 | 10.834 | 0.9234
| ACV0.1330.045 | ACV | 33 | 330 | 0.1 | V | mV | 0.045 | 10 | Hz | 43.309 | 0.841 | 43.21 | 0.84115
| ACV1330.045 | ACV | 33 | 330 | 1 | V | mV | 0.045 | 10 | Hz | 572.72 | 0.26556 | 572.68 | 0.26538
| ACV10330.045 | ACV | 33 | 330 | 10 | V | mV | 0.045 | 10 | Hz | 5799.2 | 0.028474 | 5799.2 | 0.028386
| ACV0.0013310 | ACV | 33 | 330 | 0.001 | V | mV | 10 | 20 | kHz | 24 | 1 | 10.388 | 1.04
| ACV0.013310 | ACV | 33 | 330 | 0.01 | V | mV | 10 | 20 | kHz | 24 | 1.0332 | 10.796 | 1.0389
| ACV0.13310 | ACV | 33 | 330 | 0.1 | V | mV | 10 | 20 | kHz | 41.651 | 0.95969 | 41.554 | 0.95985
| ACV13310 | ACV | 33 | 330 | 1 | V | mV | 10 | 20 | kHz | 570.92 | 0.32864 | 570.88 | 0.32848
| ACV103310 | ACV | 33 | 330 | 10 | V | mV | 10 | 20 | kHz | 5799 | 0.035762 | 5799 | 0.035674
| ACV0.0013320 | ACV | 33 | 330 | 0.001 | V | mV | 20 | 50 | kHz* | 22 | 1.4 | 10.266 | 1.3912
| ACV0.013320 | ACV | 33 | 330 | 0.01 | V | mV | 20 | 50 | kHz* | 22 | 1.4 | 10.577 | 1.3904
| ACV0.13320 | ACV | 33 | 330 | 0.1 | V | mV | 20 | 50 | kHz* | 36.02 | 1.3575 | 37.167 | 1.3205
| ACV13320 | ACV | 33 | 330 | 1 | V | mV | 20 | 50 | kHz* | 564.6 | 0.55034 | 564.56 | 0.55021
| ACV103320 | ACV | 33 | 330 | 10 | V | mV | 20 | 50 | kHz* | 5798.2 | 0.063012 | 5798.2 | 0.062925
| ACV0.0013350 | ACV | 33 | 330 | 0.001 | V | mV | 50 | 100 | kHz* | 92 | 2.8 | 91.937 | 2.7856
| ACV0.013350 | ACV | 33 | 330 | 0.01 | V | mV | 50 | 100 | kHz* | 92.063 | 2.7853 | 92.033 | 2.7854
| ACV0.13350 | ACV | 33 | 330 | 0.1 | V | mV | 50 | 100 | kHz* | 101.77 | 2.761 | 101.69 | 2.7611
| ACV13350 | ACV | 33 | 330 | 1 | V | mV | 50 | 100 | kHz* | 546.63 | 1.8765 | 546.54 | 1.8763
| ACV103350 | ACV | 33 | 330 | 10 | V | mV | 50 | 100 | kHz* | 5793.6 | 0.28499 | 5793.5 | 0.28476
| ACV0.0013300.045 | ACV | 330 | 1020 | 0.001 | V | mV | 0.045 | 1 | Hz | 24 | 0.58 | 23.335 | 0.57834
| ACV0.013300.045 | ACV | 330 | 1020 | 0.01 | V | mV | 0.045 | 1 | Hz | 33.354 | 0.57134 | 33.321 | 0.57119
| ACV0.13300.045 | ACV | 330 | 1020 | 0.1 | V | mV | 0.045 | 1 | Hz | 510.18 | 0.32785 | 510.26 | 0.32724
| ACV13300.045 | ACV | 330 | 1020 | 1 | V | mV | 0.045 | 1 | Hz | 5790.3 | 0.041506 | 5790.4 | 0.041123
| ACV103300.045 | ACV | 330 | 1020 | 10 | V | mV | 0.045 | 1 | Hz | 57999 | 0.0042793 | 57999 | 0.004125
| ACV0.0013301 | ACV | 330 | 1020 | 0.001 | V | mV | 1 | 5 | kHz | 28 | 0.92 | 23.262 | 0.92437
| ACV0.013301 | ACV | 330 | 1020 | 0.01 | V | mV | 1 | 5 | kHz | 29.895 | 0.91967 | 29.873 | 0.91957
| ACV0.13301 | ACV | 330 | 1020 | 0.1 | V | mV | 1 | 5 | kHz | 446.23 | 0.66766 | 446.26 | 0.66723
| ACV13301 | ACV | 330 | 1020 | 1 | V | mV | 1 | 5 | kHz | 5775.4 | 0.10274 | 5775.5 | 0.10236
| ACV103301 | ACV | 330 | 1020 | 10 | V | mV | 1 | 5 | kHz | 57998 | 0.010468 | 57998 | 0.010314
| ACV0.0013305 | ACV | 330 | 1020 | 0.001 | V | mV | 5 | 10 | kHz | 64 | 1 | 23.23 | 1.0398
| ACV0.013305 | ACV | 330 | 1020 | 0.01 | V | mV | 5 | 10 | kHz | 64 | 1.0355 | 29.165 | 1.0355
| ACV0.13305 | ACV | 330 | 1020 | 0.1 | V | mV | 5 | 10 | kHz | 426.17 | 0.78768 | 426.19 | 0.7873
| ACV13305 | ACV | 330 | 1020 | 1 | V | mV | 5 | 10 | kHz | 5769 | 0.12912 | 5769.1 | 0.12875
| ACV103305 | ACV | 330 | 1020 | 10 | V | mV | 5 | 10 | kHz | 57997 | 0.013152 | 57997 | 0.012998
| Frequency0.000010.001 | Frequency | 0.001 | 0.12 | 1e-05 | kHz | Hz |  |  |  | 0.006 | 0.013 | 0.005881 | 0.012606
| Frequency0.00010.001 | Frequency | 0.001 | 0.12 | 0.0001 | kHz | Hz |  |  |  | 0.057764 | 0.0019154 | 0.057746 | 0.0014481
| Frequency0.0010.001 | Frequency | 0.001 | 0.12 | 0.001 | kHz | Hz |  |  |  | 0.57736 | 0.0003329 | 0.57736 | 0.00014506
| Frequency0.010.001 | Frequency | 0.001 | 0.12 | 0.01 | kHz | Hz |  |  |  | 1.2926 | 8.9649e-05 | 5.7736 | 1.4506e-05
| Frequency0.00010.12 | Frequency | 0.12 | 1.2 | 0.0001 | kHz | Hz |  |  |  | 0.058 | 0.0095 | 0.056807 | 0.0092702
| Frequency0.0010.12 | Frequency | 0.12 | 1.2 | 0.001 | kHz | Hz |  |  |  | 0.57742 | 0.0011566 | 0.57725 | 0.0010094
| Frequency0.010.12 | Frequency | 0.12 | 1.2 | 0.01 | kHz | Hz |  |  |  | 5.7736 | 0.00016017 | 5.7735 | 0.00010104
| Frequency0.10.12 | Frequency | 0.12 | 1.2 | 0.1 | kHz | Hz |  |  |  | 57.736 | 3.376e-05 | 57.736 | 1.0104e-05
| Frequency0.0011.2 | Frequency | 1.2 | 12 | 0.001 | kHz | Hz |  |  |  | 0.58 | 0.0091 | 0.56785 | 0.0088368
| Frequency0.011.2 | Frequency | 1.2 | 12 | 0.01 | kHz | Hz |  |  |  | 5.7742 | 0.0011047 | 5.7725 | 0.00095754
| Frequency0.11.2 | Frequency | 1.2 | 12 | 0.1 | kHz | Hz |  |  |  | 57.736 | 0.00015498 | 57.735 | 9.584e-05
| Frequency11.2 | Frequency | 1.2 | 12 | 1 | kHz | Hz |  |  |  | 577.36 | 3.324e-05 | 577.36 | 9.5841e-06
| Frequency0.0112 | Frequency | 12 | 120 | 0.01 | kHz | Hz |  |  |  | 19 | 0.028 | 18.255 | 0.027641
| Frequency0.112 | Frequency | 12 | 120 | 0.1 | kHz | Hz |  |  |  | 60.338 | 0.0096628 | 60.269 | 0.0091178
| Frequency112 | Frequency | 12 | 120 | 1 | kHz | Hz |  |  |  | 577.65 | 0.0011971 | 577.61 | 0.00096075
| Frequency1012 | Frequency | 12 | 120 | 10 | kHz | Hz |  |  |  | 5773.6 | 0.00019075 | 5773.6 | 9.6128e-05
| Frequency0.1120 | Frequency | 120 | 1200 | 0.1 | kHz | Hz |  |  |  | 60 | 0.016 | 59.55 | 0.015104
| Frequency1120 | Frequency | 120 | 1200 | 1 | kHz | Hz |  |  |  | 577.64 | 0.0023288 | 577.51 | 0.0018144
| Frequency10120 | Frequency | 120 | 1200 | 10 | kHz | Hz |  |  |  | 5773.6 | 0.00038884 | 5773.6 | 0.00018186
| Frequency100120 | Frequency | 120 | 1200 | 100 | kHz | Hz |  |  |  | 57736 | 0.00010098 | 57736 | 1.8187e-05
| Frequency11200 | Frequency | 1200 | 2000 | 1 | kHz | kHz |  |  |  | 0.43 | 0.13 | 0.42565 | 0.12836
| Frequency101200 | Frequency | 1200 | 2000 | 10 | kHz | kHz |  |  |  | 5.7564 | 0.016002 | 5.757 | 0.014015
| Frequency1001200 | Frequency | 1200 | 2000 | 100 | kHz | kHz |  |  |  | 57.734 | 0.0022013 | 57.734 | 0.0014029
| Frequency10001200 | Frequency | 1200 | 2000 | 1000 | kHz | kHz |  |  |  | 577.36 | 0.00045965 | 577.36 | 0.0001403
| ACA0.00010.0000295 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 5 | 10 | kHz | 57.729 | 302.48 | 57.729 | 302.01
| ACA0.000000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-08 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46198 | 18475
| ACA0.00000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-07 | A | µA | 10 | 30 | kHz | 0.62 | 18000 | 0.46389 | 18470
| ACA0.0000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-06 | A | µA | 10 | 30 | kHz | 0.6298 | 18044 | 0.62958 | 18044
| ACA0.000010.00002910 | ACA | 2.9e-05 | 0.00033 | 1e-05 | A | µA | 10 | 30 | kHz | 5.582 | 9564.6 | 5.5818 | 9564.1
| ACA0.00010.00002910 | ACA | 2.9e-05 | 0.00033 | 0.0001 | A | µA | 10 | 30 | kHz | 57.709 | 1205.5 | 57.709 | 1205.1
| ACA0.000000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 0.01 | 0.02 | kHz | 0.23 | 2300 | 0.17326 | 2315.4
| ACA0.00000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 0.01 | 0.02 | kHz | 0.23 | 2304.7 | 0.17481 | 2314.9
| ACA0.0000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 0.01 | 0.02 | kHz | 0.35294 | 2267.3 | 0.35227 | 2267.5
| ACA0.000010.000330.01 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 0.01 | 0.02 | kHz | 5.4199 | 1301.8 | 5.4195 | 1301.8
| ACA0.00010.000330.01 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 0.01 | 0.02 | kHz | 57.686 | 174.69 | 57.685 | 174.66
| ACA0.000000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 0.32 | 1400 | 0.17327 | 1443.8
| ACA0.00000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 0.32 | 1400 | 0.17604 | 1443.1
| ACA0.0000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 0.41399 | 1381.1 | 0.41349 | 1381.2
| ACA0.000010.000330.02 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 5.6118 | 601.75 | 5.6114 | 601.72
| ACA0.00010.000330.02 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 57.716 | 69.766 | 57.716 | 69.735
| ACA0.000000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17375 | 1155.2
| ACA0.00000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 0.045 | 1 | kHz | 0.18 | 1200 | 0.17651 | 1154.4
| ACA0.0000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 0.045 | 1 | kHz | 0.43013 | 1124.2 | 0.44201 | 1086.4
| ACA0.000010.000330.045 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 0.045 | 1 | kHz | 5.6654 | 409.29 | 5.6651 | 409.25
| ACA0.00010.000330.045 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 0.045 | 1 | kHz | 57.723 | 45.397 | 57.723 | 45.366
| ACA0.000000010.000331 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 1 | 5 | kHz | 0.27 | 2300 | 0.23112 | 2309.7
| ACA0.00000010.000331 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 1 | 5 | kHz | 0.27 | 2300 | 0.2328 | 2309.2
| ACA0.0000010.000331 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 1 | 5 | kHz | 0.40201 | 2264.3 | 0.40156 | 2264.4
| ACA0.000010.000331 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 1 | 5 | kHz | 5.4266 | 1309.3 | 5.4263 | 1309.3
| ACA0.00010.000331 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 1 | 5 | kHz | 57.686 | 176.15 | 57.686 | 176.12
| ACA0.000000010.000335 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 5 | 10 | kHz | 0.35 | 5800 | 0.34654 | 5773.6
| ACA0.00000010.000335 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 5 | 10 | kHz | 0.35 | 5800 | 0.34724 | 5773.4
| ACA0.0000010.000335 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 5 | 10 | kHz | 0.41769 | 5779.5 | 0.4264 | 5752
| ACA0.000010.000335 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 5 | 10 | kHz | 4.6373 | 4728.3 | 4.6369 | 4728.3
| ACA0.00010.000335 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 5 | 10 | kHz | 57.432 | 1053.3 | 57.432 | 1053.2
| ACA0.000000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-08 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69332 | 11547
| ACA0.00000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-07 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.69339 | 11547
| ACA0.0000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-06 | A | µA | 10 | 30 | kHz | 0.7 | 12000 | 0.7339 | 11536
| ACA0.000010.0003310 | ACA | 0.00033 | 0.0033 | 1e-05 | A | µA | 10 | 30 | kHz | 3.6592 | 11103 | 3.7773 | 10742
| ACA0.00010.0003310 | ACA | 0.00033 | 0.0033 | 0.0001 | A | µA | 10 | 30 | kHz | 56.617 | 3922.7 | 56.616 | 3922.6
| ACA0.000000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 0.01 | 0.02 | kHz | 2.4 | 2100 | 2.3005 | 2087.6
| ACA0.00000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 0.01 | 0.02 | kHz | 2.4 | 2100 | 2.3157 | 2087.1
| ACA0.0000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 0.01 | 0.02 | kHz | 4.0897 | 2048.8 | 4.1185 | 2039.5
| ACA0.000010.00330.01 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 0.01 | 0.02 | kHz | 54.774 | 1117.8 | 54.773 | 1117.7
| ACA0.00010.00330.01 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 0.01 | 0.02 | kHz | 576.95 | 144.81 | 576.95 | 144.75
| ACA0.000000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 0.02 | 0.045 | kHz | 3.8 | 1000 | 2.2968 | 1043.6
| ACA0.00000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 0.02 | 0.045 | kHz | 3.8 | 1001.4 | 2.3282 | 1042.8
| ACA0.0000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 0.02 | 0.045 | kHz | 4.917 | 977.92 | 4.9145 | 977.93
| ACA0.000010.00330.02 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 0.02 | 0.045 | kHz | 56.865 | 350.66 | 56.863 | 350.55
| ACA0.00010.00330.02 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 0.02 | 0.045 | kHz | 577.26 | 38.407 | 577.26 | 38.349
| ACA0.000000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 0.045 | 1 | kHz | 2.5 | 460 | 2.2998 | 465.62
| ACA0.00000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 0.045 | 1 | kHz | 2.5 | 463.22 | 2.3464 | 464.49
| ACA0.0000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 0.045 | 1 | kHz | 5.6396 | 392.41 | 5.6369 | 392.36
| ACA0.000010.00330.045 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 0.045 | 1 | kHz | 57.584 | 84.81 | 57.583 | 84.672
| ACA0.00010.00330.045 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 0.045 | 1 | kHz | 577.34 | 8.7266 | 577.34 | 8.6681
| ACA0.000000010.00331 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 1 | 5 | kHz | 2.5 | 920 | 2.3048 | 925.75
| ACA0.00000010.00331 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 1 | 5 | kHz | 2.5 | 920.01 | 2.3386 | 924.88
| ACA0.0000010.00331 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 1 | 5 | kHz | 5.0491 | 857.92 | 5.0466 | 857.92
| ACA0.000010.00331 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 1 | 5 | kHz | 57.047 | 284.47 | 57.045 | 284.35
| ACA0.00010.00331 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 1 | 5 | kHz | 577.28 | 30.67 | 577.28 | 30.612
| ACA0.000000010.00335 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 5 | 10 | kHz | 3.8 | 2300 | 3.4644 | 2309.4
| ACA0.00000010.00335 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 5 | 10 | kHz | 3.8 | 2300 | 3.4807 | 2309
| ACA0.0000010.00335 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 5 | 10 | kHz | 5.0186 | 2268.7 | 5.0129 | 2268.8
| ACA0.000010.00335 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 5 | 10 | kHz | 54.397 | 1333.5 | 54.39 | 1333.5
| ACA0.00010.00335 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 5 | 10 | kHz | 576.87 | 180.75 | 576.87 | 180.65
| ACA0.000000010.003310 | ACA | 0.0033 | 0.033 | 1e-08 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6163 | 4619.9
| ACA0.00000010.003310 | ACA | 0.0033 | 0.033 | 1e-07 | A | µA | 10 | 30 | kHz | 5.3 | 4600 | 4.6262 | 4619.6
| ACA0.0000010.003310 | ACA | 0.0033 | 0.033 | 1e-06 | A | µA | 10 | 30 | kHz | 5.5294 | 4595.5 | 5.5179 | 4595.8
| ACA0.000010.003310 | ACA | 0.0033 | 0.033 | 1e-05 | A | µA | 10 | 30 | kHz | 49.259 | 3578.8 | 49.247 | 3578.8
| ACA0.00010.003310 | ACA | 0.0033 | 0.033 | 0.0001 | A | µA | 10 | 30 | kHz | 575.41 | 695.3 | 575.4 | 695.06
| ACA0.0000010.0330.01 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 0.01 | 0.02 | kHz | 0.024 | 2.1 | 0.022917 | 2.0907
| ACA0.000010.0330.01 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 0.01 | 0.02 | kHz | 0.024 | 2.1 | 0.023129 | 2.0901
| ACA0.00010.0330.01 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 0.01 | 0.02 | kHz | 0.040941 | 2.0487 | 0.041148 | 2.0425
| ACA0.0010.0330.01 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 0.01 | 0.02 | kHz | 0.54767 | 1.1202 | 0.54766 | 1.1202
| ACA0.010.0330.01 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 0.01 | 0.02 | kHz | 5.7695 | 0.14517 | 5.7695 | 0.14515
| ACA0.0000010.0330.02 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 0.02 | 0.045 | kHz | 0.038 | 1 | 0.023094 | 1.0436
| ACA0.000010.0330.02 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 0.02 | 0.045 | kHz | 0.038 | 1.0016 | 0.023358 | 1.0428
| ACA0.00010.0330.02 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 0.02 | 0.045 | kHz | 0.049203 | 0.97804 | 0.049192 | 0.97804
| ACA0.0010.0330.02 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 0.02 | 0.045 | kHz | 0.56864 | 0.35073 | 0.56864 | 0.35068
| ACA0.010.0330.02 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 0.02 | 0.045 | kHz | 5.7726 | 0.03839 | 5.7726 | 0.038365
| ACA0.0000010.0330.045 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.47 | 0.02282 | 0.47204
| ACA0.000010.0330.045 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.024 | 0.46973 | 0.023285 | 0.47091
| ACA0.00010.0330.045 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.056197 | 0.39853 | 0.056185 | 0.39851
| ACA0.0010.0330.045 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.57578 | 0.086645 | 0.57577 | 0.086584
| ACA0.010.0330.045 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7734 | 0.0088941 | 5.7734 | 0.0088682
| ACA0.0000010.0331 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 1 | 5 | kHz | 0.058 | 1.2 | 0.057624 | 1.1588
| ACA0.000010.0331 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 1 | 5 | kHz | 0.058 | 1.2 | 0.057825 | 1.1583
| ACA0.00010.0331 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 1 | 5 | kHz | 0.073934 | 1.1517 | 0.075029 | 1.1175
| ACA0.0010.0331 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 1 | 5 | kHz | 0.56965 | 0.47368 | 0.56963 | 0.47361
| ACA0.010.0331 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 1 | 5 | kHz | 5.7726 | 0.053742 | 5.7726 | 0.053697
| ACA0.0000010.0335 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 5 | 10 | kHz | 0.12 | 2.3 | 0.11543 | 2.3114
| ACA0.000010.0335 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 5 | 10 | kHz | 0.12 | 2.3 | 0.11552 | 2.3112
| ACA0.00010.0335 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 5 | 10 | kHz | 0.12471 | 2.2891 | 0.12466 | 2.2892
| ACA0.0010.0335 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 5 | 10 | kHz | 0.55922 | 1.4904 | 0.55916 | 1.4904
| ACA0.010.0335 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 5 | 10 | kHz | 5.7697 | 0.21298 | 5.7697 | 0.21288
| ACA0.0000010.03310 | ACA | 0.033 | 0.33 | 1e-06 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23085 | 4.6219
| ACA0.000010.03310 | ACA | 0.033 | 0.33 | 1e-05 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.2309 | 4.6218
| ACA0.00010.03310 | ACA | 0.033 | 0.33 | 0.0001 | A | mA | 10 | 30 | kHz | 0.24 | 4.6 | 0.23556 | 4.6105
| ACA0.0010.03310 | ACA | 0.033 | 0.33 | 0.001 | A | mA | 10 | 30 | kHz | 0.56478 | 3.8905 | 0.56466 | 3.8906
| ACA0.010.03310 | ACA | 0.033 | 0.33 | 0.01 | A | mA | 10 | 30 | kHz | 5.7587 | 0.83674 | 5.7587 | 0.83651
| ACA0.000010.330.01 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.01 | 0.045 | kHz | 0.12 | 2.1 | 0.11617 | 2.0784
| ACA0.00010.330.01 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.01 | 0.045 | kHz | 0.12 | 2.1 | 0.11882 | 2.0766
| ACA0.0010.330.01 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 0.35291 | 1.9255 | 0.35284 | 1.9254
| ACA0.010.330.01 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 5.6471 | 0.55145 | 5.6471 | 0.55121
| ACA0.10.330.01 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 57.722 | 0.057755 | 57.722 | 0.057649
| ACA0.000010.330.045 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.12 | 0.58 | 0.11734 | 0.57739
| ACA0.00010.330.045 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.12267 | 0.57758 | 0.124 | 0.57333
| ACA0.0010.330.045 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.52825 | 0.38235 | 0.52823 | 0.382
| ACA0.010.330.045 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.7643 | 0.053023 | 5.7643 | 0.05276
| ACA0.10.330.045 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.735 | 0.005408 | 57.735 | 0.0053018
| ACA0.000010.331 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 1 | 5 | kHz | 1.2 | 6.9 | 1.1549 | 6.9282
| ACA0.00010.331 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 1 | 5 | kHz | 1.2 | 6.9 | 1.1555 | 6.9278
| ACA0.0010.331 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 1 | 5 | kHz | 1.2164 | 6.8903 | 1.2155 | 6.8904
| ACA0.010.331 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 1 | 5 | kHz | 5.1011 | 4.9149 | 5.0998 | 4.9136
| ACA0.10.331 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 1 | 5 | kHz | 57.598 | 0.72965 | 57.598 | 0.72819
| ACA0.000010.335 | ACA | 0.33 | 1.1 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7734 | 28.868
| ACA0.00010.335 | ACA | 0.33 | 1.1 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7742 | 28.867
| ACA0.0010.335 | ACA | 0.33 | 1.1 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7874 | 28.859
| ACA0.010.335 | ACA | 0.33 | 1.1 | 0.01 | A | mA | 5 | 10 | kHz | 7.0908 | 28.073 | 7.0891 | 28.073
| ACA0.10.335 | ACA | 0.33 | 1.1 | 0.1 | A | mA | 5 | 10 | kHz | 55.816 | 11.862 | 55.815 | 11.86
| ACA0.000011.10.01 | ACA | 1.1 | 3 | 1e-05 | A | mA | 0.01 | 0.045 | kHz | 0.12 | 2.1 | 0.11589 | 2.0787
| ACA0.00011.10.01 | ACA | 1.1 | 3 | 0.0001 | A | mA | 0.01 | 0.045 | kHz | 0.12 | 2.1 | 0.11676 | 2.0785
| ACA0.0011.10.01 | ACA | 1.1 | 3 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 0.18892 | 2.077 | 0.20868 | 2.0565
| ACA0.011.10.01 | ACA | 1.1 | 3 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 4.9071 | 1.2281 | 4.9043 | 1.2265
| ACA0.11.10.01 | ACA | 1.1 | 3 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 57.613 | 0.15838 | 57.613 | 0.15706
| ACA0.000011.10.045 | ACA | 1.1 | 3 | 1e-05 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11619 | 0.69369
| ACA0.00011.10.045 | ACA | 1.1 | 3 | 0.0001 | A | mA | 0.045 | 1 | kHz | 0.13 | 0.69 | 0.11868 | 0.6931
| ACA0.0011.10.045 | ACA | 1.1 | 3 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.35192 | 0.64234 | 0.34551 | 0.6421
| ACA0.011.10.045 | ACA | 1.1 | 3 | 0.01 | A | mA | 0.045 | 1 | kHz | 5.6463 | 0.18057 | 5.6448 | 0.17758
| ACA0.11.10.045 | ACA | 1.1 | 3 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.722 | 0.019812 | 57.722 | 0.018474
| ACA0.000011.11 | ACA | 1.1 | 3 | 1e-05 | A | mA | 1 | 5 | kHz | 1.3 | 6.9 | 1.1548 | 6.9283
| ACA0.00011.11 | ACA | 1.1 | 3 | 0.0001 | A | mA | 1 | 5 | kHz | 1.3 | 6.9 | 1.1552 | 6.9282
| ACA0.0011.11 | ACA | 1.1 | 3 | 0.001 | A | mA | 1 | 5 | kHz | 1.3 | 6.9 | 1.1804 | 6.9223
| ACA0.011.11 | ACA | 1.1 | 3 | 0.01 | A | mA | 1 | 5 | kHz | 3.4547 | 6.4116 | 3.4521 | 6.4115
| ACA0.11.11 | ACA | 1.1 | 3 | 0.1 | A | mA | 1 | 5 | kHz | 56.451 | 1.7722 | 56.451 | 1.771
| ACA0.000011.15 | ACA | 1.1 | 3 | 1e-05 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.773 | 28.868
| ACA0.00011.15 | ACA | 1.1 | 3 | 0.0001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.7731 | 28.868
| ACA0.0011.15 | ACA | 1.1 | 3 | 0.001 | A | mA | 5 | 10 | kHz | 5.8 | 29 | 5.78 | 28.866
| ACA0.011.15 | ACA | 1.1 | 3 | 0.01 | A | mA | 5 | 10 | kHz | 6.2284 | 28.857 | 6.3664 | 28.73
| ACA0.11.15 | ACA | 1.1 | 3 | 0.1 | A | mA | 5 | 10 | kHz | 45.662 | 21.093 | 45.659 | 21.092
| ACA0.000130.045 | ACA | 3 | 11 | 0.0001 | A | mA | 0.045 | 0.1 | kHz | 2.4 | 0.69 | 2.3298 | 0.69373
| ACA0.00130.045 | ACA | 3 | 11 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 2.4 | 0.69 | 2.3709 | 0.69112
| ACA0.0130.045 | ACA | 3 | 11 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 5.6644 | 0.53041 | 5.6622 | 0.53026
| ACA0.130.045 | ACA | 3 | 11 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 57.647 | 0.085551 | 57.646 | 0.085343
| ACA130.045 | ACA | 3 | 11 | 1 | A | mA | 0.045 | 0.1 | kHz | 577.35 | 0.0086938 | 577.35 | 0.0086087
| ACA0.000130.1 | ACA | 3 | 11 | 0.0001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.324 | 1.1553
| ACA0.00130.1 | ACA | 3 | 11 | 0.001 | A | mA | 0.1 | 1 | kHz | 2.4 | 1.2 | 2.3551 | 1.1531
| ACA0.0130.1 | ACA | 3 | 11 | 0.01 | A | mA | 0.1 | 1 | kHz | 5.1972 | 0.99041 | 5.1951 | 0.99036
| ACA0.130.1 | ACA | 3 | 11 | 0.1 | A | mA | 0.1 | 1 | kHz | 57.412 | 0.20433 | 57.411 | 0.20413
| ACA130.1 | ACA | 3 | 11 | 1 | A | mA | 0.1 | 1 | kHz | 577.32 | 0.020884 | 577.32 | 0.020799
| ACA0.000131 | ACA | 3 | 11 | 0.0001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3103 | 34.641
| ACA0.00131 | ACA | 3 | 11 | 0.001 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.3112 | 34.641
| ACA0.0131 | ACA | 3 | 11 | 0.01 | A | mA | 1 | 5 | kHz | 2.5 | 35 | 2.5084 | 34.627
| ACA0.131 | ACA | 3 | 11 | 0.1 | A | mA | 1 | 5 | kHz | 20.888 | 33.346 | 20.867 | 33.347
| ACA131 | ACA | 3 | 11 | 1 | A | mA | 1 | 5 | kHz | 547.31 | 13.25 | 547.3 | 13.249
| ACA0.0001110.045 | ACA | 11 | 20.5 | 0.0001 | A | mA | 0.045 | 0.1 | kHz | 5.8 | 1.4 | 5.7745 | 1.386
| ACA0.001110.045 | ACA | 11 | 20.5 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 5.8 | 1.4 | 5.7855 | 1.3857
| ACA0.01110.045 | ACA | 11 | 20.5 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 6.9249 | 1.3567 | 6.9028 | 1.3548
| ACA0.1110.045 | ACA | 11 | 20.5 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 55.153 | 0.59996 | 55.168 | 0.5931
| ACA1110.045 | ACA | 11 | 20.5 | 1 | A | mA | 0.045 | 0.1 | kHz | 579.65 | 0.069317 | 579.66 | 0.065887
| ACA0.0001110.1 | ACA | 11 | 20.5 | 0.0001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7754 | 1.7324
| ACA0.001110.1 | ACA | 11 | 20.5 | 0.001 | A | mA | 0.1 | 1 | kHz | 6.5 | 1.7 | 5.7852 | 1.7321
| ACA0.01110.1 | ACA | 11 | 20.5 | 0.01 | A | mA | 0.1 | 1 | kHz | 6.7671 | 1.7062 | 6.7466 | 1.7047
| ACA0.1110.1 | ACA | 11 | 20.5 | 0.1 | A | mA | 0.1 | 1 | kHz | 53.699 | 0.85925 | 53.709 | 0.85296
| ACA1110.1 | ACA | 11 | 20.5 | 1 | A | mA | 0.1 | 1 | kHz | 579.44 | 0.102 | 579.45 | 0.098572
| ACA0.0001111 | ACA | 11 | 20.5 | 0.0001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7742 | 34.641
| ACA0.001111 | ACA | 11 | 20.5 | 0.001 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.7745 | 34.641
| ACA0.01111 | ACA | 11 | 20.5 | 0.01 | A | mA | 1 | 5 | kHz | 5.9 | 35 | 5.8395 | 34.639
| ACA0.1111 | ACA | 11 | 20.5 | 0.1 | A | mA | 1 | 5 | kHz | 6.5765 | 34.962 | 12.386 | 34.433
| ACA1111 | ACA | 11 | 20.5 | 1 | A | mA | 1 | 5 | kHz | 437.55 | 23.603 | 437.54 | 23.602
| ACA 10 turn0.000110.01 | ACA 10 turn | 1 | 11 | 0.0001 | A | mA | 0.01 | 0.045 | kHz | 3.3 | 2.7 | 0.97318 | 2.873
| ACA 10 turn0.00110.01 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 4.3265 | 2.6149 | 4.3204 | 2.6148
| ACA 10 turn0.0110.01 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 57.287 | 0.84059 | 57.285 | 0.83952
| ACA 10 turn0.110.01 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 579.92 | 0.090643 | 579.92 | 0.090136
| ACA 10 turn110.01 | ACA 10 turn | 1 | 11 | 1 | A | mA | 0.01 | 0.045 | kHz | 5800 | 0.0092241 | 5800 | 0.0090207
| ACA 10 turn1010.01 | ACA 10 turn | 1 | 11 | 10 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.00098346 | 58000 | 0.00090208
| ACA 10 turn0.00110.045 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.66 | 2.1 | 0.65508 | 2.0543
| ACA 10 turn0.0110.045 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 0.045 | 1 | kHz | 4.6183 | 1.7587 | 4.6169 | 1.7583
| ACA 10 turn0.110.045 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 0.045 | 1 | kHz | 57.618 | 0.44323 | 57.618 | 0.44244
| ACA 10 turn110.045 | ACA 10 turn | 1 | 11 | 1 | A | mA | 0.045 | 1 | kHz | 579.96 | 0.046302 | 579.96 | 0.045959
| ACA 10 turn1010.045 | ACA 10 turn | 1 | 11 | 10 | A | mA | 0.045 | 1 | kHz | 5800 | 0.004735 | 5800 | 0.0045977
| ACA 10 turn0.00111 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 1 | 5 | kHz | 12 | 7.2 | 11.405 | 7.1882
| ACA 10 turn0.0111 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 1 | 5 | kHz | 12.361 | 7.1187 | 12.349 | 7.1191
| ACA 10 turn0.111 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 1 | 5 | kHz | 56.262 | 4.6574 | 56.248 | 4.6564
| ACA 10 turn111 | ACA 10 turn | 1 | 11 | 1 | A | mA | 1 | 5 | kHz | 579.63 | 0.67288 | 579.63 | 0.67163
| ACA 10 turn1011 | ACA 10 turn | 1 | 11 | 10 | A | mA | 1 | 5 | kHz | 5800 | 0.068091 | 5800 | 0.067582
| ACA 10 turn0.00115 | ACA 10 turn | 1 | 11 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.703 | 28.929
| ACA 10 turn0.0115 | ACA 10 turn | 1 | 11 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 28.914 | 57.905 | 28.915
| ACA 10 turn0.115 | ACA 10 turn | 1 | 11 | 0.1 | A | mA | 5 | 10 | kHz | 76.668 | 27.611 | 76.638 | 27.612
| ACA 10 turn115 | ACA 10 turn | 1 | 11 | 1 | A | mA | 5 | 10 | kHz | 575.98 | 10.476 | 575.96 | 10.474
| ACA 10 turn1015 | ACA 10 turn | 1 | 11 | 10 | A | mA | 5 | 10 | kHz | 5799.5 | 1.1535 | 5799.5 | 1.1523
| ACA 10 turn0.001110.01 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 0.87 | 2.9 | 0.85389 | 2.8839
| ACA 10 turn0.01110.01 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 1.4281 | 2.8814 | 1.5429 | 2.8673
| ACA 10 turn0.1110.01 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 44.394 | 2.0188 | 44.317 | 2.0184
| ACA 10 turn1110.01 | ACA 10 turn | 11 | 30 | 1 | A | mA | 0.01 | 0.045 | kHz | 577.68 | 0.29774 | 577.66 | 0.2963
| ACA 10 turn10110.01 | ACA 10 turn | 11 | 30 | 10 | A | mA | 0.01 | 0.045 | kHz | 5799.8 | 0.030408 | 5799.8 | 0.029819
| ACA 10 turn0.001110.045 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 0.045 | 1 | kHz | 0.9 | 2.1 | 0.42297 | 2.1155
| ACA 10 turn0.01110.045 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 0.045 | 1 | kHz | 1.3908 | 2.0927 | 1.3648 | 2.0928
| ACA 10 turn0.1110.045 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 0.045 | 1 | kHz | 48.986 | 1.2457 | 48.967 | 1.244
| ACA 10 turn1110.045 | ACA 10 turn | 11 | 30 | 1 | A | mA | 0.045 | 1 | kHz | 578.74 | 0.1605 | 578.73 | 0.15918
| ACA 10 turn10110.045 | ACA 10 turn | 11 | 30 | 10 | A | mA | 0.045 | 1 | kHz | 5799.9 | 0.016507 | 5799.9 | 0.015972
| ACA 10 turn0.001111 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 1 | 5 | kHz | 12 | 7.2 | 11.174 | 7.2092
| ACA 10 turn0.01111 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 1 | 5 | kHz | 12 | 7.2 | 11.422 | 7.2034
| ACA 10 turn0.1111 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 1 | 5 | kHz | 33.827 | 6.6978 | 33.792 | 6.6979
| ACA 10 turn1111 | ACA 10 turn | 11 | 30 | 1 | A | mA | 1 | 5 | kHz | 566.19 | 1.8955 | 566.18 | 1.8942
| ACA 10 turn10111 | ACA 10 turn | 11 | 30 | 10 | A | mA | 1 | 5 | kHz | 5798.5 | 0.19808 | 5798.5 | 0.1975
| ACA 10 turn0.001115 | ACA 10 turn | 11 | 30 | 0.001 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.627 | 28.936
| ACA 10 turn0.01115 | ACA 10 turn | 11 | 30 | 0.01 | A | mA | 5 | 10 | kHz | 58 | 29 | 57.682 | 28.935
| ACA 10 turn0.1115 | ACA 10 turn | 11 | 30 | 0.1 | A | mA | 5 | 10 | kHz | 63.353 | 28.822 | 63.613 | 28.797
| ACA 10 turn1115 | ACA 10 turn | 11 | 30 | 1 | A | mA | 5 | 10 | kHz | 458.92 | 21.118 | 458.89 | 21.117
| ACA 10 turn10115 | ACA 10 turn | 11 | 30 | 10 | A | mA | 5 | 10 | kHz | 5776.7 | 3.2241 | 5776.7 | 3.2231
| ACA 10 turn0.01300.045 | ACA 10 turn | 30 | 110 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 12 | 2.1 | 11.984 | 2.086
| ACA 10 turn0.1300.045 | ACA 10 turn | 30 | 110 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 36.579 | 1.9251 | 36.551 | 1.9245
| ACA 10 turn1300.045 | ACA 10 turn | 30 | 110 | 1 | A | mA | 0.045 | 0.1 | kHz | 568.43 | 0.54491 | 568.44 | 0.54347
| ACA 10 turn10300.045 | ACA 10 turn | 30 | 110 | 10 | A | mA | 0.045 | 0.1 | kHz | 5798.8 | 0.057436 | 5798.8 | 0.056801
| ACA 10 turn100300.045 | ACA 10 turn | 30 | 110 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0059371 | 58000 | 0.0056828
| ACA 10 turn0.01300.1 | ACA 10 turn | 30 | 110 | 0.01 | A | mA | 0.1 | 1 | kHz | 15 | 2.3 | 14.872 | 2.2866
| ACA 10 turn0.1300.1 | ACA 10 turn | 30 | 110 | 0.1 | A | mA | 0.1 | 1 | kHz | 37.348 | 2.1392 | 37.319 | 2.1387
| ACA 10 turn1300.1 | ACA 10 turn | 30 | 110 | 1 | A | mA | 0.1 | 1 | kHz | 566.33 | 0.65492 | 566.34 | 0.65351
| ACA 10 turn10300.1 | ACA 10 turn | 30 | 110 | 10 | A | mA | 0.1 | 1 | kHz | 5798.5 | 0.069561 | 5798.5 | 0.068926
| ACA 10 turn100300.1 | ACA 10 turn | 30 | 110 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0071509 | 58000 | 0.0068966
| ACA 10 turn0.01301 | ACA 10 turn | 30 | 110 | 0.01 | A | mA | 1 | 5 | kHz | 24 | 35 | 23.064 | 34.699
| ACA 10 turn0.1301 | ACA 10 turn | 30 | 110 | 0.1 | A | mA | 1 | 5 | kHz | 24 | 35 | 25.112 | 34.684
| ACA 10 turn1301 | ACA 10 turn | 30 | 110 | 1 | A | mA | 1 | 5 | kHz | 210.04 | 33.396 | 210.02 | 33.395
| ACA 10 turn10301 | ACA 10 turn | 30 | 110 | 10 | A | mA | 1 | 5 | kHz | 5499.6 | 13.241 | 5499.7 | 13.239
| ACA 10 turn100301 | ACA 10 turn | 30 | 110 | 100 | A | mA | 1 | 5 | kHz | 57966 | 1.4661 | 57966 | 1.4652
| ACA 10 turn0.011100.045 | ACA 10 turn | 110 | 205 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 42 | 2.4 | 38.553 | 2.4159
| ACA 10 turn0.11100.045 | ACA 10 turn | 110 | 205 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 46.787 | 2.3924 | 46.65 | 2.3916
| ACA 10 turn11100.045 | ACA 10 turn | 110 | 205 | 1 | A | mA | 0.045 | 0.1 | kHz | 500.59 | 1.4072 | 500.63 | 1.403
| ACA 10 turn101100.045 | ACA 10 turn | 110 | 205 | 10 | A | mA | 0.045 | 0.1 | kHz | 5788.7 | 0.17673 | 5788.8 | 0.17407
| ACA 10 turn1001100.045 | ACA 10 turn | 110 | 205 | 100 | A | mA | 0.045 | 0.1 | kHz | 57999 | 0.018526 | 57999 | 0.017455
| ACA 10 turn0.011100.1 | ACA 10 turn | 110 | 205 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 42.202 | 2.6324
| ACA 10 turn0.11100.1 | ACA 10 turn | 110 | 205 | 0.1 | A | mA | 0.1 | 1 | kHz | 49.765 | 2.6108 | 49.649 | 2.61
| ACA 10 turn11100.1 | ACA 10 turn | 110 | 205 | 1 | A | mA | 0.1 | 1 | kHz | 490.55 | 1.6184 | 490.58 | 1.6144
| ACA 10 turn101100.1 | ACA 10 turn | 110 | 205 | 10 | A | mA | 0.1 | 1 | kHz | 5786.7 | 0.2093 | 5786.8 | 0.20664
| ACA 10 turn1001100.1 | ACA 10 turn | 110 | 205 | 100 | A | mA | 0.1 | 1 | kHz | 57999 | 0.021803 | 57999 | 0.020732
| ACA 10 turn0.011101 | ACA 10 turn | 110 | 205 | 0.01 | A | mA | 1 | 5 | kHz | 58 | 35 | 57.612 | 34.699
| ACA 10 turn0.11101 | ACA 10 turn | 110 | 205 | 0.1 | A | mA | 1 | 5 | kHz | 58 | 35 | 58.262 | 34.697
| ACA 10 turn11101 | ACA 10 turn | 110 | 205 | 1 | A | mA | 1 | 5 | kHz | 79.07 | 34.897 | 123.66 | 34.491
| ACA 10 turn101101 | ACA 10 turn | 110 | 205 | 10 | A | mA | 1 | 5 | kHz | 4372.5 | 23.662 | 4372.4 | 23.661
| ACA 10 turn1001101 | ACA 10 turn | 110 | 205 | 100 | A | mA | 1 | 5 | kHz | 57768 | 3.2889 | 57768 | 3.2878
| ACA 20 turn0.001220.01 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 0.01 | 0.045 | kHz | -5.2 | 3 | -6.3515 | 3.0358
| ACA 20 turn0.01220.01 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 0.01 | 0.045 | kHz | -5.2 | 3.0309 | -6.0166 | 3.031
| ACA 20 turn0.1220.01 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 24.504 | 2.6609 | 24.465 | 2.6608
| ACA 20 turn1220.01 | ACA 20 turn | 22 | 60 | 1 | A | mA | 0.01 | 0.045 | kHz | 569.88 | 0.59872 | 569.88 | 0.59787
| ACA 20 turn10220.01 | ACA 20 turn | 22 | 60 | 10 | A | mA | 0.01 | 0.045 | kHz | 5799 | 0.061619 | 5799 | 0.061255
| ACA 20 turn0.001220.045 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -5.078 | 2.227
| ACA 20 turn0.01220.045 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 0.045 | 1 | kHz | 15 | 2.3 | -4.615 | 2.2203
| ACA 20 turn0.1220.045 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 0.045 | 1 | kHz | 25.632 | 2.1228 | 32.901 | 1.7903
| ACA 20 turn1220.045 | ACA 20 turn | 22 | 60 | 1 | A | mA | 0.045 | 1 | kHz | 574.48 | 0.32468 | 574.48 | 0.3238
| ACA 20 turn10220.045 | ACA 20 turn | 22 | 60 | 10 | A | mA | 0.045 | 1 | kHz | 5799.4 | 0.033173 | 5799.4 | 0.032809
| ACA 20 turn0.001221 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 1 | 5 | kHz | 2.3 | 7.6 | 2.2215 | 7.5886
| ACA 20 turn0.01221 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 1 | 5 | kHz | 2.3482 | 7.5869 | 2.3433 | 7.5869
| ACA 20 turn0.1221 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 1 | 5 | kHz | 14.11 | 7.4273 | 14.091 | 7.4273
| ACA 20 turn1221 | ACA 20 turn | 22 | 60 | 1 | A | mA | 1 | 5 | kHz | 526.05 | 3.5184 | 526.05 | 3.5178
| ACA 20 turn10221 | ACA 20 turn | 22 | 60 | 10 | A | mA | 1 | 5 | kHz | 5793.5 | 0.40619 | 5793.5 | 0.40583
| ACA 20 turn0.001225 | ACA 20 turn | 22 | 60 | 0.001 | A | mA | 5 | 10 | kHz | 52 | 30 | 34.529 | 30.459
| ACA 20 turn0.01225 | ACA 20 turn | 22 | 60 | 0.01 | A | mA | 5 | 10 | kHz | 52 | 30.458 | 34.545 | 30.459
| ACA 20 turn0.1225 | ACA 20 turn | 22 | 60 | 0.1 | A | mA | 5 | 10 | kHz | 52 | 30.419 | 37.46 | 30.419
| ACA 20 turn1225 | ACA 20 turn | 22 | 60 | 1 | A | mA | 5 | 10 | kHz | 304.91 | 27.255 | 304.84 | 27.255
| ACA 20 turn10225 | ACA 20 turn | 22 | 60 | 10 | A | mA | 5 | 10 | kHz | 5698.4 | 6.4999 | 5698.4 | 6.4989
| ACA 20 turn0.01600.045 | ACA 20 turn | 60 | 220 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 24 | 2.1 | 23.534 | 2.0875
| ACA 20 turn0.1600.045 | ACA 20 turn | 60 | 220 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 37.146 | 2.0415 | 37.093 | 2.0415
| ACA 20 turn1600.045 | ACA 20 turn | 60 | 220 | 1 | A | mA | 0.045 | 0.1 | kHz | 540.32 | 0.97471 | 540.3 | 0.97417
| ACA 20 turn10600.045 | ACA 20 turn | 60 | 220 | 10 | A | mA | 0.045 | 0.1 | kHz | 5795.1 | 0.11376 | 5795.1 | 0.11344
| ACA 20 turn100600.045 | ACA 20 turn | 60 | 220 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.011493 | 58000 | 0.011365
| ACA 20 turn0.01600.1 | ACA 20 turn | 60 | 220 | 0.01 | A | mA | 0.1 | 1 | kHz | 30 | 2.3 | 29.364 | 2.2879
| ACA 20 turn0.1600.1 | ACA 20 turn | 60 | 220 | 0.1 | A | mA | 0.1 | 1 | kHz | 41.544 | 2.2475 | 41.547 | 2.2467
| ACA 20 turn1600.1 | ACA 20 turn | 60 | 220 | 1 | A | mA | 0.1 | 1 | kHz | 534.45 | 1.1508 | 534.42 | 1.1503
| ACA 20 turn10600.1 | ACA 20 turn | 60 | 220 | 10 | A | mA | 0.1 | 1 | kHz | 5794.1 | 0.13793 | 5794.1 | 0.13761
| ACA 20 turn100600.1 | ACA 20 turn | 60 | 220 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.01392 | 58000 | 0.013793
| ACA 20 turn0.01601 | ACA 20 turn | 60 | 220 | 0.01 | A | mA | 1 | 5 | kHz | 47 | 35 | 46.103 | 34.699
| ACA 20 turn0.1601 | ACA 20 turn | 60 | 220 | 0.1 | A | mA | 1 | 5 | kHz | 47 | 35 | 47.146 | 34.695
| ACA 20 turn1601 | ACA 20 turn | 60 | 220 | 1 | A | mA | 1 | 5 | kHz | 127.73 | 34.633 | 144.7 | 34.35
| ACA 20 turn10601 | ACA 20 turn | 60 | 220 | 10 | A | mA | 1 | 5 | kHz | 4885.9 | 21.537 | 4885.9 | 21.537
| ACA 20 turn100601 | ACA 20 turn | 60 | 220 | 100 | A | mA | 1 | 5 | kHz | 57864 | 2.9205 | 57864 | 2.9201
| ACA 20 turn0.012200.045 | ACA 20 turn | 220 | 410 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 84 | 2.4 | 76.978 | 2.4161
| ACA 20 turn0.12200.045 | ACA 20 turn | 220 | 410 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 84 | 2.4069 | 81.077 | 2.4099
| ACA 20 turn12200.045 | ACA 20 turn | 220 | 410 | 1 | A | mA | 0.045 | 0.1 | kHz | 407.38 | 1.9712 | 407.21 | 1.9701
| ACA 20 turn102200.045 | ACA 20 turn | 220 | 410 | 10 | A | mA | 0.045 | 0.1 | kHz | 5755.8 | 0.34673 | 5755.9 | 0.34526
| ACA 20 turn1002200.045 | ACA 20 turn | 220 | 410 | 100 | A | mA | 0.045 | 0.1 | kHz | 57996 | 0.03551 | 57996 | 0.034907
| ACA 20 turn0.012200.1 | ACA 20 turn | 220 | 410 | 0.01 | A | mA | 0.1 | 1 | kHz | 98 | 2.6 | 84.28 | 2.6326
| ACA 20 turn0.12200.1 | ACA 20 turn | 220 | 410 | 0.1 | A | mA | 0.1 | 1 | kHz | 98 | 2.6065 | 88.046 | 2.6269
| ACA 20 turn12200.1 | ACA 20 turn | 220 | 410 | 1 | A | mA | 0.1 | 1 | kHz | 396.27 | 2.2061 | 396.12 | 2.205
| ACA 20 turn102200.1 | ACA 20 turn | 220 | 410 | 10 | A | mA | 0.1 | 1 | kHz | 5747.7 | 0.41071 | 5747.8 | 0.40925
| ACA 20 turn1002200.1 | ACA 20 turn | 220 | 410 | 100 | A | mA | 0.1 | 1 | kHz | 57995 | 0.042063 | 57995 | 0.04146
| ACA 20 turn0.012201 | ACA 20 turn | 220 | 410 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.22 | 34.699
| ACA 20 turn0.12201 | ACA 20 turn | 220 | 410 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.69 | 34.698
| ACA 20 turn12201 | ACA 20 turn | 220 | 410 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 148.62 | 34.646
| ACA 20 turn102201 | ACA 20 turn | 220 | 410 | 10 | A | mA | 1 | 5 | kHz | 2974.4 | 30.479 | 2973.8 | 30.479
| ACA 20 turn1002201 | ACA 20 turn | 220 | 410 | 100 | A | mA | 1 | 5 | kHz | 57090 | 6.4831 | 57090 | 6.4816
| ACA 50 turn0.001550.01 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8 | 4.2301 | 2.8841
| ACA 50 turn0.01550.01 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.7673 | 4.3733 | 2.8834
| ACA 50 turn0.1550.01 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 0.01 | 0.045 | kHz | 22 | 2.8076 | 17.829 | 2.819
| ACA 50 turn1550.01 | ACA 50 turn | 55 | 150 | 1 | A | mA | 0.01 | 0.045 | kHz | 530.85 | 1.3025 | 530.83 | 1.3019
| ACA 50 turn10550.01 | ACA 50 turn | 55 | 150 | 10 | A | mA | 0.01 | 0.045 | kHz | 5794.1 | 0.14924 | 5794.1 | 0.14886
| ACA 50 turn100550.01 | ACA 50 turn | 55 | 150 | 100 | A | mA | 0.01 | 0.045 | kHz | 58000 | 0.01506 | 57992 | 0.01491
| ACA 50 turn0.001550.045 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
| ACA 50 turn0.01550.045 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
| ACA 50 turn0.1550.045 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 49 | 2.6 | 20.257 | 2.0293
| ACA 50 turn1550.045 | ACA 50 turn | 55 | 150 | 1 | A | mA | 0.045 | 0.1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
| ACA 50 turn10550.045 | ACA 50 turn | 55 | 150 | 10 | A | mA | 0.045 | 0.1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
| ACA 50 turn100550.045 | ACA 50 turn | 55 | 150 | 100 | A | mA | 0.045 | 0.1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
| ACA 50 turn0.001550.1 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.0611 | 2.1158
| ACA 50 turn0.01550.1 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 2.2593 | 2.1148
| ACA 50 turn0.1550.1 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 0.1 | 1 | kHz | 49 | 2.6 | 20.257 | 2.0293
| ACA 50 turn1550.1 | ACA 50 turn | 55 | 150 | 1 | A | mA | 0.1 | 1 | kHz | 551.36 | 0.73945 | 551.34 | 0.7387
| ACA 50 turn10550.1 | ACA 50 turn | 55 | 150 | 10 | A | mA | 0.1 | 1 | kHz | 5796.8 | 0.080168 | 5796.8 | 0.079794
| ACA 50 turn100550.1 | ACA 50 turn | 55 | 150 | 100 | A | mA | 0.1 | 1 | kHz | 58000 | 0.0081364 | 57996 | 0.0079862
| ACA 50 turn0.001551 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 1 | 5 | kHz | 58 | 7.2 | 55.853 | 7.2093
| ACA 50 turn0.01551 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 1 | 5 | kHz | 58 | 7.209 | 55.901 | 7.2091
| ACA 50 turn0.1551 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 1 | 5 | kHz | 60.896 | 7.1858 | 60.843 | 7.1859
| ACA 50 turn1551 | ACA 50 turn | 55 | 150 | 1 | A | mA | 1 | 5 | kHz | 422.41 | 5.6956 | 422.3 | 5.6954
| ACA 50 turn10551 | ACA 50 turn | 55 | 150 | 10 | A | mA | 1 | 5 | kHz | 5763.9 | 0.97779 | 5763.9 | 0.97715
| ACA 50 turn100551 | ACA 50 turn | 55 | 150 | 100 | A | mA | 1 | 5 | kHz | 57996 | 0.099047 | 57948 | 0.098783
| ACA 50 turn0.001555 | ACA 50 turn | 55 | 150 | 0.001 | A | mA | 5 | 10 | kHz | 290 | 29 | 288.13 | 28.936
| ACA 50 turn0.01555 | ACA 50 turn | 55 | 150 | 0.01 | A | mA | 5 | 10 | kHz | 290 | 28.936 | 288.14 | 28.936
| ACA 50 turn0.1555 | ACA 50 turn | 55 | 150 | 0.1 | A | mA | 5 | 10 | kHz | 290 | 28.931 | 289.4 | 28.93
| ACA 50 turn1555 | ACA 50 turn | 55 | 150 | 1 | A | mA | 5 | 10 | kHz | 405.34 | 28.398 | 405.23 | 28.397
| ACA 50 turn10555 | ACA 50 turn | 55 | 150 | 10 | A | mA | 5 | 10 | kHz | 5330.6 | 13.938 | 5330.7 | 13.932
| ACA 50 turn100555 | ACA 50 turn | 55 | 150 | 100 | A | mA | 5 | 10 | kHz | 57941 | 1.6237 | 57264 | 1.6205
| ACA 50 turn0.011500.045 | ACA 50 turn | 150 | 550 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 59 | 2.1 | 58.542 | 2.0879
| ACA 50 turn0.11500.045 | ACA 50 turn | 150 | 550 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 62.205 | 2.0942 | 64.156 | 2.0802
| ACA 50 turn11500.045 | ACA 50 turn | 150 | 550 | 1 | A | mA | 0.045 | 0.1 | kHz | 445.36 | 1.6257 | 445.1 | 1.6253
| ACA 50 turn101500.045 | ACA 50 turn | 150 | 550 | 10 | A | mA | 0.045 | 0.1 | kHz | 5769.8 | 0.28152 | 5769.8 | 0.28084
| ACA 50 turn1001500.045 | ACA 50 turn | 150 | 550 | 100 | A | mA | 0.045 | 0.1 | kHz | 57997 | 0.02869 | 57997 | 0.028411
| ACA 50 turn0.011500.1 | ACA 50 turn | 150 | 550 | 0.01 | A | mA | 0.1 | 1 | kHz | 74 | 2.3 | 73.171 | 2.2882
| ACA 50 turn0.11500.1 | ACA 50 turn | 150 | 550 | 0.1 | A | mA | 0.1 | 1 | kHz | 76.076 | 2.2962 | 78.162 | 2.2814
| ACA 50 turn11500.1 | ACA 50 turn | 150 | 550 | 1 | A | mA | 0.1 | 1 | kHz | 437.31 | 1.8466 | 437.03 | 1.8463
| ACA 50 turn101500.1 | ACA 50 turn | 150 | 550 | 10 | A | mA | 0.1 | 1 | kHz | 5764 | 0.34064 | 5763.9 | 0.33997
| ACA 50 turn1001500.1 | ACA 50 turn | 150 | 550 | 100 | A | mA | 0.1 | 1 | kHz | 57996 | 0.034758 | 57996 | 0.034478
| ACA 50 turn0.011501 | ACA 50 turn | 150 | 550 | 0.01 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.24 | 34.699
| ACA 50 turn0.11501 | ACA 50 turn | 150 | 550 | 0.1 | A | mA | 1 | 5 | kHz | 120 | 35 | 115.76 | 34.698
| ACA 50 turn11501 | ACA 50 turn | 150 | 550 | 1 | A | mA | 1 | 5 | kHz | 120 | 35 | 155.32 | 34.642
| ACA 50 turn101501 | ACA 50 turn | 150 | 550 | 10 | A | mA | 1 | 5 | kHz | 3301.1 | 30.465 | 3300.8 | 30.465
| ACA 50 turn1001501 | ACA 50 turn | 150 | 550 | 100 | A | mA | 1 | 5 | kHz | 57174 | 7.1298 | 57174 | 7.1293
| ACA 50 turn0.015500.045 | ACA 50 turn | 550 | 1025 | 0.01 | A | mA | 0.045 | 0.1 | kHz | 210 | 2.4 | 192.41 | 2.4161
| ACA 50 turn0.15500.045 | ACA 50 turn | 550 | 1025 | 0.1 | A | mA | 0.045 | 0.1 | kHz | 210 | 2.4 | 193.98 | 2.4152
| ACA 50 turn15500.045 | ACA 50 turn | 550 | 1025 | 1 | A | mA | 0.045 | 0.1 | kHz | 350.89 | 2.3236 | 350.74 | 2.3224
| ACA 50 turn105500.045 | ACA 50 turn | 550 | 1025 | 10 | A | mA | 0.045 | 0.1 | kHz | 5545.6 | 0.82078 | 5546.4 | 0.81772
| ACA 50 turn1005500.045 | ACA 50 turn | 550 | 1025 | 100 | A | mA | 0.045 | 0.1 | kHz | 57972 | 0.088607 | 57972 | 0.087215
| ACA 50 turn0.015500.1 | ACA 50 turn | 550 | 1025 | 0.01 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 210.67 | 2.6326
| ACA 50 turn0.15500.1 | ACA 50 turn | 550 | 1025 | 0.1 | A | mA | 0.1 | 1 | kHz | 250 | 2.6 | 212.17 | 2.6317
| ACA 50 turn15500.1 | ACA 50 turn | 550 | 1025 | 1 | A | mA | 0.1 | 1 | kHz | 357.02 | 2.5469 | 356.89 | 2.5458
| ACA 50 turn105500.1 | ACA 50 turn | 550 | 1025 | 10 | A | mA | 0.1 | 1 | kHz | 5503.5 | 0.96337 | 5504.3 | 0.96038
| ACA 50 turn1005500.1 | ACA 50 turn | 550 | 1025 | 100 | A | mA | 0.1 | 1 | kHz | 57966 | 0.10497 | 57967 | 0.10358
| ACA 50 turn0.015501 | ACA 50 turn | 550 | 1025 | 0.01 | A | mA | 1 | 5 | kHz | 290 | 35 | 288.04 | 34.699
| ACA 50 turn0.15501 | ACA 50 turn | 550 | 1025 | 0.1 | A | mA | 1 | 5 | kHz | 290 | 35 | 288.13 | 34.699
| ACA 50 turn15501 | ACA 50 turn | 550 | 1025 | 1 | A | mA | 1 | 5 | kHz | 290 | 35 | 301.82 | 34.69
| ACA 50 turn105501 | ACA 50 turn | 550 | 1025 | 10 | A | mA | 1 | 5 | kHz | 1583.2 | 33.891 | 1582.2 | 33.891
| ACA 50 turn1005501 | ACA 50 turn | 550 | 1025 | 100 | A | mA | 1 | 5 | kHz | 53001 | 14.818 | 53001 | 14.816
| Frequency0.000010.001 | Frequency | 0.001 | 0.12 | 1e-05 | kHz | Hz |  |  |  | 0.006 | 0.013 | 0.005881 | 0.012606
| Frequency0.00010.001 | Frequency | 0.001 | 0.12 | 0.0001 | kHz | Hz |  |  |  | 0.057764 | 0.0019154 | 0.057746 | 0.0014481
| Frequency0.0010.001 | Frequency | 0.001 | 0.12 | 0.001 | kHz | Hz |  |  |  | 0.57736 | 0.0003329 | 0.57736 | 0.00014506
| Frequency0.010.001 | Frequency | 0.001 | 0.12 | 0.01 | kHz | Hz |  |  |  | 1.2926 | 8.9649e-05 | 5.7736 | 1.4506e-05
| Frequency0.00010.12 | Frequency | 0.12 | 1.2 | 0.0001 | kHz | Hz |  |  |  | 0.058 | 0.0095 | 0.056807 | 0.0092702
| Frequency0.0010.12 | Frequency | 0.12 | 1.2 | 0.001 | kHz | Hz |  |  |  | 0.57742 | 0.0011566 | 0.57725 | 0.0010094
| Frequency0.010.12 | Frequency | 0.12 | 1.2 | 0.01 | kHz | Hz |  |  |  | 5.7736 | 0.00016017 | 5.7735 | 0.00010104
| Frequency0.10.12 | Frequency | 0.12 | 1.2 | 0.1 | kHz | Hz |  |  |  | 57.736 | 3.376e-05 | 57.736 | 1.0104e-05
| Frequency0.0011.2 | Frequency | 1.2 | 12 | 0.001 | kHz | Hz |  |  |  | 0.58 | 0.0091 | 0.56785 | 0.0088368
| Frequency0.011.2 | Frequency | 1.2 | 12 | 0.01 | kHz | Hz |  |  |  | 5.7742 | 0.0011047 | 5.7725 | 0.00095754
| Frequency0.11.2 | Frequency | 1.2 | 12 | 0.1 | kHz | Hz |  |  |  | 5.9215 | 0.00015498 | 57.735 | 9.584e-05
| Frequency11.2 | Frequency | 1.2 | 12 | 1 | kHz | Hz |  |  |  | 5.9216 | 3.324e-05 | 577.36 | 9.5841e-06
| Frequency0.0112 | Frequency | 12 | 120 | 0.01 | kHz | Hz |  |  |  | 19 | 0.028 | 18.255 | 0.027641
| Frequency0.112 | Frequency | 12 | 120 | 0.1 | kHz | Hz |  |  |  | 60.338 | 0.0096628 | 60.269 | 0.0091178
| Frequency112 | Frequency | 12 | 120 | 1 | kHz | Hz |  |  |  | 577.65 | 0.0011971 | 577.61 | 0.00096075
| Frequency1012 | Frequency | 12 | 120 | 10 | kHz | Hz |  |  |  | 5773.6 | 0.00019075 | 5773.6 | 9.6128e-05
| Frequency0.1120 | Frequency | 120 | 1200 | 0.1 | kHz | Hz |  |  |  | 60 | 0.016 | 59.55 | 0.015104
| Frequency1120 | Frequency | 120 | 1200 | 1 | kHz | Hz |  |  |  | 577.64 | 0.0023288 | 577.51 | 0.0018144
| Frequency10120 | Frequency | 120 | 1200 | 10 | kHz | Hz |  |  |  | 5773.6 | 0.00038884 | 5773.6 | 0.00018186
| Frequency100120 | Frequency | 120 | 1200 | 100 | kHz | Hz |  |  |  | 57736 | 0.00010098 | 57736 | 1.8187e-05
| Frequency11200 | Frequency | 1200 | 2000 | 1 | kHz | kHz |  |  |  | 0.43 | 0.13 | 0.42565 | 0.12836
| Frequency101200 | Frequency | 1200 | 2000 | 10 | kHz | kHz |  |  |  | 5.7564 | 0.016002 | 5.757 | 0.014015
| Frequency1001200 | Frequency | 1200 | 2000 | 100 | kHz | kHz |  |  |  | 57.734 | 0.0022013 | 57.734 | 0.0014029
| Frequency10001200 | Frequency | 1200 | 2000 | 1000 | kHz | kHz |  |  |  | 577.36 | 0.00045965 | 577.36 | 0.0001403
| Capacitance0.000010.00033 | Capacitance | 0.00033 | 0.0011 | 1e-05 | mF | nF |  |  |  | 5.6891 | 1394 | 5.6874 | 1390.9
| Capacitance0.00010.00033 | Capacitance | 0.00033 | 0.0011 | 0.0001 | mF | nF |  |  |  | 57.722 | 162.48 | 57.721 | 160.69
| Capacitance0.00000010.0011 | Capacitance | 0.0011 | 0.0033 | 1e-07 | mF | nF |  |  |  | 3.5 | 2900 | 3.4649 | 2886.7
| Capacitance0.0000010.0011 | Capacitance | 0.0011 | 0.0033 | 1e-06 | mF | nF |  |  |  | 3.5 | 2900 | 3.4958 | 2881.2
| Capacitance0.000010.0011 | Capacitance | 0.0011 | 0.0033 | 1e-05 | mF | nF |  |  |  | 6.0937 | 2463.2 | 6.0906 | 2462.3
| Capacitance0.00010.0011 | Capacitance | 0.0011 | 0.0033 | 0.0001 | mF | nF |  |  |  | 57.585 | 484.68 | 57.585 | 483.14
| Capacitance0.0000010.0033 | Capacitance | 0.0033 | 0.011 | 1e-06 | mF | nF |  |  |  | 12 | 2900 | 11.566 | 2886.2
| Capacitance0.000010.0033 | Capacitance | 0.0033 | 0.011 | 1e-05 | mF | nF |  |  |  | 12.429 | 2861.1 | 12.501 | 2835.6
| Capacitance0.00010.0033 | Capacitance | 0.0033 | 0.011 | 0.0001 | mF | nF |  |  |  | 56.88 | 1394.3 | 56.874 | 1390.9
| Capacitance0.0010.0033 | Capacitance | 0.0033 | 0.011 | 0.001 | mF | nF |  |  |  | 577.21 | 162.54 | 577.21 | 160.69
| Capacitance0.0000010.011 | Capacitance | 0.011 | 0.033 | 1e-06 | mF | nF |  |  |  | 36 | 4600 | 34.648 | 4618.7
| Capacitance0.000010.011 | Capacitance | 0.011 | 0.033 | 1e-05 | mF | nF |  |  |  | 36 | 4600 | 34.894 | 4613.9
| Capacitance0.00010.011 | Capacitance | 0.011 | 0.033 | 0.0001 | mF | nF |  |  |  | 56.814 | 4211.3 | 56.808 | 4211
| Capacitance0.0010.011 | Capacitance | 0.011 | 0.033 | 0.001 | mF | nF |  |  |  | 572.02 | 1057.8 | 572.02 | 1057.2
| Capacitance0.000010.033 | Capacitance | 0.033 | 0.11 | 1e-05 | mF | nF |  |  |  | 120 | 5200 | 115.61 | 5195.6
| Capacitance0.00010.033 | Capacitance | 0.033 | 0.11 | 0.0001 | mF | nF |  |  |  | 121.85 | 5183.2 | 122.72 | 5152.8
| Capacitance0.0010.033 | Capacitance | 0.033 | 0.11 | 0.001 | mF | nF |  |  |  | 536.63 | 3284.2 | 536.45 | 3282
| Capacitance0.010.033 | Capacitance | 0.033 | 0.11 | 0.01 | mF | nF |  |  |  | 5766.3 | 438.5 | 5766.3 | 436.53
| Capacitance0.000010.11 | Capacitance | 0.11 | 0.33 | 1e-05 | mF | µF |  |  |  | 0.35 | 5.2 | 0.34613 | 5.1998
| Capacitance0.00010.11 | Capacitance | 0.11 | 0.33 | 0.0001 | mF | µF |  |  |  | 0.35 | 5.2 | 0.34836 | 5.1953
| Capacitance0.0010.11 | Capacitance | 0.11 | 0.33 | 0.001 | mF | µF |  |  |  | 0.55623 | 4.8038 | 0.55608 | 4.8037
| Capacitance0.010.11 | Capacitance | 0.11 | 0.33 | 0.01 | mF | µF |  |  |  | 5.7038 | 1.2944 | 5.7038 | 1.2938
| Capacitance0.00010.33 | Capacitance | 0.33 | 1.1 | 0.0001 | mF | µF |  |  |  | 1.2 | 5.2 | 1.1558 | 5.1983
| Capacitance0.0010.33 | Capacitance | 0.33 | 1.1 | 0.001 | mF | µF |  |  |  | 1.2193 | 5.1825 | 1.2268 | 5.1555
| Capacitance0.010.33 | Capacitance | 0.33 | 1.1 | 0.01 | mF | µF |  |  |  | 5.3658 | 3.2865 | 5.3641 | 3.2842
| Capacitance0.10.33 | Capacitance | 0.33 | 1.1 | 0.1 | mF | µF |  |  |  | 57.663 | 0.43886 | 57.663 | 0.43689
| Capacitance0.00011.1 | Capacitance | 1.1 | 3.3 | 0.0001 | mF | µF |  |  |  | 3.5 | 5.2 | 3.4636 | 5.1989
| Capacitance0.0011.1 | Capacitance | 1.1 | 3.3 | 0.001 | mF | µF |  |  |  | 3.5 | 5.2 | 3.4866 | 5.1944
| Capacitance0.011.1 | Capacitance | 1.1 | 3.3 | 0.01 | mF | µF |  |  |  | 5.5761 | 4.8039 | 5.5634 | 4.8029
| Capacitance0.11.1 | Capacitance | 1.1 | 3.3 | 0.1 | mF | µF |  |  |  | 57.043 | 1.2991 | 57.038 | 1.2937
| Capacitance0.0013.3 | Capacitance | 3.3 | 11 | 0.001 | mF | µF |  |  |  | 13 | 5.2 | 12.076 | 5.1983
| Capacitance0.013.3 | Capacitance | 3.3 | 11 | 0.01 | mF | µF |  |  |  | 13 | 5.2 | 12.713 | 5.1556
| Capacitance0.13.3 | Capacitance | 3.3 | 11 | 0.1 | mF | µF |  |  |  | 53.666 | 3.2879 | 53.64 | 3.2843
| Capacitance13.3 | Capacitance | 3.3 | 11 | 1 | mF | µF |  |  |  | 576.63 | 0.43999 | 576.63 | 0.4369
| Capacitance0.00111 | Capacitance | 11 | 33 | 0.001 | mF | µF |  |  |  | 35 | 8.7 | 34.642 | 8.6619
| Capacitance0.0111 | Capacitance | 11 | 33 | 0.01 | mF | µF |  |  |  | 35 | 8.7 | 34.806 | 8.6585
| Capacitance0.111 | Capacitance | 11 | 33 | 0.1 | mF | µF |  |  |  | 50.508 | 8.3397 | 50.437 | 8.3396
| Capacitance111 | Capacitance | 11 | 33 | 1 | mF | µF |  |  |  | 557.55 | 3.1193 | 557.52 | 3.1158
| Capacitance0.0133 | Capacitance | 33 | 110 | 0.01 | mF | µF* |  |  |  | 120 | 13 | 115.63 | 12.701
| Capacitance0.133 | Capacitance | 33 | 110 | 0.1 | mF | µF* |  |  |  | 120 | 13 | 119.5 | 12.676
| Capacitance133 | Capacitance | 33 | 110 | 1 | mF | µF* |  |  |  | 430.99 | 10.809 | 430.28 | 10.808
| Capacitance1033 | Capacitance | 33 | 110 | 10 | mF | µF* |  |  |  | 5725.6 | 2.2156 | 5725.3 | 2.2101
| Frequency0.000010.001 | Frequency | 0.001 | 0.12 | 1e-05 | kHz | Hz |  |  |  | 0.006 | 0.013 | 0.0059069 | 0.012562
| Frequency0.00010.001 | Frequency | 0.001 | 0.12 | 0.0001 | kHz | Hz |  |  |  | 0.058029 | 0.0019067 | 0.058011 | 0.0014415
| Frequency0.0010.001 | Frequency | 0.001 | 0.12 | 0.001 | kHz | Hz |  |  |  | 0.58001 | 0.00033138 | 0.58 | 0.0001444
| Frequency0.010.001 | Frequency | 0.001 | 0.12 | 0.01 | kHz | Hz |  |  |  | 1.2926 | 8.924e-05 | 5.8 | 1.444e-05
| Frequency0.00010.12 | Frequency | 0.12 | 1.2 | 0.0001 | kHz | Hz |  |  |  | 0.058 | 0.0095 | 0.057075 | 0.0092344
| Frequency0.0010.12 | Frequency | 0.12 | 1.2 | 0.001 | kHz | Hz |  |  |  | 0.58007 | 0.0011513 | 0.5799 | 0.0010048
| Frequency0.010.12 | Frequency | 0.12 | 1.2 | 0.01 | kHz | Hz |  |  |  | 5.8001 | 0.00015944 | 5.8 | 0.00010057
| Frequency0.10.12 | Frequency | 0.12 | 1.2 | 0.1 | kHz | Hz |  |  |  | 58 | 3.3606e-05 | 58 | 1.0058e-05
| Frequency0.0011.2 | Frequency | 1.2 | 12 | 0.001 | kHz | Hz |  |  |  | 0.58 | 0.0091 | 0.57054 | 0.0088022
| Frequency0.011.2 | Frequency | 1.2 | 12 | 0.01 | kHz | Hz |  |  |  | 5.8007 | 0.0010997 | 5.799 | 0.00095317
| Frequency0.11.2 | Frequency | 1.2 | 12 | 0.1 | kHz | Hz |  |  |  | 5.9215 | 0.00015427 | 58 | 9.5402e-05
| Frequency11.2 | Frequency | 1.2 | 12 | 1 | kHz | Hz |  |  |  | 5.9216 | 3.3088e-05 | 580 | 9.5403e-06
| Frequency0.0112 | Frequency | 12 | 120 | 0.01 | kHz | Hz |  |  |  | 19 | 0.028 | 18.263 | 0.02763
| Frequency0.112 | Frequency | 12 | 120 | 0.1 | kHz | Hz |  |  |  | 60.591 | 0.0096232 | 60.522 | 0.0090804
| Frequency112 | Frequency | 12 | 120 | 1 | kHz | Hz |  |  |  | 580.29 | 0.0011916 | 580.26 | 0.00095637
| Frequency1012 | Frequency | 12 | 120 | 10 | kHz | Hz |  |  |  | 5800 | 0.00018988 | 5800 | 9.569e-05
| Frequency0.1120 | Frequency | 120 | 1200 | 0.1 | kHz | Hz |  |  |  | 61 | 0.016 | 59.805 | 0.015056
| Frequency1120 | Frequency | 120 | 1200 | 1 | kHz | Hz |  |  |  | 580.28 | 0.0023182 | 580.16 | 0.0018062
| Frequency10120 | Frequency | 120 | 1200 | 10 | kHz | Hz |  |  |  | 5800.1 | 0.00038706 | 5800 | 0.00018103
| Frequency100120 | Frequency | 120 | 1200 | 100 | kHz | Hz |  |  |  | 58000 | 0.00010052 | 58000 | 1.8104e-05
| Frequency11200 | Frequency | 1200 | 2000 | 1 | kHz | kHz |  |  |  | 0.44 | 0.13 | 0.42887 | 0.12787
| Frequency101200 | Frequency | 1200 | 2000 | 10 | kHz | kHz |  |  |  | 5.783 | 0.015929 | 5.7835 | 0.013952
| Frequency1001200 | Frequency | 1200 | 2000 | 100 | kHz | kHz |  |  |  | 57.998 | 0.0021913 | 57.998 | 0.0013965
| Frequency10001200 | Frequency | 1200 | 2000 | 1000 | kHz | kHz |  |  |  | 580 | 0.00045756 | 580 | 0.00013966
| TcindicatorE0.1-418 | TcindicatorE | -418 | -148 | 0.1 | °F | °F |  |  |  | 1.0462 | 0 | 1.0433 | 0
| TcindicatorE0.1-148 | TcindicatorE | -148 | -13 | 0.1 | °F | °F |  |  |  | 0.3466 | 0 | 0.34484 | 0
| TcindicatorE0.1-13 | TcindicatorE | -13 | 662 | 0.1 | °F | °F |  |  |  | 0.30437 | 0 | 0.30237 | 0
| TcindicatorE0.1662 | TcindicatorE | 662 | 1202 | 0.1 | °F | °F |  |  |  | 0.34352 | 0 | 0.34251 | 0
| TcindicatorE0.11202 | TcindicatorE | 1202 | 1832 | 0.1 | °F | °F |  |  |  | 0.44489 | 0 | 0.44411 | 0
| TcindicatorE1-418 | TcindicatorE | -418 | -148 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
| TcindicatorE1-148 | TcindicatorE | -148 | -13 | 1 | °F | °F |  |  |  | 1.2014 | 0 | 1.191 | 0
| TcindicatorE1-13 | TcindicatorE | -13 | 662 | 1 | °F | °F |  |  |  | 0.6529 | 0 | 0.64917 | 0
| TcindicatorE1662 | TcindicatorE | 662 | 1202 | 1 | °F | °F |  |  |  | 0.67089 | 0 | 0.66882 | 0
| TcindicatorE11202 | TcindicatorE | 1202 | 1832 | 1 | °F | °F |  |  |  | 0.72802 | 0 | 0.72611 | 0
| TcindicatorE0.1-250 | TcindicatorE | -250 | -100 | 0.1 | °C | °C |  |  |  | 0.58423 | 0 | 0.58258 | 0
| TcindicatorE0.1-100 | TcindicatorE | -100 | -25 | 0.1 | °C | °C |  |  |  | 0.19859 | 0 | 0.19765 | 0
| TcindicatorE0.1-25 | TcindicatorE | -25 | 350 | 0.1 | °C | °C |  |  |  | 0.17588 | 0 | 0.17482 | 0
| TcindicatorE0.1350 | TcindicatorE | 350 | 650 | 0.1 | °C | °C |  |  |  | 0.19683 | 0 | 0.19628 | 0
| TcindicatorE0.1650 | TcindicatorE | 650 | 1000 | 0.1 | °C | °C |  |  |  | 0.25181 | 0 | 0.25139 | 0
| TcindicatorE1-250 | TcindicatorE | -250 | -100 | 1 | °C | °C |  |  |  | 0.82286 | 0 | 0.81817 | 0
| TcindicatorE1-100 | TcindicatorE | -100 | -25 | 1 | °C | °C |  |  |  | 0.60874 | 0 | 0.60751 | 0
| TcindicatorE1-25 | TcindicatorE | -25 | 350 | 1 | °C | °C |  |  |  | 0.60171 | 0 | 0.60047 | 0
| TcindicatorE1350 | TcindicatorE | 350 | 650 | 1 | °C | °C |  |  |  | 0.60777 | 0 | 0.60706 | 0
| TcindicatorE1650 | TcindicatorE | 650 | 1000 | 1 | °C | °C |  |  |  | 0.62773 | 0 | 0.62705 | 0
| TcindicatorJ0.1-346 | TcindicatorJ | -346 | -148 | 0.1 | °F | °F |  |  |  | 0.57727 | 0 | 0.56758 | 0
| TcindicatorJ0.1-148 | TcindicatorJ | -148 | -22 | 0.1 | °F | °F |  |  |  | 0.35408 | 0 | 0.34263 | 0
| TcindicatorJ0.1-22 | TcindicatorJ | -22 | 302 | 0.1 | °F | °F |  |  |  | 0.31531 | 0 | 0.3024 | 0
| TcindicatorJ0.1302 | TcindicatorJ | 302 | 1400 | 0.1 | °F | °F |  |  |  | 0.36584 | 0 | 0.36275 | 0
| TcindicatorJ0.11400 | TcindicatorJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 0.48736 | 0 | 0.48504 | 0
| TcindicatorJ1-346 | TcindicatorJ | -346 | -148 | 1 | °F | °F |  |  |  | 0.83458 | 0 | 0.80755 | 0
| TcindicatorJ1-148 | TcindicatorJ | -148 | -22 | 1 | °F | °F |  |  |  | 0.69231 | 0 | 0.66888 | 0
| TcindicatorJ1-22 | TcindicatorJ | -22 | 302 | 1 | °F | °F |  |  |  | 0.67331 | 0 | 0.64919 | 0
| TcindicatorJ1302 | TcindicatorJ | 302 | 1400 | 1 | °F | °F |  |  |  | 0.686 | 0 | 0.6794 | 0
| TcindicatorJ11400 | TcindicatorJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 0.75781 | 0 | 0.75184 | 0
| TcindicatorJ0.1-210 | TcindicatorJ | -210 | -100 | 0.1 | °C | °C |  |  |  | 0.32482 | 0 | 0.31951 | 0
| TcindicatorJ0.1-100 | TcindicatorJ | -100 | -30 | 0.1 | °C | °C |  |  |  | 0.20266 | 0 | 0.19649 | 0
| TcindicatorJ0.1-30 | TcindicatorJ | -30 | 150 | 0.1 | °C | °C |  |  |  | 0.18177 | 0 | 0.17487 | 0
| TcindicatorJ0.1150 | TcindicatorJ | 150 | 760 | 0.1 | °C | °C |  |  |  | 0.2089 | 0 | 0.20722 | 0
| TcindicatorJ0.1760 | TcindicatorJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 0.27503 | 0 | 0.27376 | 0
| TcindicatorJ1-210 | TcindicatorJ | -210 | -100 | 1 | °C | °C |  |  |  | 0.66767 | 0 | 0.65733 | 0
| TcindicatorJ1-100 | TcindicatorJ | -100 | -30 | 1 | °C | °C |  |  |  | 0.61519 | 0 | 0.60713 | 0
| TcindicatorJ1-30 | TcindicatorJ | -30 | 150 | 1 | °C | °C |  |  |  | 0.60863 | 0 | 0.60048 | 0
| TcindicatorJ1150 | TcindicatorJ | 150 | 760 | 1 | °C | °C |  |  |  | 0.61296 | 0 | 0.61069 | 0
| TcindicatorJ1760 | TcindicatorJ | 760 | 1200 | 1 | °C | °C |  |  |  | 0.63854 | 0 | 0.63635 | 0
| TcindicatorK0.1-328 | TcindicatorK | -328 | -148 | 0.1 | °F | °F |  |  |  | 0.70186 | 0 | 0.69128 | 0
| TcindicatorK0.1-148 | TcindicatorK | -148 | -13 | 0.1 | °F | °F |  |  |  | 0.40198 | 0 | 0.38321 | 0
| TcindicatorK0.1-13 | TcindicatorK | -13 | 248 | 0.1 | °F | °F |  |  |  | 0.35702 | 0 | 0.34265 | 0
| TcindicatorK0.1248 | TcindicatorK | 248 | 1832 | 0.1 | °F | °F |  |  |  | 0.55155 | 0 | 0.54665 | 0
| TcindicatorK0.11832 | TcindicatorK | 1832 | 2500 | 0.1 | °F | °F |  |  |  | 0.83993 | 0 | 0.83548 | 0
| TcindicatorK1-328 | TcindicatorK | -328 | -148 | 1 | °F | °F |  |  |  | 0.93103 | 0 | 0.89881 | 0
| TcindicatorK1-148 | TcindicatorK | -148 | -13 | 1 | °F | °F |  |  |  | 0.73198 | 0 | 0.69054 | 0
| TcindicatorK1-13 | TcindicatorK | -13 | 248 | 1 | °F | °F |  |  |  | 0.69831 | 0 | 0.66889 | 0
| TcindicatorK1248 | TcindicatorK | 248 | 1832 | 1 | °F | °F |  |  |  | 0.80643 | 0 | 0.79299 | 0
| TcindicatorK11832 | TcindicatorK | 1832 | 2500 | 1 | °F | °F |  |  |  | 1.0285 | 0 | 1.0139 | 0
| TcindicatorK0.1-200 | TcindicatorK | -200 | -100 | 0.1 | °C | °C |  |  |  | 0.39349 | 0 | 0.38767 | 0
| TcindicatorK0.1-100 | TcindicatorK | -100 | -25 | 0.1 | °C | °C |  |  |  | 0.22869 | 0 | 0.21853 | 0
| TcindicatorK0.1-25 | TcindicatorK | -25 | 120 | 0.1 | °C | °C |  |  |  | 0.20428 | 0 | 0.19653 | 0
| TcindicatorK0.1120 | TcindicatorK | 120 | 1000 | 0.1 | °C | °C |  |  |  | 0.31026 | 0 | 0.30758 | 0
| TcindicatorK0.11000 | TcindicatorK | 1000 | 1372 | 0.1 | °C | °C |  |  |  | 0.46918 | 0 | 0.46672 | 0
| TcindicatorK1-200 | TcindicatorK | -200 | -100 | 1 | °C | °C |  |  |  | 0.70603 | 0 | 0.69303 | 0
| TcindicatorK1-100 | TcindicatorK | -100 | -25 | 1 | °C | °C |  |  |  | 0.62924 | 0 | 0.61462 | 0
| TcindicatorK1-25 | TcindicatorK | -25 | 120 | 1 | °C | °C |  |  |  | 0.61728 | 0 | 0.60715 | 0
| TcindicatorK1120 | TcindicatorK | 120 | 1000 | 1 | °C | °C |  |  |  | 0.65669 | 0 | 0.65162 | 0
| TcindicatorK11000 | TcindicatorK | 1000 | 1372 | 1 | °C | °C |  |  |  | 0.74634 | 0 | 0.74015 | 0
| TcindicatorN0.1-328 | TcindicatorN | -328 | -148 | 0.1 | °F | °F |  |  |  | 0.84492 | 0 | 0.83553 | 0
| TcindicatorN0.1-148 | TcindicatorN | -148 | -13 | 0.1 | °F | °F |  |  |  | 0.47674 | 0 | 0.46496 | 0
| TcindicatorN0.1-13 | TcindicatorN | -13 | 248 | 0.1 | °F | °F |  |  |  | 0.41715 | 0 | 0.40364 | 0
| TcindicatorN0.1248 | TcindicatorN | 248 | 770 | 0.1 | °F | °F |  |  |  | 0.3915 | 0 | 0.38322 | 0
| TcindicatorN0.1770 | TcindicatorN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 0.57285 | 0 | 0.56722 | 0
| TcindicatorN1-328 | TcindicatorN | -328 | -148 | 1 | °F | °F |  |  |  | 1.0446 | 0 | 1.014 | 0
| TcindicatorN1-148 | TcindicatorN | -148 | -13 | 1 | °F | °F |  |  |  | 0.76848 | 0 | 0.73904 | 0
| TcindicatorN1-13 | TcindicatorN | -13 | 248 | 1 | °F | °F |  |  |  | 0.73301 | 0 | 0.70208 | 0
| TcindicatorN1248 | TcindicatorN | 248 | 770 | 1 | °F | °F |  |  |  | 0.70888 | 0 | 0.69055 | 0
| TcindicatorN1770 | TcindicatorN | 770 | 2372 | 1 | °F | °F |  |  |  | 0.82304 | 0 | 0.8073 | 0
| TcindicatorN0.1-200 | TcindicatorN | -200 | -100 | 0.1 | °C | °C |  |  |  | 0.47201 | 0 | 0.46682 | 0
| TcindicatorN0.1-100 | TcindicatorN | -100 | -25 | 0.1 | °C | °C |  |  |  | 0.26968 | 0 | 0.26325 | 0
| TcindicatorN0.1-25 | TcindicatorN | -25 | 120 | 0.1 | °C | °C |  |  |  | 0.23706 | 0 | 0.22973 | 0
| TcindicatorN0.1120 | TcindicatorN | 120 | 410 | 0.1 | °C | °C |  |  |  | 0.22303 | 0 | 0.21855 | 0
| TcindicatorN0.1410 | TcindicatorN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 0.32197 | 0 | 0.31888 | 0
| TcindicatorN1-200 | TcindicatorN | -200 | -100 | 1 | °C | °C |  |  |  | 0.75326 | 0 | 0.74022 | 0
| TcindicatorN1-100 | TcindicatorN | -100 | -25 | 1 | °C | °C |  |  |  | 0.64265 | 0 | 0.6319 | 0
| TcindicatorN1-25 | TcindicatorN | -25 | 120 | 1 | °C | °C |  |  |  | 0.62966 | 0 | 0.61869 | 0
| TcindicatorN1120 | TcindicatorN | 120 | 410 | 1 | °C | °C |  |  |  | 0.62103 | 0 | 0.61463 | 0
| TcindicatorN1410 | TcindicatorN | 410 | 1300 | 1 | °C | °C |  |  |  | 0.66303 | 0 | 0.65703 | 0
| TcindicatorR0.132 | TcindicatorR | 32 | 482 | 0.1 | °F | °F |  |  |  | 1.1966 | 0 | 1.19 | 0
| TcindicatorR0.1482 | TcindicatorR | 482 | 752 | 0.1 | °F | °F |  |  |  | 0.74217 | 0 | 0.73324 | 0
| TcindicatorR0.1752 | TcindicatorR | 752 | 1832 | 0.1 | °F | °F |  |  |  | 0.70146 | 0 | 0.69201 | 0
| TcindicatorR0.11832 | TcindicatorR | 1832 | 3213 | 0.1 | °F | °F |  |  |  | 0.96695 | 0 | 0.96011 | 0
| TcindicatorR132 | TcindicatorR | 32 | 482 | 1 | °F | °F |  |  |  | 1.3448 | 0 | 1.3214 | 0
| TcindicatorR1482 | TcindicatorR | 482 | 752 | 1 | °F | °F |  |  |  | 0.95934 | 0 | 0.93147 | 0
| TcindicatorR1752 | TcindicatorR | 752 | 1832 | 1 | °F | °F |  |  |  | 0.92821 | 0 | 0.89938 | 0
| TcindicatorR11832 | TcindicatorR | 1832 | 3213 | 1 | °F | °F |  |  |  | 1.1422 | 0 | 1.1188 | 0
| TcindicatorR0.10 | TcindicatorR | 0 | 250 | 0.1 | °C | °C |  |  |  | 0.66953 | 0 | 0.66592 | 0
| TcindicatorR0.1250 | TcindicatorR | 250 | 400 | 0.1 | °C | °C |  |  |  | 0.41654 | 0 | 0.41163 | 0
| TcindicatorR0.1400 | TcindicatorR | 400 | 1000 | 0.1 | °C | °C |  |  |  | 0.39417 | 0 | 0.38897 | 0
| TcindicatorR0.11000 | TcindicatorR | 1000 | 1767 | 0.1 | °C | °C |  |  |  | 0.53998 | 0 | 0.5362 | 0
| TcindicatorR10 | TcindicatorR | 0 | 250 | 1 | °C | °C |  |  |  | 0.89034 | 0 | 0.87946 | 0
| TcindicatorR1250 | TcindicatorR | 250 | 400 | 1 | °C | °C |  |  |  | 0.71813 | 0 | 0.70671 | 0
| TcindicatorR1400 | TcindicatorR | 400 | 1000 | 1 | °C | °C |  |  |  | 0.70538 | 0 | 0.69376 | 0
| TcindicatorR11000 | TcindicatorR | 1000 | 1767 | 1 | °C | °C |  |  |  | 0.7961 | 0 | 0.78582 | 0
| TcindicatorS0.132 | TcindicatorS | 32 | 482 | 0.1 | °F | °F |  |  |  | 0.99102 | 0 | 0.98329 | 0
| TcindicatorS0.1482 | TcindicatorS | 482 | 1832 | 0.1 | °F | °F |  |  |  | 0.76311 | 0 | 0.75443 | 0
| TcindicatorS0.11832 | TcindicatorS | 1832 | 2552 | 0.1 | °F | °F |  |  |  | 0.78274 | 0 | 0.77428 | 0
| TcindicatorS0.12552 | TcindicatorS | 2552 | 3213 | 0.1 | °F | °F |  |  |  | 0.96728 | 0 | 0.96044 | 0
| TcindicatorS132 | TcindicatorS | 32 | 482 | 1 | °F | °F |  |  |  | 1.1653 | 0 | 1.1388 | 0
| TcindicatorS1482 | TcindicatorS | 482 | 1832 | 1 | °F | °F |  |  |  | 0.97706 | 0 | 0.94824 | 0
| TcindicatorS11832 | TcindicatorS | 1832 | 2552 | 1 | °F | °F |  |  |  | 0.99106 | 0 | 0.96411 | 0
| TcindicatorS12552 | TcindicatorS | 2552 | 3213 | 1 | °F | °F |  |  |  | 1.1424 | 0 | 1.1191 | 0
| TcindicatorS0.10 | TcindicatorS | 0 | 250 | 0.1 | °C | °C |  |  |  | 0.55632 | 0 | 0.55207 | 0
| TcindicatorS0.1250 | TcindicatorS | 250 | 1000 | 0.1 | °C | °C |  |  |  | 0.42876 | 0 | 0.42399 | 0
| TcindicatorS0.11000 | TcindicatorS | 1000 | 1400 | 0.1 | °C | °C |  |  |  | 0.4386 | 0 | 0.43394 | 0
| TcindicatorS0.11400 | TcindicatorS | 1400 | 1767 | 0.1 | °C | °C |  |  |  | 0.54057 | 0 | 0.53679 | 0
| TcindicatorS10 | TcindicatorS | 0 | 250 | 1 | °C | °C |  |  |  | 0.80847 | 0 | 0.79673 | 0
| TcindicatorS1250 | TcindicatorS | 250 | 1000 | 1 | °C | °C |  |  |  | 0.72528 | 0 | 0.71398 | 0
| TcindicatorS11000 | TcindicatorS | 1000 | 1400 | 1 | °C | °C |  |  |  | 0.73114 | 0 | 0.71993 | 0
| TcindicatorS11400 | TcindicatorS | 1400 | 1767 | 1 | °C | °C |  |  |  | 0.7965 | 0 | 0.78622 | 0
| TcindicatorT0.1-418 | TcindicatorT | -418 | -238 | 0.1 | °F | °F |  |  |  | 0.99115 | 0 | 0.98245 | 0
| TcindicatorT0.1-238 | TcindicatorT | -238 | 32 | 0.1 | °F | °F |  |  |  | 0.51969 | 0 | 0.50583 | 0
| TcindicatorT0.132 | TcindicatorT | 32 | 248 | 0.1 | °F | °F |  |  |  | 0.31004 | 0 | 0.3024 | 0
| TcindicatorT0.1248 | TcindicatorT | 248 | 752 | 0.1 | °F | °F |  |  |  | 0.31004 | 0 | 0.3024 | 0
| TcindicatorT1-418 | TcindicatorT | -418 | -238 | 1 | °F | °F |  |  |  | 1.1678 | 0 | 1.1381 | 0
| TcindicatorT1-238 | TcindicatorT | -238 | 32 | 1 | °F | °F |  |  |  | 0.8017 | 0 | 0.76542 | 0
| TcindicatorT132 | TcindicatorT | 32 | 248 | 1 | °F | °F |  |  |  | 0.68273 | 0 | 0.66888 | 0
| TcindicatorT1248 | TcindicatorT | 248 | 752 | 1 | °F | °F |  |  |  | 0.66345 | 0 | 0.64919 | 0
| TcindicatorT0.1-250 | TcindicatorT | -250 | -150 | 0.1 | °C | °C |  |  |  | 0.55536 | 0 | 0.55058 | 0
| TcindicatorT0.1-150 | TcindicatorT | -150 | 0 | 0.1 | °C | °C |  |  |  | 0.29308 | 0 | 0.28549 | 0
| TcindicatorT0.10 | TcindicatorT | 0 | 120 | 0.1 | °C | °C |  |  |  | 0.20015 | 0 | 0.19651 | 0
| TcindicatorT0.1120 | TcindicatorT | 120 | 400 | 0.1 | °C | °C |  |  |  | 0.17895 | 0 | 0.17487 | 0
| TcindicatorT1-250 | TcindicatorT | -250 | -150 | 1 | °C | °C |  |  |  | 0.8089 | 0 | 0.7957 | 0
| TcindicatorT1-150 | TcindicatorT | -150 | 0 | 1 | °C | °C |  |  |  | 0.65502 | 0 | 0.64149 | 0
| TcindicatorT10 | TcindicatorT | 0 | 120 | 1 | °C | °C |  |  |  | 0.61188 | 0 | 0.60714 | 0
| TcindicatorT1120 | TcindicatorT | 120 | 400 | 1 | °C | °C |  |  |  | 0.60528 | 0 | 0.60048 | 0

## Standard2
Updated: | 2021-02-23 | Transmille
| Main Name | Item | Range1 Low | Range1 High | Res | Meas Units | Uncert Units | Range 2 Low | Range2 High | Range2 Units | Base | Mult | Base 2 | Mult 2
| DCV0.0000010 | DCV | 0 | 0.1 | 1e-06 | V | µV |  |  |  | 12 | 92 | 11.624 | 92.068
| DCV0.000010 | DCV | 0 | 0.1 | 1e-05 | V | µV |  |  |  | 12.963 | 86.429 | 12.956 | 86.458
| DCV0.00010 | DCV | 0 | 0.1 | 0.0001 | V | µV |  |  |  | 58.889 | 24.899 | 58.883 | 24.902
| DCV0.0010 | DCV | 0 | 0.1 | 0.001 | V | µV |  |  |  | 577.47 | 2.5928 | 577.47 | 2.5928
| DCV0.010 | DCV | 0 | 0.1 | 0.01 | V | µV |  |  |  | 5773.6 | 0.25939 | 5773.6 | 0.25939
| DCV0.000010.1 | DCV | 0.1 | 1 | 1e-05 | V | µV |  |  |  | 36 | 92 | 35.885 | 91.648
| DCV0.00010.1 | DCV | 0.1 | 1 | 0.0001 | V | µV |  |  |  | 65.685 | 74.286 | 65.574 | 74.301
| DCV0.0010.1 | DCV | 0.1 | 1 | 0.001 | V | µV |  |  |  | 577.78 | 13.555 | 577.73 | 13.515
| DCV0.010.1 | DCV | 0.1 | 1 | 0.01 | V | µV |  |  |  | 5773.6 | 1.3866 | 5773.6 | 1.3696
| DCV0.10.1 | DCV | 0.1 | 1 | 0.1 | V | µV |  |  |  | 57736 | 0.14379 | 57736 | 0.13698
| DCV0.00011 | DCV | 1 | 10 | 0.0001 | V | µV |  |  |  | 360 | 92 | 354.54 | 91.829
| DCV0.0011 | DCV | 1 | 10 | 0.001 | V | µV |  |  |  | 653.72 | 74.371 | 653.17 | 74.329
| DCV0.011 | DCV | 1 | 10 | 0.01 | V | µV |  |  |  | 5777.2 | 13.559 | 5777 | 13.491
| DCV0.11 | DCV | 1 | 10 | 0.1 | V | µV |  |  |  | 57736 | 1.3954 | 57736 | 13.496
| DCV11 | DCV | 1 | 10 | 1 | V | µV |  |  |  | 577360 | 0.14808 | 577360 | 13.497
| DCV0.00110 | DCV | 10 | 100 | 0.001 | V | mV |  |  |  | 6.8 | 0.082 | 6.7338 | 0.082465
| DCV0.0110 | DCV | 10 | 100 | 0.01 | V | mV |  |  |  | 8.774 | 0.072816 | 8.766 | 0.072781
| DCV0.110 | DCV | 10 | 100 | 0.1 | V | mV |  |  |  | 58.072 | 0.015845 | 58.068 | 0.015769
| DCV110 | DCV | 10 | 100 | 1 | V | mV |  |  |  | 577.39 | 0.0016415 | 577.39 | 0.0016093
| DCV1010 | DCV | 10 | 100 | 10 | V | mV |  |  |  | 5773.6 | 0.00017383 | 5773.6 | 0.00016097
| DCV0.01100 | DCV | 100 | 1020 | 0.01 | V | mV |  |  |  | 36 | 0.092 | 35.843 | 0.091772
| DCV0.1100 | DCV | 100 | 1020 | 0.1 | V | mV |  |  |  | 65.632 | 0.07463 | 65.525 | 0.074608
| DCV1100 | DCV | 100 | 1020 | 1 | V | mV |  |  |  | 577.76 | 0.013752 | 577.71 | 0.013677
| DCV10100 | DCV | 100 | 1020 | 10 | V | mV |  |  |  | 5773.6 | 0.0014179 | 5773.6 | 0.0013864
| DCV100100 | DCV | 100 | 1020 | 100 | V | mV |  |  |  | 57736 | 0.00015125 | 57736 | 0.00013866
| DCA0.0000000010 | DCA | 0 | 0.0001 | 1e-09 | A | nA |  |  |  | 35 | 350000 | 34.933 | 349170
| DCA0.000000010 | DCA | 0 | 0.0001 | 1e-08 | A | nA |  |  |  | 35.368 | 346810 | 35.366 | 346810
| DCA0.00000010 | DCA | 0 | 0.0001 | 1e-07 | A | nA |  |  |  | 67.34 | 230750 | 67.336 | 230720
| DCA0.0000010 | DCA | 0 | 0.0001 | 1e-06 | A | nA |  |  |  | 578.4 | 31405 | 578.39 | 31376
| DCA0.000010 | DCA | 0 | 0.0001 | 1e-05 | A | nA |  |  |  | 5773.7 | 3163.2 | 5773.7 | 3151.6
| DCA0.000000010.0001 | DCA | 0.0001 | 0.001 | 1e-08 | A | nA |  |  |  | 120 | 350000 | 115.86 | 347000
| DCA0.00000010.0001 | DCA | 0.0001 | 0.001 | 1e-07 | A | nA |  |  |  | 127.29 | 339170 | 127.24 | 339180
| DCA0.0000010.0001 | DCA | 0.0001 | 0.001 | 1e-06 | A | nA |  |  |  | 580.76 | 159300 | 580.72 | 159250
| DCA0.000010.0001 | DCA | 0.0001 | 0.001 | 1e-05 | A | nA |  |  |  | 5773.7 | 18430 | 5773.7 | 18401
| DCA0.00010.0001 | DCA | 0.0001 | 0.001 | 0.0001 | A | nA |  |  |  | 57736 | 1854.7 | 57736 | 1843.4
| DCA0.00000010.001 | DCA | 0.001 | 0.01 | 1e-07 | A | µA |  |  |  | 1.2 | 350 | 1.1597 | 346.89
| DCA0.0000010.001 | DCA | 0.001 | 0.01 | 1e-06 | A | µA |  |  |  | 1.2739 | 339.07 | 1.2734 | 339.08
| DCA0.000010.001 | DCA | 0.001 | 0.01 | 1e-05 | A | µA |  |  |  | 5.8079 | 159.28 | 5.8074 | 159.23
| DCA0.00010.001 | DCA | 0.001 | 0.01 | 0.0001 | A | µA |  |  |  | 57.737 | 18.431 | 57.737 | 18.399
| DCA0.0010.001 | DCA | 0.001 | 0.01 | 0.001 | A | µA |  |  |  | 577.36 | 1.856 | 577.36 | 1.8431
| DCA0.0000010.01 | DCA | 0.01 | 0.1 | 1e-06 | A | µA |  |  |  | 12 | 350 | 11.589 | 348.45
| DCA0.000010.01 | DCA | 0.01 | 0.1 | 1e-05 | A | µA |  |  |  | 12.73 | 340.62 | 12.725 | 340.63
| DCA0.00010.01 | DCA | 0.01 | 0.1 | 0.0001 | A | µA |  |  |  | 58.071 | 160.28 | 58.066 | 160.23
| DCA0.0010.01 | DCA | 0.01 | 0.1 | 0.001 | A | µA |  |  |  | 577.37 | 18.56 | 577.37 | 18.528
| DCA0.010.01 | DCA | 0.01 | 0.1 | 0.01 | A | µA |  |  |  | 5773.6 | 1.8689 | 5773.6 | 1.856
| DCA0.000010.1 | DCA | 0.1 | 1 | 1e-05 | A | µA |  |  |  | 180 | 370 | 172.98 | 365.23
| DCA0.00010.1 | DCA | 0.1 | 1 | 0.0001 | A | µA |  |  |  | 181.28 | 360.03 | 181.24 | 360.03
| DCA0.0010.1 | DCA | 0.1 | 1 | 0.001 | A | µA |  |  |  | 594.75 | 194.63 | 594.71 | 194.58
| DCA0.010.1 | DCA | 0.1 | 1 | 0.01 | A | µA |  |  |  | 5775 | 23.623 | 5775 | 23.591
| DCA0.10.1 | DCA | 0.1 | 1 | 0.1 | A | µA |  |  |  | 57736 | 2.3778 | 57736 | 2.3649
| DCA0.00011 | DCA | 1 | 10.2 | 0.0001 | A | µA |  |  |  | 2.4 | 0.59 | 2.3095 | 0.57975
| DCA0.0011 | DCA | 1 | 10.2 | 0.001 | A | µA |  |  |  | 2.4 | 0.63081 | 2.3701 | 0.57578
| DCA0.011 | DCA | 1 | 10.2 | 0.01 | A | µA |  |  |  | 5.8935 | 0.56293 | 6.0655 | 0.39037
| DCA0.11 | DCA | 1 | 10.2 | 0.1 | A | µA |  |  |  | 57.625 | 0.18253 | 57.752 | 0.055477
| DCA11 | DCA | 1 | 10.2 | 1 | A | µA |  |  |  | 577.31 | 0.057402 | 577.36 | 0.0055789
| DCA 10 turn0.000110 | DCA 10 turn | 10 | 102 | 0.0001 | A | mA |  |  |  | 16 | 2 | 15.052 | 2.0088
| DCA 10 turn0.00110 | DCA 10 turn | 10 | 102 | 0.001 | A | mA |  |  |  | 16 | 2.0045 | 15.566 | 2.0045
| DCA 10 turn0.0110 | DCA 10 turn | 10 | 102 | 0.01 | A | mA |  |  |  | 50.243 | 1.7371 | 50.216 | 1.7371
| DCA 10 turn0.110 | DCA 10 turn | 10 | 102 | 0.1 | A | mA |  |  |  | 574.15 | 0.42863 | 574.14 | 0.42836
| DCA 10 turn110 | DCA 10 turn | 10 | 102 | 1 | A | mA |  |  |  | 5773.2 | 0.044479 | 5773.2 | 0.044361
| DCA 10 turn1010 | DCA 10 turn | 10 | 102 | 10 | A | mA |  |  |  | 57735 | 0.0044851 | 57735 | 0.0044377
| DCA 20 turn0.00120 | DCA 20 turn | 20 | 204 | 0.001 | A | mA |  |  |  | 32 | 2 | 30.087 | 2.0089
| DCA 20 turn0.0120 | DCA 20 turn | 20 | 204 | 0.01 | A | mA |  |  |  | 32 | 2.0077 | 30.344 | 2.0078
| DCA 20 turn0.120 | DCA 20 turn | 20 | 204 | 0.1 | A | mA |  |  |  | 52.695 | 1.9168 | 52.601 | 1.917
| DCA 20 turn120 | DCA 20 turn | 20 | 204 | 1 | A | mA |  |  |  | 565.98 | 0.78427 | 565.94 | 0.78385
| DCA 20 turn1020 | DCA 20 turn | 20 | 204 | 10 | A | mA |  |  |  | 5772.2 | 0.088858 | 5772.2 | 0.088623
| DCA 50 turn0.0015 | DCA 50 turn | 5 | 51 | 0.001 | A | mA |  |  |  | 4.7 | 2 | 4.5993 | 1.98
| DCA 50 turn0.015 | DCA 50 turn | 5 | 51 | 0.01 | A | mA |  |  |  | 5.9279 | 1.9581 | 5.8001 | 1.9595
| DCA 50 turn0.15 | DCA 50 turn | 5 | 51 | 0.1 | A | mA |  |  |  | 53.039 | 1.3233 | 52.916 | 1.3219
| DCA 50 turn15 | DCA 50 turn | 5 | 51 | 1 | A | mA |  |  |  | 576.56 | 0.20649 | 576.52 | 0.20417
| DCA 50 turn105 | DCA 50 turn | 5 | 51 | 10 | A | mA |  |  |  | 5773.5 | 0.021536 | 5773.5 | 0.020588
| DCA 50 turn0.00150 | DCA 50 turn | 50 | 510 | 0.001 | A | mA |  |  |  | 80 | 2 | 75.208 | 2.0089
| DCA 50 turn0.0150 | DCA 50 turn | 50 | 510 | 0.01 | A | mA |  |  |  | 80 | 2.0086 | 75.318 | 2.0087
| DCA 50 turn0.150 | DCA 50 turn | 50 | 510 | 0.1 | A | mA |  |  |  | 85.569 | 1.9918 | 85.292 | 1.9921
| DCA 50 turn150 | DCA 50 turn | 50 | 510 | 1 | A | mA |  |  |  | 534.37 | 1.3886 | 534.07 | 1.3883
| DCA 50 turn1050 | DCA 50 turn | 50 | 510 | 10 | A | mA |  |  |  | 5765.3 | 0.22044 | 5765.2 | 0.21986
| ACV0.00000100.01 | ACV | 0 | 0.1 | 1e-06 | V | µV | 0.01 | 2 | KHz | 35 | 940 | 34.74 | 937.4
| ACV0.0000100.01 | ACV | 0 | 0.1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 35.162 | 934.08 | 35.157 | 934.05
| ACV0.000100.01 | ACV | 0 | 0.1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 67.116 | 737.26 | 67.105 | 737.07
| ACV0.00100.01 | ACV | 0 | 0.1 | 0.001 | V | µV | 0.01 | 2 | KHz | 578.33 | 131.69 | 578.32 | 131.46
| ACV0.0100.01 | ACV | 0 | 0.1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5773.7 | 13.415 | 5773.6 | 13.317
| ACV0.00000102 | ACV | 0 | 0.1 | 1e-06 | V | µV | 2 | 20 | KHz | 36 | 930 | 35.471 | 926.49
| ACV0.0000102 | ACV | 0 | 0.1 | 1e-05 | V | µV | 2 | 20 | KHz | 36 | 923.18 | 35.865 | 923.15
| ACV0.000102 | ACV | 0 | 0.1 | 0.0001 | V | µV | 2 | 20 | KHz | 67.126 | 727.32 | 67.116 | 727.12
| ACV0.00102 | ACV | 0 | 0.1 | 0.001 | V | µV | 2 | 20 | KHz | 578.33 | 129.33 | 578.32 | 129.1
| ACV0.0102 | ACV | 0 | 0.1 | 0.01 | V | µV | 2 | 20 | KHz | 5773.7 | 13.173 | 5773.6 | 13.075
| ACV0.000010.10.01 | ACV | 0.1 | 1 | 1e-05 | V | µV | 0.01 | 2 | KHz | 350 | 940 | 347.15 | 937.18
| ACV0.00010.10.01 | ACV | 0.1 | 1 | 0.0001 | V | µV | 0.01 | 2 | KHz | 351.19 | 934.51 | 351.16 | 934.46
| ACV0.0010.10.01 | ACV | 0.1 | 1 | 0.001 | V | µV | 0.01 | 2 | KHz | 650.74 | 757.68 | 650.66 | 757.46
| ACV0.010.10.01 | ACV | 0.1 | 1 | 0.01 | V | µV | 0.01 | 2 | KHz | 5776.6 | 138.39 | 5776.5 | 138.13
| ACV0.10.10.01 | ACV | 0.1 | 1 | 0.1 | V | µV | 0.01 | 2 | KHz | 57736 | 14.108 | 57736 | 14
| ACV0.000010.12 | ACV | 0.1 | 1 | 1e-05 | V | µV | 2 | 20 | KHz | 850 | 1700 | 808.55 | 1738.9
| ACV0.00010.12 | ACV | 0.1 | 1 | 0.0001 | V | µV | 2 | 20 | KHz | 850 | 1737.8 | 810.33 | 1737.8
| ACV0.0010.12 | ACV | 0.1 | 1 | 0.001 | V | µV | 2 | 20 | KHz | 976.25 | 1636.3 | 975.89 | 1636.2
| ACV0.010.12 | ACV | 0.1 | 1 | 0.01 | V | µV | 2 | 20 | KHz | 5806.3 | 505.05 | 5806.1 | 504.49
| ACV0.10.12 | ACV | 0.1 | 1 | 0.1 | V | µV | 2 | 20 | KHz | 57739 | 53.381 | 57739 | 53.128
| ACV0.000110.01 | ACV | 1 | 10 | 0.0001 | V | mV | 0.01 | 2 | KHz | 3.5 | 0.94 | 3.4755 | 0.93718
| ACV0.00110.01 | ACV | 1 | 10 | 0.001 | V | mV | 0.01 | 2 | KHz | 3.5127 | 0.93465 | 3.5116 | 0.93446
| ACV0.0110.01 | ACV | 1 | 10 | 0.01 | V | mV | 0.01 | 2 | KHz | 6.5091 | 0.7583 | 6.5066 | 0.75746
| ACV0.110.01 | ACV | 1 | 10 | 0.1 | V | mV | 0.01 | 2 | KHz | 57.766 | 0.1391 | 57.765 | 0.13813
| ACV110.01 | ACV | 1 | 10 | 1 | V | mV | 0.01 | 2 | KHz | 577.36 | 0.014397 | 577.36 | 0.014
| ACV0.000112 | ACV | 1 | 10 | 0.0001 | V | mV | 2 | 20 | KHz | 8.2 | 1.8 | 8.1347 | 1.7537
| ACV0.00112 | ACV | 1 | 10 | 0.001 | V | mV | 2 | 20 | KHz | 8.2 | 1.7525 | 8.1525 | 1.7526
| ACV0.0112 | ACV | 1 | 10 | 0.01 | V | mV | 2 | 20 | KHz | 9.8126 | 1.6513 | 9.799 | 1.6514
| ACV0.112 | ACV | 1 | 10 | 0.1 | V | mV | 2 | 20 | KHz | 58.073 | 0.51337 | 58.064 | 0.51216
| ACV112 | ACV | 1 | 10 | 1 | V | mV | 2 | 20 | KHz | 577.39 | 0.054542 | 577.39 | 0.053976
| ACV0.001100.04 | ACV | 10 | 100 | 0.001 | V | mV | 0.04 | 1 | KHz | 35 | 0.95 | 34.943 | 0.94799
| ACV0.01100.04 | ACV | 10 | 100 | 0.01 | V | mV | 0.04 | 1 | KHz | 35.349 | 0.94557 | 35.282 | 0.94527
| ACV0.1100.04 | ACV | 10 | 100 | 0.1 | V | mV | 0.04 | 1 | KHz | 65.076 | 0.76909 | 64.997 | 0.7675
| ACV1100.04 | ACV | 10 | 100 | 1 | V | mV | 0.04 | 1 | KHz | 577.66 | 0.14257 | 577.63 | 0.14057
| ACV10100.04 | ACV | 10 | 100 | 10 | V | mV | 0.04 | 1 | KHz | 5773.6 | 0.015076 | 5773.6 | 0.01425
| ACV0.011000.04 | ACV | 100 | 1020 | 0.01 | V | mV | 0.04 | 1 | KHz | 360 | 1.1 | 337.15 | 1.1167
| ACV0.11000.04 | ACV | 100 | 1020 | 0.1 | V | mV | 0.04 | 1 | KHz | 360 | 1.1141 | 341.11 | 1.1139
| ACV11000.04 | ACV | 100 | 1020 | 1 | V | mV | 0.04 | 1 | KHz | 639.27 | 0.92932 | 638.46 | 0.92803
| ACV101000.04 | ACV | 100 | 1020 | 10 | V | mV | 0.04 | 1 | KHz | 5773 | 0.18489 | 5772.7 | 0.18294
| ACV1001000.04 | ACV | 100 | 1020 | 100 | V | mV | 0.04 | 1 | KHz | 57736 | 0.019421 | 57735 | 0.018613
| ACA0.0000000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-09 | A | µA | 0.01 | 2 | KHz | 0.47 | 1200 | 0.46304 | 1233.9
| ACA0.000000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.47 | 1234 | 0.46308 | 1233.9
| ACA0.00000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.47 | 1227.5 | 0.4666 | 1227
| ACA0.0000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 0.73981 | 857.29 | 0.73979 | 855.77
| ACA0.000010.000010.01 | ACA | 1e-05 | 0.000104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8088 | 394.38 | 5.8088 | 393.5
| ACA0.000000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-08 | A | µA | 0.01 | 2 | KHz | 0.97 | 1200 | 0.9653 | 1192.1
| ACA0.00000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 0.97 | 1191.5 | 0.96691 | 1191.3
| ACA0.0000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 1.1172 | 1119.7 | 1.1168 | 1118.6
| ACA0.000010.00010.01 | ACA | 0.0001 | 0.00104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 5.8401 | 353.65 | 5.8398 | 351.6
| ACA0.00010.00010.01 | ACA | 0.0001 | 0.00104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 57.912 | 316.08 | 57.912 | 315.2
| ACA0.00000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-07 | A | µA | 0.01 | 2 | KHz | 9.7 | 1200 | 9.653 | 1192.1
| ACA0.0000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 9.7 | 1191.5 | 9.6691 | 1191.3
| ACA0.000010.0010.01 | ACA | 0.001 | 0.0104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 11.172 | 1119.7 | 11.168 | 1118.6
| ACA0.00010.0010.01 | ACA | 0.001 | 0.0104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 58.401 | 353.65 | 58.398 | 351.6
| ACA0.0010.0010.01 | ACA | 0.001 | 0.0104 | 0.001 | A | µA | 0.01 | 2 | KHz | 579.12 | 316.08 | 579.12 | 315.2
| ACA0.0000010.010.01 | ACA | 0.01 | 0.104 | 1e-06 | A | µA | 0.01 | 2 | KHz | 97 | 1200 | 96.53 | 1192.1
| ACA0.000010.010.01 | ACA | 0.01 | 0.104 | 1e-05 | A | µA | 0.01 | 2 | KHz | 97 | 1191.5 | 96.691 | 1191.3
| ACA0.00010.010.01 | ACA | 0.01 | 0.104 | 0.0001 | A | µA | 0.01 | 2 | KHz | 111.72 | 1119.7 | 111.68 | 1118.6
| ACA0.0010.010.01 | ACA | 0.01 | 0.104 | 0.001 | A | µA | 0.01 | 2 | KHz | 583.96 | 353.65 | 583.98 | 351.6
| ACA0.010.010.01 | ACA | 0.01 | 0.104 | 0.01 | A | µA | 0.01 | 2 | KHz | 5791.2 | 316.08 | 5791.2 | 315.2
| ACA0.000010.10.01 | ACA | 0.1 | 1.04 | 1e-05 | A | mA | 0.01 | 2 | KHz | 1 | 1.5 | 0.99589 | 1.489
| ACA0.00010.10.01 | ACA | 0.1 | 1.04 | 0.0001 | A | mA | 0.01 | 2 | KHz | 1 | 1.4884 | 0.99738 | 1.4882
| ACA0.0010.10.01 | ACA | 0.1 | 1.04 | 0.001 | A | mA | 0.01 | 2 | KHz | 1.1829 | 1.4133 | 1.1409 | 1.4124
| ACA0.010.10.01 | ACA | 0.1 | 1.04 | 0.01 | A | mA | 0.01 | 2 | KHz | 5.9116 | 0.47821 | 5.8383 | 0.47621
| ACA0.10.10.01 | ACA | 0.1 | 1.04 | 0.1 | A | mA | 0.01 | 2 | KHz | 57.947 | 0.32961 | 57.91 | 0.32873
| ACA0.000110.01 | ACA | 1 | 10.4 | 0.0001 | A | mA | 0.01 | 2 | KHz | 18 | 1.4 | 17.128 | 1.4001
| ACA0.00110.01 | ACA | 1 | 10.4 | 0.001 | A | mA | 0.01 | 2 | KHz | 18 | 1.3999 | 17.138 | 1.3997
| ACA0.0110.01 | ACA | 1 | 10.4 | 0.01 | A | mA | 0.01 | 2 | KHz | 18.047 | 1.3635 | 18.045 | 1.3626
| ACA0.110.01 | ACA | 1 | 10.4 | 0.1 | A | mA | 0.01 | 2 | KHz | 60.057 | 0.58243 | 60.055 | 0.58052
| ACA110.01 | ACA | 1 | 10.4 | 1 | A | mA | 0.01 | 2 | KHz | 579.27 | 0.34275 | 579.27 | 0.34186
| Capacitance0.0000000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-09 | mF | nF |  |  |  | 0.036 | 6600 | 0.035476 | 6559.3
| Capacitance0.000000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-08 | mF | nF |  |  |  | 0.036 | 6536.7 | 0.035893 | 6534.2
| Capacitance0.00000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-07 | mF | nF |  |  |  | 0.066425 | 5023.2 | 0.066402 | 5012.6
| Capacitance0.0000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-06 | mF | nF |  |  |  | 0.57863 | 1104.8 | 0.57863 | 1095.3
| Capacitance0.000010.000001 | Capacitance | 1e-06 | 1e-05 | 1e-05 | mF | nF |  |  |  | 5.7903 | 3024.4 | 5.7903 | 3020.5
| Capacitance0.000000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-08 | mF | nF |  |  |  | 0.5 | 5800 | 0.48989 | 5801.6
| Capacitance0.00000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-07 | mF | nF |  |  |  | 0.5 | 5788.2 | 0.49306 | 5785.5
| Capacitance0.0000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-06 | mF | nF |  |  |  | 0.74933 | 4690.3 | 0.74914 | 4679.8
| Capacitance0.000010.00001 | Capacitance | 1e-05 | 0.0001 | 1e-05 | mF | nF |  |  |  | 5.797 | 1102.8 | 5.797 | 1093.3
| Capacitance0.00010.00001 | Capacitance | 1e-05 | 0.0001 | 0.0001 | mF | nF |  |  |  | 57.904 | 3024.3 | 57.904 | 3020.4
| Capacitance0.00000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-07 | mF | nF |  |  |  | 6.2 | 5200 | 6.1439 | 5209.5
| Capacitance0.0000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-06 | mF | nF |  |  |  | 6.2 | 5201.1 | 6.1698 | 5198.4
| Capacitance0.000010.0001 | Capacitance | 0.0001 | 0.001 | 1e-05 | mF | nF |  |  |  | 8.3825 | 4378.3 | 8.3808 | 4368.2
| Capacitance0.00010.0001 | Capacitance | 0.0001 | 0.001 | 0.0001 | mF | nF |  |  |  | 58.094 | 1100.5 | 58.094 | 1091
| Capacitance0.0010.0001 | Capacitance | 0.0001 | 0.001 | 0.001 | mF | nF |  |  |  | 579.05 | 3024.2 | 579.05 | 3020.3
| Frequency0.0010.001 | Frequency | 0.001 | 100 | 0.001 | kHz | Hz |  |  |  | 0.87 | 0.016 | 0.81638 | 0.016364
| Frequency0.010.001 | Frequency | 0.001 | 100 | 0.01 | kHz | Hz |  |  |  | 5.8379 | 0.0048237 | 5.8243 | 0.0046826
| Frequency0.10.001 | Frequency | 0.001 | 100 | 0.1 | kHz | Hz |  |  |  | 58.005 | 0.0031699 | 58 | 0.0031065
| Frequency10.001 | Frequency | 0.001 | 100 | 1 | kHz | Hz |  |  |  | 579.98 | 0.026531 | 579.97 | 0.026505
| Frequency100.001 | Frequency | 0.001 | 100 | 10 | kHz | Hz |  |  |  | 5799.7 | 0.26461 | 5799.7 | 0.2646
| Insulation DCV11 | Insulation DCV | 1 | 100 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.4054 | 0.010453
| Insulation DCV101 | Insulation DCV | 1 | 100 | 10 | V | V |  |  |  | 5.9149 | 0.0038256 | 5.9069 | 0.0036668
| Insulation DCV1001 | Insulation DCV | 1 | 100 | 100 | V | V |  |  |  | 57.918 | 0.0033656 | 57.915 | 0.0032952
| Insulation DCV1100 | Insulation DCV | 100 | 250 | 1 | V | V |  |  |  | 1.4 | 0.011 | 1.3333 | 0.011162
| Insulation DCV10100 | Insulation DCV | 100 | 250 | 10 | V | V |  |  |  | 5.6899 | 0.0057678 | 5.6798 | 0.0056527
| Insulation DCV100100 | Insulation DCV | 100 | 250 | 100 | V | V |  |  |  | 57.719 | 0.002458 | 57.716 | 0.0023936
| Insulation DCV1250 | Insulation DCV | 250 | 500 | 1 | V | V |  |  |  | 1.5 | 0.011 | 1.2696 | 0.011414
| Insulation DCV10250 | Insulation DCV | 250 | 500 | 10 | V | V |  |  |  | 5.1032 | 0.0080276 | 5.0881 | 0.007932
| Insulation DCV100250 | Insulation DCV | 250 | 500 | 100 | V | V |  |  |  | 57.497 | 0.0022225 | 57.502 | 0.0021425
| Insulation DCV 1500 | Insulation DCV | 500 | 1000 | 1 | V | V |  |  |  | 1.8 | 0.011 | 1.2242 | 0.011504
| Insulation DCV 10500 | Insulation DCV | 500 | 1000 | 10 | V | V |  |  |  | 4.1693 | 0.0098749 | 4.0987 | 0.0098767
| Insulation DCV 100500 | Insulation DCV | 500 | 1000 | 100 | V | V |  |  |  | 57.066 | 0.0024926 | 57.041 | 0.0024497
| Insulation Res10250 | Insulation Res | 250 | 100000 | 10 | kΩ | kΩ |  |  |  | 18 | 0.0092 | 13.773 | 0.0092419
| Insulation Res100250 | Insulation Res | 250 | 100000 | 100 | kΩ | kΩ |  |  |  | 57.523 | 0.0088223 | 57.453 | 0.0088228
| Insulation Res1000250 | Insulation Res | 250 | 100000 | 1000 | kΩ | kΩ |  |  |  | 576.26 | 0.0052659 | 576.24 | 0.0052655
| Insulation Res10000250 | Insulation Res | 250 | 100000 | 10000 | kΩ | kΩ |  |  |  | 5775 | 0.0010208 | 5775 | 0.0010204
| Insulation Res10100000 | Insulation Res | 100000 | 250000 | 10 | kΩ | kΩ |  |  |  | 14 | 0.011 | 13.511 | 0.010788
| Insulation Res100100000 | Insulation Res | 100000 | 250000 | 100 | kΩ | kΩ |  |  |  | 57.138 | 0.010616 | 57.125 | 0.010616
| Insulation Res1000100000 | Insulation Res | 100000 | 250000 | 1000 | kΩ | kΩ |  |  |  | 575.4 | 0.0087861 | 575.35 | 0.0087861
| Insulation Res10000100000 | Insulation Res | 100000 | 250000 | 10000 | kΩ | kΩ |  |  |  | 5773 | 0.002517 | 5772.9 | 0.0025167
| Insulation Res100000100000 | Insulation Res | 100000 | 250000 | 100000 | kΩ | kΩ |  |  |  | 57896 | 0.0013129 | 57896 | 0.0013128
| Insulation Res10250000 | Insulation Res | 250000 | 500000 | 10 | kΩ | kΩ |  |  |  | 18 | 0.011 | 12.764 | 0.011009
| Insulation Res100250000 | Insulation Res | 250000 | 500000 | 100 | kΩ | kΩ |  |  |  | 55.079 | 0.010925 | 54.77 | 0.010926
| Insulation Res1000250000 | Insulation Res | 250000 | 500000 | 1000 | kΩ | kΩ |  |  |  | 572.67 | 0.0099501 | 572.64 | 0.0099501
| Insulation Res10000250000 | Insulation Res | 250000 | 500000 | 10000 | kΩ | kΩ |  |  |  | 5771.4 | 0.0044676 | 5771.3 | 0.0044674
| Insulation Res10500000 | Insulation Res | 500000 | 1000000 | 10 | kΩ | kΩ |  |  |  | 25 | 0.011 | -76.51 | 0.011211
| Insulation Res100500000 | Insulation Res | 500000 | 1000000 | 100 | kΩ | kΩ |  |  |  | 25 | 0.011209 | -75.391 | 0.011209
| Insulation Res1000500000 | Insulation Res | 500000 | 1000000 | 1000 | kΩ | kΩ |  |  |  | 25 | 0.011045 | 24.396 | 0.011045
| Insulation Res10000500000 | Insulation Res | 500000 | 1000000 | 10000 | kΩ | kΩ |  |  |  | 5114.5 | 0.007343 | 5114.5 | 0.0073429
| Insulation Res100000500000 | Insulation Res | 500000 | 1000000 | 100000 | kΩ | kΩ |  |  |  | 57638 | 0.0014395 | 57638 | 0.0014395
| TcindicatorJ0.1-346 | TcindicatorJ | -346 | -148 | 0.1 | °F | °F |  |  |  | 2.4168 | 0 | 2.4145 | 0
| TcindicatorJ0.1-148 | TcindicatorJ | -148 | 302 | 0.1 | °F | °F |  |  |  | 1.9165 | 0 | 1.9145 | 0
| TcindicatorJ0.1302 | TcindicatorJ | 302 | 1400 | 0.1 | °F | °F |  |  |  | 2.0393 | 0 | 2.0387 | 0
| TcindicatorJ0.11400 | TcindicatorJ | 1400 | 2192 | 0.1 | °F | °F |  |  |  | 2.2469 | 0 | 2.2464 | 0
| TcindicatorJ1-346 | TcindicatorJ | -346 | -148 | 1 | °F | °F |  |  |  | 2.4908 | 0 | 2.4819 | 0
| TcindicatorJ1-148 | TcindicatorJ | -148 | 302 | 1 | °F | °F |  |  |  | 2.0067 | 0 | 1.9988 | 0
| TcindicatorJ1302 | TcindicatorJ | 302 | 1400 | 1 | °F | °F |  |  |  | 2.1202 | 0 | 2.1181 | 0
| TcindicatorJ11400 | TcindicatorJ | 1400 | 2192 | 1 | °F | °F |  |  |  | 2.3207 | 0 | 2.3187 | 0
| TcindicatorJ0.1-210 | TcindicatorJ | -210 | -100 | 0.1 | °C | °C |  |  |  | 1.3461 | 0 | 1.3448 | 0
| TcindicatorJ0.1-100 | TcindicatorJ | -100 | 150 | 0.1 | °C | °C |  |  |  | 1.0665 | 0 | 1.0653 | 0
| TcindicatorJ0.1150 | TcindicatorJ | 150 | 760 | 0.1 | °C | °C |  |  |  | 1.1342 | 0 | 1.1339 | 0
| TcindicatorJ0.1760 | TcindicatorJ | 760 | 1200 | 0.1 | °C | °C |  |  |  | 1.2495 | 0 | 1.2492 | 0
| TcindicatorJ1-210 | TcindicatorJ | -210 | -100 | 1 | °C | °C |  |  |  | 1.467 | 0 | 1.4624 | 0
| TcindicatorJ1-100 | TcindicatorJ | -100 | 150 | 1 | °C | °C |  |  |  | 1.2144 | 0 | 1.2103 | 0
| TcindicatorJ1150 | TcindicatorJ | 150 | 760 | 1 | °C | °C |  |  |  | 1.2722 | 0 | 1.2711 | 0
| TcindicatorJ1760 | TcindicatorJ | 760 | 1200 | 1 | °C | °C |  |  |  | 1.376 | 0 | 1.375 | 0
| TcindicatorK0.1-328 | TcindicatorK | -328 | -148 | 0.1 | °F | °F |  |  |  | 2.5841 | 0 | 2.5813 | 0
| TcindicatorK0.1-148 | TcindicatorK | -148 | 248 | 0.1 | °F | °F |  |  |  | 2.0844 | 0 | 2.0809 | 0
| TcindicatorK0.1248 | TcindicatorK | 248 | 2500 | 0.1 | °F | °F |  |  |  | 2.5391 | 0 | 2.5376 | 0
| TcindicatorK1-328 | TcindicatorK | -328 | -148 | 1 | °F | °F |  |  |  | 2.6555 | 0 | 2.6444 | 0
| TcindicatorK1-148 | TcindicatorK | -148 | 248 | 1 | °F | °F |  |  |  | 2.1723 | 0 | 2.1587 | 0
| TcindicatorK1248 | TcindicatorK | 248 | 2500 | 1 | °F | °F |  |  |  | 2.6075 | 0 | 2.6018 | 0
| TcindicatorK0.1-200 | TcindicatorK | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.4398 | 0 | 1.4382 | 0
| TcindicatorK0.1-100 | TcindicatorK | -100 | 120 | 0.1 | °C | °C |  |  |  | 1.16 | 0 | 1.1581 | 0
| TcindicatorK0.1120 | TcindicatorK | 120 | 1370 | 0.1 | °C | °C |  |  |  | 1.4121 | 0 | 1.4113 | 0
| TcindicatorK1-200 | TcindicatorK | -200 | -100 | 1 | °C | °C |  |  |  | 1.5545 | 0 | 1.5487 | 0
| TcindicatorK1-100 | TcindicatorK | -100 | 120 | 1 | °C | °C |  |  |  | 1.2998 | 0 | 1.2927 | 0
| TcindicatorK1120 | TcindicatorK | 120 | 1370 | 1 | °C | °C |  |  |  | 1.5267 | 0 | 1.5237 | 0
| TcindicatorN0.1-328 | TcindicatorN | -328 | -148 | 0.1 | °F | °F |  |  |  | 3.205 | 0 | 3.2025 | 0
| TcindicatorN0.1-148 | TcindicatorN | -148 | 770 | 0.1 | °F | °F |  |  |  | 2.2918 | 0 | 2.2894 | 0
| TcindicatorN0.1770 | TcindicatorN | 770 | 2372 | 0.1 | °F | °F |  |  |  | 2.4558 | 0 | 2.4545 | 0
| TcindicatorN1-328 | TcindicatorN | -328 | -148 | 1 | °F | °F |  |  |  | 3.2633 | 0 | 3.2536 | 0
| TcindicatorN1-148 | TcindicatorN | -148 | 770 | 1 | °F | °F |  |  |  | 2.3698 | 0 | 2.3604 | 0
| TcindicatorN1770 | TcindicatorN | 770 | 2372 | 1 | °F | °F |  |  |  | 2.5259 | 0 | 2.5208 | 0
| TcindicatorN0.1-200 | TcindicatorN | -200 | -100 | 0.1 | °C | °C |  |  |  | 1.782 | 0 | 1.7806 | 0
| TcindicatorN0.1-100 | TcindicatorN | -100 | 410 | 0.1 | °C | °C |  |  |  | 1.2762 | 0 | 1.2749 | 0
| TcindicatorN0.1410 | TcindicatorN | 410 | 1300 | 0.1 | °C | °C |  |  |  | 1.3658 | 0 | 1.3651 | 0
| TcindicatorN1-200 | TcindicatorN | -200 | -100 | 1 | °C | °C |  |  |  | 1.8762 | 0 | 1.871 | 0
| TcindicatorN1-100 | TcindicatorN | -100 | 410 | 1 | °C | °C |  |  |  | 1.4032 | 0 | 1.3983 | 0
| TcindicatorN1410 | TcindicatorN | 410 | 1300 | 1 | °C | °C |  |  |  | 1.4837 | 0 | 1.481 | 0
| TcindicatorR0.132 | TcindicatorR | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
| TcindicatorR0.1482 | TcindicatorR | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5823 | 0 | 3.5805 | 0
| TcindicatorR132 | TcindicatorR | 32 | 482 | 1 | °F | °F |  |  |  | 4.8335 | 0 | 4.8271 | 0
| TcindicatorR1482 | TcindicatorR | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6336 | 0 | 3.6263 | 0
| TcindicatorR0.10 | TcindicatorR | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6784 | 0 | 2.6775 | 0
| TcindicatorR0.1250 | TcindicatorR | 250 | 1760 | 0.1 | °C | °C |  |  |  | 1.9965 | 0 | 1.9955 | 0
| TcindicatorR10 | TcindicatorR | 0 | 250 | 1 | °C | °C |  |  |  | 2.7419 | 0 | 2.7384 | 0
| TcindicatorR1250 | TcindicatorR | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0804 | 0 | 2.0765 | 0
| TcindicatorS0.132 | TcindicatorS | 32 | 482 | 0.1 | °F | °F |  |  |  | 4.7944 | 0 | 4.7928 | 0
| TcindicatorS0.1482 | TcindicatorS | 482 | 3200 | 0.1 | °F | °F |  |  |  | 3.5846 | 0 | 3.5828 | 0
| TcindicatorS132 | TcindicatorS | 32 | 482 | 1 | °F | °F |  |  |  | 4.8334 | 0 | 4.8271 | 0
| TcindicatorS1482 | TcindicatorS | 482 | 3200 | 1 | °F | °F |  |  |  | 3.6358 | 0 | 3.6285 | 0
| TcindicatorS0.10 | TcindicatorS | 0 | 250 | 0.1 | °C | °C |  |  |  | 2.6783 | 0 | 2.6775 | 0
| TcindicatorS0.1250 | TcindicatorS | 250 | 1760 | 0.1 | °C | °C |  |  |  | 2.0006 | 0 | 1.9996 | 0
| TcindicatorS10 | TcindicatorS | 0 | 250 | 1 | °C | °C |  |  |  | 2.7418 | 0 | 2.7384 | 0
| TcindicatorS1250 | TcindicatorS | 250 | 1760 | 1 | °C | °C |  |  |  | 2.0844 | 0 | 2.0805 | 0
| TcindicatorT0.1-418 | TcindicatorT | -418 | -238 | 0.1 | °F | °F |  |  |  | 3.9624 | 0 | 3.9602 | 0
| TcindicatorT0.1-238 | TcindicatorT | -238 | 752 | 0.1 | °F | °F |  |  |  | 1.9193 | 0 | 1.9156 | 0
| TcindicatorT1-418 | TcindicatorT | -418 | -238 | 1 | °F | °F |  |  |  | 4.0103 | 0 | 4.0017 | 0
| TcindicatorT1-238 | TcindicatorT | -238 | 752 | 1 | °F | °F |  |  |  | 2.0141 | 0 | 1.9999 | 0
| TcindicatorT0.1-250 | TcindicatorT | -250 | -150 | 0.1 | °C | °C |  |  |  | 2.2146 | 0 | 2.2134 | 0
| TcindicatorT0.1-150 | TcindicatorT | -150 | 400 | 0.1 | °C | °C |  |  |  | 1.0694 | 0 | 1.0674 | 0
| TcindicatorT1-250 | TcindicatorT | -250 | -150 | 1 | °C | °C |  |  |  | 2.2914 | 0 | 2.2868 | 0
| TcindicatorT1-150 | TcindicatorT | -150 | 400 | 1 | °C | °C |  |  |  | 1.2194 | 0 | 1.2121 | 0

## Software Validation
Date: | Action
2015-12-02 | Validated all functions and calculations, repeat readings, uncertainty derivations from table.
2017-02-16 | Updated Page numbering, changed calculation to mark unaccredited readings with an asterisk